## CERTIFICATION WORKSHEET
Pages 1-5 |  | Report Number: | 072513A |  |  | password: jgi
ITEM |  |  | CONTROLLING SPECS | STANDARD VOLTMETER | TOLERANCE SOURCE |  | AMS2750D | TRUE | TRUE | 1 | FALSE | TRUE
[7 empty] | BAC5621K | FALSE | FALSE | 2 | FALSE | TRUE
|  |  |  | Additel 286 SN:6851022030037 |  |  | PS23401 | FALSE | FALSE | 3 |  | FALSE
[7 empty] | ASTME-230 | TRUE |  | 1 |  | FALSE
[7 empty] | BAC5602AF | FALSE
[7 empty] | BPS4139M | FALSE |  | 1
[7 empty] | BAERD GEN-007 | FALSE
[5 empty] | Written Procedure(Other than JGI) |  | OTHER | FALSE
Thermocouple Wire Probe |  |  |  | METER/SN
|  |  |  | CAL: | COMMENTS SECTION
DUE:
T/N:
TEMP | 72 |  | TOLERANCE
%RH | 25
Wire Gauge |  | (If not "N/A") |  |  | Additional Comments
Inspector: | KF
Location |  | (If not "Production")
CAL DATE | 2025-04-09 | (If not today's date)
DUE DATE |  | (If not "N/A")
±1.0°F or 0.1% of reading
Note on F29: DO NOT USE THE SIGN FOR "Greater than or equal to"
ITEM
Thermocouple Test/Working Sensor(s)
Thermocouple Probe(s)
Thermocouple Wire
Thermocouple Surface Probe(s)
Other (Specify Below)
Greater of ±2.0°F or ±0.4% of Reading (AMS2750,PS23401,BAC5602,BPS4139,BAERD GEN)
Greater of ±4°F or ±0.75% of Reading (AMS2750)
</= 1000°F ±2.0°F, >1000°F ±0.4% (BAC5621L)
Greater of ±1°F or ±0.1% of Reading (AMS2750)
Type T: ±1.8°F / Type J,K,N: ±4.0°F (ASTME-230)
OTHER (Enter Below)
Tolerance
Special Limits
Standard Limits
NEW
USED
New/Used
Tolerance Source
Customer
Specification
Manufacturer
JGI Recommended
Found In Tolerance. No Action Required.
Comments Section
Found out of Tolerance.
Unable to calibrate due to a break in continuity.
OTHER.
AMS2750H
BAC5621L
PS23401
ASTME220-19
BAC5602AF
BPS4139M
BAERD GEN-007D
OTHER(s)

## CT
1 | Test Temperature |  |  | -50 |  |  |  | 2 | Test Temperature |  |  | 200 |  |  |  | 3 | Test Temperature |  |  | 450 |  |  |  | 4 | Test Temperature |  |  |  |  | 5 | Test Temperature |  |  |  |  | 6 | Test Temperature |  |  |  |  | 7 | Test Temperature |  |  |  |  | 8 | Test Temperature |  |  |  |  | 9 | Test Temperature |  |  |  |  | 10 | Test Temperature |  |  |  |  | 11 | Test Temperature |  |  |  | 12 | Test Temperature |  |  |  |  | 13 | Test Temperature |  |  |  |  | 14 | Test Temperature |  |  |  |  | 15 | Test Temperature |  |  |  |  | 16 | Test Temperature |  |  |  |  | 17 | Test Temperature |  |  |  |  | 18 | Test Temperature |  |  |  |  | 19 | Test Temperature |  |  |  |  | 20 | Test Temperature
PAGE INFO/INSTRUCTIONS
| STANDARD(s) USED: |  | PRT | PRT | Standard |  |  |  | STANDARD(s) USED: |  | PRT | PRT | Standard |  |  |  | STANDARD(s) USED: |  | PRT | PRT | Standard |  |  |  | STANDARD(s) USED: |  |  | Standard |  |  | STANDARD(s) USED: |  |  | Standard |  |  | STANDARD(s) USED: |  |  | Standard |  |  | STANDARD(s) USED: |  |  | Standard |  |  | STANDARD(s) USED: |  |  | Standard |  |  | STANDARD(s) USED: |  |  | Standard |  |  | STANDARD(s) USED: |  |  | Standard |  |  | STANDARD(s) USED: |  |  | Standard |  | STANDARD(s) USED: |  |  |  | Standard |  | STANDARD(s) USED: |  |  |  | Standard |  | STANDARD(s) USED: |  |  |  | Standard |  | STANDARD(s) USED: |  |  |  | Standard |  | STANDARD(s) USED: |  |  |  | Standard |  | STANDARD(s) USED: |  |  |  | Standard |  | STANDARD(s) USED: |  |  |  | Standard |  | STANDARD(s) USED: |  |  |  | Standard |  | STANDARD(s) USED: |  |  |  | Standard
| Standard mvDC |  |  | -50 | -50 |  |  |  | Standard mvDC |  |  | 200 | 200 |  |  |  | Standard mvDC |  |  | 450 | 450 |  |  |  | Standard mvDC |  |  | #N/A |  |  | Standard mvDC |  |  | #N/A |  |  | Standard mvDC |  |  | #N/A |  |  | Standard mvDC |  |  | #N/A |  |  | Standard mvDC |  |  | #N/A |  |  | Standard mvDC |  |  | #N/A |  |  | Standard mvDC |  |  | #N/A |  |  | Standard mvDC |  |  | #N/A |  | Standard mvDC |  |  |  | #N/A |  | Standard mvDC |  |  |  | #N/A |  | Standard mvDC |  |  |  | #N/A |  | Standard mvDC |  |  |  | #N/A |  | Standard mvDC |  |  |  | #N/A |  | Standard mvDC |  |  |  | #N/A |  | Standard mvDC |  |  |  | #N/A |  | Standard mvDC |  |  |  | #N/A |  | Standard mvDC |  |  |  | #N/A
|  |  | T/C |  | Unit Under Test (°F) | 1 |  |  |  |  | T/C |  | Unit Under Test (°F) | 1 |  |  |  |  | T/C |  | Unit Under Test (°F) | 1 |  |  |  |  | T/C | Unit Under Test (°F) | 1 |  |  |  | T/C | Unit Under Test (°F) | 1 |  |  |  | T/C | Unit Under Test (°F) | 1 |  |  |  | T/C | Unit Under Test (°F) | 1 |  |  |  | T/C | Unit Under Test (°F) | 1 |  |  |  | T/C | Unit Under Test (°F) | 1 |  |  |  | T/C | Unit Under Test (°F) | 1 |  |  |  | T/C | Unit Under Test (°F) |  |  |  | T/C |  | Unit Under Test (°F) |  |  |  | T/C |  | Unit Under Test (°F) |  |  |  | T/C |  | Unit Under Test (°F) |  |  |  | T/C |  | Unit Under Test (°F) |  |  |  | T/C |  | Unit Under Test (°F) |  |  |  | T/C |  | Unit Under Test (°F) |  |  |  | T/C |  | Unit Under Test (°F) |  |  |  | T/C |  | Unit Under Test (°F) |  |  |  | T/C |  | Unit Under Test (°F)
Pg #s |  | Serial Number | Type | [6 empty] | Serial Number | Type | [6 empty] | Serial Number | Type | [6 empty] | Serial Number | Type |  |  |  |  | Serial Number | Type |  |  |  |  | Serial Number | Type |  |  |  |  | Serial Number | Type |  |  |  |  | Serial Number | Type |  |  |  |  | Serial Number | Type |  |  |  |  | Serial Number | Type |  |  |  |  | Serial Number | Type |  |  |  | Serial Number | Type | mvDC |  |  |  | Serial Number | Type | mvDC |  |  |  | Serial Number | Type | mvDC |  |  |  | Serial Number | Type | mvDC |  |  |  | Serial Number | Type | mvDC |  |  |  | Serial Number | Type | mvDC |  |  |  | Serial Number | Type | mvDC |  |  |  | Serial Number | Type | mvDC |  |  |  | Serial Number | Type | mvDC
21 | 1 | 072513A | J |  | 32 | -51.1 | 1.1 | 21 | 1 | 072513A | J |  | 32 | 199.1 | 0.9 | 21 | 1 | 072513A | J |  | 32 | 448.9 | 1.1 | 21 | 1 | 072513A | J | 32 |  | 21 | 1 | 072513A | J | 32 |  | 21 | 5 |  | J | 32 |  | 21 | 5 |  | J | 32 |  | 21 | 5 |  | J | 32 |  | 21 | 5 |  | J | 32 |  | 21 | 5 |  | J | 32 |  | 21 | 9 |  | J | 32 | 21 | 9 |  | J |  | 32 | 21 | 9 |  | J |  | 32 | 21 | 9 |  | J |  | 32 | 21 | 9 |  | J |  | 32 | 21 | 13 |  | J |  | 32 | 21 | 13 |  | J |  | 32 | 21 | 13 |  | J |  | 32 | 21 | 13 |  | J |  | 32 | 21 | 13 |  | J |  | 32
22 | 1 |  | J |  | 32 | -51 | 1 | 22 | 1 |  | J |  | 32 | 199.1 | 0.9 | 22 | 1 |  | J |  | 32 | 449 | 1 | 22 | 1 |  | J | 32 |  | 22 | 1 |  | J | 32 |  | 22 | 5 |  | J | 32 |  | 22 | 5 |  | J | 32 |  | 22 | 5 |  | J | 32 |  | 22 | 5 |  | J | 32 |  | 22 | 5 |  | J | 32 |  | 22 | 9 |  | J | 32 | 22 | 9 |  | J |  | 32 | 22 | 9 |  | J |  | 32 | 22 | 9 |  | J |  | 32 | 22 | 9 |  | J |  | 32 | 22 | 13 |  | J |  | 32 | 22 | 13 |  | J |  | 32 | 22 | 13 |  | J |  | 32 | 22 | 13 |  | J |  | 32 | 22 | 13 |  | J |  | 32
23 | 1 |  |  |  | #N/A |  |  | 23 | 1 |  |  |  | #N/A |  |  | 23 | 1 |  |  |  | #N/A |  |  | 23 | 1 |  |  | #N/A |  | 23 | 1 |  |  | #N/A |  | 23 | 5 |  |  | #N/A |  | 23 | 5 |  |  | #N/A |  | 23 | 5 |  |  | #N/A |  | 23 | 5 |  |  | #N/A |  | 23 | 5 |  |  | #N/A |  | 23 | 9 |  |  | #N/A | 23 | 9 |  |  |  | #N/A | 23 | 9 |  |  |  | #N/A | 23 | 9 |  |  |  | #N/A | 23 | 9 |  |  |  | #N/A | 23 | 13 |  |  |  | #N/A | 23 | 13 |  |  |  | #N/A | 23 | 13 |  |  |  | #N/A | 23 | 13 |  |  |  | #N/A | 23 | 13 |  |  |  | #N/A
24 | 1 |  |  |  | #N/A |  |  | 24 | 1 |  |  |  | #N/A |  |  | 24 | 1 |  |  |  | #N/A |  |  | 24 | 1 |  |  | #N/A |  | 24 | 1 |  |  | #N/A |  | 24 | 5 |  |  | #N/A |  | 24 | 5 |  |  | #N/A |  | 24 | 5 |  |  | #N/A |  | 24 | 5 |  |  | #N/A |  | 24 | 5 |  |  | #N/A |  | 24 | 9 |  |  | #N/A | 24 | 9 |  |  |  | #N/A | 24 | 9 |  |  |  | #N/A | 24 | 9 |  |  |  | #N/A | 24 | 9 |  |  |  | #N/A | 24 | 13 |  |  |  | #N/A | 24 | 13 |  |  |  | #N/A | 24 | 13 |  |  |  | #N/A | 24 | 13 |  |  |  | #N/A | 24 | 13 |  |  |  | #N/A
25 | 1 |  |  |  | #N/A |  |  | 25 | 1 |  |  |  | #N/A |  |  | 25 | 1 |  |  |  | #N/A |  |  | 25 | 1 |  |  | #N/A |  | 25 | 1 |  |  | #N/A |  | 25 | 5 |  |  | #N/A |  | 25 | 5 |  |  | #N/A |  | 25 | 5 |  |  | #N/A |  | 25 | 5 |  |  | #N/A |  | 25 | 5 |  |  | #N/A |  | 25 | 9 |  |  | #N/A | 25 | 9 |  |  |  | #N/A | 25 | 9 |  |  |  | #N/A | 25 | 9 |  |  |  | #N/A | 25 | 9 |  |  |  | #N/A | 25 | 13 |  |  |  | #N/A | 25 | 13 |  |  |  | #N/A | 25 | 13 |  |  |  | #N/A | 25 | 13 |  |  |  | #N/A | 25 | 13 |  |  |  | #N/A
26 | 2 |  |  |  | #N/A |  |  | 26 | 2 |  |  |  | #N/A |  |  | 26 | 2 |  |  |  | #N/A |  |  | 26 | 2 |  |  | #N/A |  | 26 | 2 |  |  | #N/A |  | 26 | 6 |  |  | #N/A |  | 26 | 6 |  |  | #N/A |  | 26 | 6 |  |  | #N/A |  | 26 | 6 |  |  | #N/A |  | 26 | 6 |  |  | #N/A |  | 26 | 10 |  |  | #N/A | 26 | 10 |  |  |  | #N/A | 26 | 10 |  |  |  | #N/A | 26 | 10 |  |  |  | #N/A | 26 | 10 |  |  |  | #N/A | 26 | 14 |  |  |  | #N/A | 26 | 14 |  |  |  | #N/A | 26 | 14 |  |  |  | #N/A | 26 | 14 |  |  |  | #N/A | 26 | 14 |  |  |  | #N/A
27 | 2 |  |  |  | #N/A |  |  | 27 | 2 |  |  |  | #N/A |  |  | 27 | 2 |  |  |  | #N/A |  |  | 27 | 2 |  |  | #N/A |  | 27 | 2 |  |  | #N/A |  | 27 | 6 |  |  | #N/A |  | 27 | 6 |  |  | #N/A |  | 27 | 6 |  |  | #N/A |  | 27 | 6 |  |  | #N/A |  | 27 | 6 |  |  | #N/A |  | 27 | 10 |  |  | #N/A | 27 | 10 |  |  |  | #N/A | 27 | 10 |  |  |  | #N/A | 27 | 10 |  |  |  | #N/A | 27 | 10 |  |  |  | #N/A | 27 | 14 |  |  |  | #N/A | 27 | 14 |  |  |  | #N/A | 27 | 14 |  |  |  | #N/A | 27 | 14 |  |  |  | #N/A | 27 | 14 |  |  |  | #N/A
28 | 2 |  |  |  | #N/A |  |  | 28 | 2 |  |  |  | #N/A |  |  | 28 | 2 |  |  |  | #N/A |  |  | 28 | 2 |  |  | #N/A |  | 28 | 2 |  |  | #N/A |  | 28 | 6 |  |  | #N/A |  | 28 | 6 |  |  | #N/A |  | 28 | 6 |  |  | #N/A |  | 28 | 6 |  |  | #N/A |  | 28 | 6 |  |  | #N/A |  | 28 | 10 |  |  | #N/A | 28 | 10 |  |  |  | #N/A | 28 | 10 |  |  |  | #N/A | 28 | 10 |  |  |  | #N/A | 28 | 10 |  |  |  | #N/A | 28 | 14 |  |  |  | #N/A | 28 | 14 |  |  |  | #N/A | 28 | 14 |  |  |  | #N/A | 28 | 14 |  |  |  | #N/A | 28 | 14 |  |  |  | #N/A
29 | 2 |  |  |  | #N/A |  |  | 29 | 2 |  |  |  | #N/A |  |  | 29 | 2 |  |  |  | #N/A |  |  | 29 | 2 |  |  | #N/A |  | 29 | 2 |  |  | #N/A |  | 29 | 6 |  |  | #N/A |  | 29 | 6 |  |  | #N/A |  | 29 | 6 |  |  | #N/A |  | 29 | 6 |  |  | #N/A |  | 29 | 6 |  |  | #N/A |  | 29 | 10 |  |  | #N/A | 29 | 10 |  |  |  | #N/A | 29 | 10 |  |  |  | #N/A | 29 | 10 |  |  |  | #N/A | 29 | 10 |  |  |  | #N/A | 29 | 14 |  |  |  | #N/A | 29 | 14 |  |  |  | #N/A | 29 | 14 |  |  |  | #N/A | 29 | 14 |  |  |  | #N/A | 29 | 14 |  |  |  | #N/A
30 | 2 |  |  |  | #N/A |  |  | 30 | 2 |  |  |  | #N/A |  |  | 30 | 2 |  |  |  | #N/A |  |  | 30 | 2 |  |  | #N/A |  | 30 | 2 |  |  | #N/A |  | 30 | 6 |  |  | #N/A |  | 30 | 6 |  |  | #N/A |  | 30 | 6 |  |  | #N/A |  | 30 | 6 |  |  | #N/A |  | 30 | 6 |  |  | #N/A |  | 30 | 10 |  |  | #N/A | 30 | 10 |  |  |  | #N/A | 30 | 10 |  |  |  | #N/A | 30 | 10 |  |  |  | #N/A | 30 | 10 |  |  |  | #N/A | 30 | 14 |  |  |  | #N/A | 30 | 14 |  |  |  | #N/A | 30 | 14 |  |  |  | #N/A | 30 | 14 |  |  |  | #N/A | 30 | 14 |  |  |  | #N/A
31 | 3 |  |  |  | #N/A |  |  | 31 | 3 |  |  |  | #N/A |  |  | 31 | 3 |  |  |  | #N/A |  |  | 31 | 3 |  |  | #N/A |  | 31 | 3 |  |  | #N/A |  | 31 | 7 |  |  | #N/A |  | 31 | 7 |  |  | #N/A |  | 31 | 7 |  |  | #N/A |  | 31 | 7 |  |  | #N/A |  | 31 | 7 |  |  | #N/A |  | 31 | 11 |  |  | #N/A | 31 | 11 |  |  |  | #N/A | 31 | 11 |  |  |  | #N/A | 31 | 11 |  |  |  | #N/A | 31 | 11 |  |  |  | #N/A | 31 | 15 |  |  |  | #N/A | 31 | 15 |  |  |  | #N/A | 31 | 15 |  |  |  | #N/A | 31 | 15 |  |  |  | #N/A | 31 | 15 |  |  |  | #N/A
32 | 3 |  |  |  | #N/A |  |  | 32 | 3 |  |  |  | #N/A |  |  | 32 | 3 |  |  |  | #N/A |  |  | 32 | 3 |  |  | #N/A |  | 32 | 3 |  |  | #N/A |  | 32 | 7 |  |  | #N/A |  | 32 | 7 |  |  | #N/A |  | 32 | 7 |  |  | #N/A |  | 32 | 7 |  |  | #N/A |  | 32 | 7 |  |  | #N/A |  | 32 | 11 |  |  | #N/A | 32 | 11 |  |  |  | #N/A | 32 | 11 |  |  |  | #N/A | 32 | 11 |  |  |  | #N/A | 32 | 11 |  |  |  | #N/A | 32 | 15 |  |  |  | #N/A | 32 | 15 |  |  |  | #N/A | 32 | 15 |  |  |  | #N/A | 32 | 15 |  |  |  | #N/A | 32 | 15 |  |  |  | #N/A
33 | 3 |  |  |  | #N/A |  |  | 33 | 3 |  |  |  | #N/A |  |  | 33 | 3 |  |  |  | #N/A |  |  | 33 | 3 |  |  | #N/A |  | 33 | 3 |  |  | #N/A |  | 33 | 7 |  |  | #N/A |  | 33 | 7 |  |  | #N/A |  | 33 | 7 |  |  | #N/A |  | 33 | 7 |  |  | #N/A |  | 33 | 7 |  |  | #N/A |  | 33 | 11 |  |  | #N/A | 33 | 11 |  |  |  | #N/A | 33 | 11 |  |  |  | #N/A | 33 | 11 |  |  |  | #N/A | 33 | 11 |  |  |  | #N/A | 33 | 15 |  |  |  | #N/A | 33 | 15 |  |  |  | #N/A | 33 | 15 |  |  |  | #N/A | 33 | 15 |  |  |  | #N/A | 33 | 15 |  |  |  | #N/A
34 | 3 |  |  |  | #N/A |  |  | 34 | 3 |  |  |  | #N/A |  |  | 34 | 3 |  |  |  | #N/A |  |  | 34 | 3 |  |  | #N/A |  | 34 | 3 |  |  | #N/A |  | 34 | 7 |  |  | #N/A |  | 34 | 7 |  |  | #N/A |  | 34 | 7 |  |  | #N/A |  | 34 | 7 |  |  | #N/A |  | 34 | 7 |  |  | #N/A |  | 34 | 11 |  |  | #N/A | 34 | 11 |  |  |  | #N/A | 34 | 11 |  |  |  | #N/A | 34 | 11 |  |  |  | #N/A | 34 | 11 |  |  |  | #N/A | 34 | 15 |  |  |  | #N/A | 34 | 15 |  |  |  | #N/A | 34 | 15 |  |  |  | #N/A | 34 | 15 |  |  |  | #N/A | 34 | 15 |  |  |  | #N/A
35 | 3 |  |  |  | #N/A |  |  | 35 | 3 |  |  |  | #N/A |  |  | 35 | 3 |  |  |  | #N/A |  |  | 35 | 3 |  |  | #N/A |  | 35 | 3 |  |  | #N/A |  | 35 | 7 |  |  | #N/A |  | 35 | 7 |  |  | #N/A |  | 35 | 7 |  |  | #N/A |  | 35 | 7 |  |  | #N/A |  | 35 | 7 |  |  | #N/A |  | 35 | 11 |  |  | #N/A | 35 | 11 |  |  |  | #N/A | 35 | 11 |  |  |  | #N/A | 35 | 11 |  |  |  | #N/A | 35 | 11 |  |  |  | #N/A | 35 | 15 |  |  |  | #N/A | 35 | 15 |  |  |  | #N/A | 35 | 15 |  |  |  | #N/A | 35 | 15 |  |  |  | #N/A | 35 | 15 |  |  |  | #N/A
36 | 4 |  |  |  | #N/A |  |  | 36 | 4 |  |  |  | #N/A |  |  | 36 | 4 |  |  |  | #N/A |  |  | 36 | 4 |  |  | #N/A |  | 36 | 4 |  |  | #N/A |  | 36 | 8 |  |  | #N/A |  | 36 | 8 |  |  | #N/A |  | 36 | 8 |  |  | #N/A |  | 36 | 8 |  |  | #N/A |  | 36 | 8 |  |  | #N/A |  | 36 | 12 |  |  | #N/A | 36 | 12 |  |  |  | #N/A | 36 | 12 |  |  |  | #N/A | 36 | 12 |  |  |  | #N/A | 36 | 12 |  |  |  | #N/A | 36 | 16 |  |  |  | #N/A | 36 | 16 |  |  |  | #N/A | 36 | 16 |  |  |  | #N/A | 36 | 16 |  |  |  | #N/A | 36 | 16 |  |  |  | #N/A
37 | 4 |  |  |  | #N/A |  |  | 37 | 4 |  |  |  | #N/A |  |  | 37 | 4 |  |  |  | #N/A |  |  | 37 | 4 |  |  | #N/A |  | 37 | 4 |  |  | #N/A |  | 37 | 8 |  |  | #N/A |  | 37 | 8 |  |  | #N/A |  | 37 | 8 |  |  | #N/A |  | 37 | 8 |  |  | #N/A |  | 37 | 8 |  |  | #N/A |  | 37 | 12 |  |  | #N/A | 37 | 12 |  |  |  | #N/A | 37 | 12 |  |  |  | #N/A | 37 | 12 |  |  |  | #N/A | 37 | 12 |  |  |  | #N/A | 37 | 16 |  |  |  | #N/A | 37 | 16 |  |  |  | #N/A | 37 | 16 |  |  |  | #N/A | 37 | 16 |  |  |  | #N/A | 37 | 16 |  |  |  | #N/A
38 | 4 |  |  |  | #N/A |  |  | 38 | 4 |  |  |  | #N/A |  |  | 38 | 4 |  |  |  | #N/A |  |  | 38 | 4 |  |  | #N/A |  | 38 | 4 |  |  | #N/A |  | 38 | 8 |  |  | #N/A |  | 38 | 8 |  |  | #N/A |  | 38 | 8 |  |  | #N/A |  | 38 | 8 |  |  | #N/A |  | 38 | 8 |  |  | #N/A |  | 38 | 12 |  |  | #N/A | 38 | 12 |  |  |  | #N/A | 38 | 12 |  |  |  | #N/A | 38 | 12 |  |  |  | #N/A | 38 | 12 |  |  |  | #N/A | 38 | 16 |  |  |  | #N/A | 38 | 16 |  |  |  | #N/A | 38 | 16 |  |  |  | #N/A | 38 | 16 |  |  |  | #N/A | 38 | 16 |  |  |  | #N/A
39 | 4 |  |  |  | #N/A |  |  | 39 | 4 |  |  |  | #N/A |  |  | 39 | 4 |  |  |  | #N/A |  |  | 39 | 4 |  |  | #N/A |  | 39 | 4 |  |  | #N/A |  | 39 | 8 |  |  | #N/A |  | 39 | 8 |  |  | #N/A |  | 39 | 8 |  |  | #N/A |  | 39 | 8 |  |  | #N/A |  | 39 | 8 |  |  | #N/A |  | 39 | 12 |  |  | #N/A | 39 | 12 |  |  |  | #N/A | 39 | 12 |  |  |  | #N/A | 39 | 12 |  |  |  | #N/A | 39 | 12 |  |  |  | #N/A | 39 | 16 |  |  |  | #N/A | 39 | 16 |  |  |  | #N/A | 39 | 16 |  |  |  | #N/A | 39 | 16 |  |  |  | #N/A | 39 | 16 |  |  |  | #N/A
40 | 4 |  |  |  | #N/A |  |  | 40 | 4 |  |  |  | #N/A |  |  | 40 | 4 |  |  |  | #N/A |  |  | 40 | 4 |  |  | #N/A |  | 40 | 4 |  |  | #N/A |  | 40 | 8 |  |  | #N/A |  | 40 | 8 |  |  | #N/A |  | 40 | 8 |  |  | #N/A |  | 40 | 8 |  |  | #N/A |  | 40 | 8 |  |  | #N/A |  | 40 | 12 |  |  | #N/A | 40 | 12 |  |  |  | #N/A | 40 | 12 |  |  |  | #N/A | 40 | 12 |  |  |  | #N/A | 40 | 12 |  |  |  | #N/A | 40 | 16 |  |  |  | #N/A | 40 | 16 |  |  |  | #N/A | 40 | 16 |  |  |  | #N/A | 40 | 16 |  |  |  | #N/A | 40 | 16 |  |  |  | #N/A
Note on C2: The thermocouple type fields colored green represent the selected thermocouple type for that page. So for page 1, the first T/C type field is green, as well as for page 2, etc. Unless you're going to put individual thermocouples on individual certs, group thermocouple types together by the color coded/numbered cells to the left of the SN field. Wire can be done (beginning and end) using the first four slots. They will appear as two beginning and end rolls on pages 17 and 18 TUS Probes Print from pg 41 For Dual Element probes, Use 1st four slots, Pgs 19&20
Note on E3: If the standard you want isn't displayed on the left when you select your temperature, insert the SN of the correct standard here. Otherwise, leave blank.
Note on AK3: Leave blank if you're using the transfer standards. If using the PRT, SP-1, or ST-6, enter here.
Note on AS3: Leave blank if you're using the transfer standards. If using the PRT, SP-1, or ST-6, enter here.
Note on BY3: Leave blank if you're using the transfer standards. If using the PRT, SP-1, or ST-6, enter here.
Note on DU3: Leave blank if you're using the transfer standards. If using the PRT, SP-1, or ST-6, enter here.
Note on G7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on O7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on W7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on AE7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on AM7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on AU7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on BC7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on BK7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on BS7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on CA7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on CI7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on CQ7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on CY7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on DG7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on DO7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on DW7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on EE7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on EM7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on EU7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert
Note on FC7: These cells show the UUT temperature after comparison with the standard, and adjusted to the setpoint, as they will appear on JGI Cert

## TC FORM
Report: |  |  | PAGE | 1 |  |  |  |  | Voltmeter Used: |  | 34420 | [5 empty] | Standard(s) Used | PRTPRTPRT | PRTPRTPRT | [18 empty] | Thermocouple Type |  | "J" |  |  |  | Thermocouple Test/Working Sensor |  |  |  | COMMENTS AND RESULTS SECTION
Customer: | JOHNSON GAGE & INSPECTION, INC. | [32 empty] | SPECS |  |  |  | Limits of Error (If Applicable) |  |  |  | " |  | Thermocouple Probe | Temp: | 72 | 72 | Function for "Contains" is : | =IF(ISERROR(SEARCH("Note",BH9)),"Miss","Hit")
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF |  |  |  | Cells in BOLD font are taken from update form | [23 empty] | Enter spec/tol. |  |  | pg1 | J | "J" | 35 | 0 | "0" |  | Thermocouple Wire | RH: | 25 | 25 |  |  | ACCEPTED
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09 | [24 empty] | Supplimental specs here: | [5 empty] | pg2 | 0 | "0" | 36 | 0 | "0" |  | Thermocouple Surface Probe |  |  | 68
Gauge: | N/A |  |  | Date Due: |  | N/A |  |  | Serial Number | Shorthand | Merged | Calibration Date | Date Due | Traceability | Type |  |  | 46756, | SP30124, | 2023-12-15 | 2023-05-16 | 2024-01-19 | 2024-01-19 | Progression of Cal date and Due Date formulas |  |  |  |  | Date Cal Field 2 | [7 empty] | pg3 | 0 | "0" | 37 | 0 | "0" |  | Other |  |  | 46 |  |  | ACCEPTED
Temp: | 72 |  |  | Relative Humidity: |  | 25 | % |  | 2823A08926 | 3458 | H.P. 3458A SN:2823A08926 | 2024-05-08 | 2025-05-08 | WO-00790034 | H.P. 3458A |  |  | G295632, | SP10124, | 2025-12-15 | 2028-05-16 | 2025-01-19 | 2025-01-19 | Date Cal field 1 |  |  |  |  | IF |  |  | AMS2750H | Greater of ±2.0°F or ±0.4% of Reading | TRUE | AMS2750H | Greater of ±2.0°F or ±0.4% of Reading | pg4 | 0 | "0" | 38 | 0 | "0" |  |  | DATE AND DATE DUE |  |  | Comments Shorthand
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading |  |  |  | B55852 | 1502 | FLUKE 1502 SN:B55852 | 2024-01-15 | 2025-01-15 | 56561-053313 | FLUKE 1502 |  |  |  |  | SN USED |  |  |  | IF |  |  |  |  | ST6 is GO |  |  | BAC5621L | <1000°F:±2.0°F,>/= 1000°F: ±0.4% Rdg. | FALSE |  |  | 21 | J | "J" | 39 | 0 | "0" |  |  |  | 2025-04-09 |  | 0 | Found in tolerance. No action required. "C.F" = "Correction Factor." | ACCEPTED
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19 |  |  | 6851022030037 | ADT286 | Additel 286 SN:6851022030037 | 2024-07-01 | 2025-07-01 | 56561-060578 | Additel 286 |  |  | PRT | PRT SN:496142 | PRT SN:496142 | 2023-12-15 |  |  | PRT is GO |  |  |  |  | AND Any thing else is GO |  |  | PS23401 | Greater of ±2.0°F or ±0.4% of Reading | FALSE |  |  | 22 | J | "J" | 40 | 0 | "0" |  |  |  | 2025-06-04 |  | Additional Comments:
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below | [10 empty] | TTS-1 | TTS-1 |  |  | 6 | BLANK | then 7-8-09 |  |  |  |  | then 9-26-08 |  |  | ASTM E220-19 | Type T: ±1.8°F / Type J,K,N: ±4.0°F | TRUE | , ASTM E220-19 | Type T: ±1.8°F / Type J,K,N: ±4.0°F | 23 | 0 | "0" | pg5 |  |  |  |  | Enter Date (if not current) | 2025-04-09 | 2025-04-09 | C.F.=Correction Factor. |  | ACCEPTED AS NOTED
Test Temperatures |  |  | -50 | 200 | 450 | [12 empty] | SP-3 | SP-3 |  |  | 1 | USED | IF PRT IS NO GO |  |  |  |  | IF |  |  | BAC5602AF | Greater of ±2.0°F or ±0.4% of Reading | FALSE |  |  | 24 | 0 | "0" | pg6 |  |  |  |  | Recall (IF there is one) | 0 | N/A | #REF!
SN | SN: 072513A |  | -51.1 | 199.1 | 448.9 | [12 empty] | PRT2SN:1624273 | PRT2SN:1624273 |  |  |  |  | and any of TS's are go |  |  |  |  | ST6 is NOGO |  |  | BPS4139M | Greater of ±2.0°F or ±0.4% of Reading | FALSE |  |  | 25 | 0 | "0" | pg7 |  |  |  | Thermocouple Wire |  | N/A |  | Correction Factors Comment: | C.F.=Correction Factor.
|  | C.F. | 1.1 | 0.9 | 1.1 | [16 empty] | TRUE |  | then 5-15-10 |  |  |  |  | AND SP1 is GO, |  |  | BAERD GEN-007D | Greater of ±2.0°F or ±0.4% of Reading | FALSE |  |  | 26 | 0 | "0" | pg8 |  |  | s | Thermocouple Wire |  |  |  | Comments S.H. | Comments Actual | Results
[18 empty] | SP-1 | SP-1 |  |  | FALSE |  | IF |  |  |  |  | AND anything else is GO |  |  | 0 |  | FALSE |  |  | 27 | 0 | "0" | pg9 | [5 empty] | 0 |  | 0 | Found in tolerance. No action required. "C.F" = "Correction Factor." | ACCEPTED
|  |  | -51 | 199.1 | 449 | [12 empty] | ST-7 | ST-7 |  |  | G295632, |  | Everything but SP1 is no go, |  |  |  |  | Then 5-7-10 | [7 empty] | 28 | 0 | "0" | 10 | [7 empty] | Reject | Found out of tolerance. "C.F." = "Correction Factor." | REJECTED
|  |  | 1 | 0.9 | 1 | [18 empty] | then 5-7-10 |  |  |  |  | IF | [7 empty] | 29 | 0 | "0" | 11
[24 empty] | IF |  |  |  |  | ST6 is NOGO |  | Supplimental tolerances here: | [5 empty] | 30 | 0 | "0" | 12 |  |  |  |  | WIRE GAUGE |  | 20 | Break | Unable to Calibrate Due to a Break in Continuity. C.F.=Correction Factor. | REJECTED
[19 empty] | Enter here | PRTPRTPRT |  |  |  | Everything but ST6 is no go, |  |  |  |  | AND SP1 IS NO GO, | [7 empty] | 31 | 0 | "0" | 13 | [6 empty] | N/A
[19 empty] | PRT SN:496142 |  |  |  |  | then 9-26-08 |  |  |  |  | AND Any Ts's are GO |  |  | TRUE | Greater of ±2.0°F or ±0.4% of Reading |  |  |  | 32 | 0 | "0" | 14 | [5 empty] | 0 | N/A
[11 empty] | VOLT METER USED |  |  | ALTERNATE METER |  |  |  | TEST SENSOR: | PRT SN:496142 | [9 empty] | AND PRT is GO |  |  | FALSE | Greater of ±4°F or ±0.75% of Reading |  |  |  | 33 | 0 | "0" | 15
[10 empty] | 1 | Additel 286 SN:6851022030037 |  | Name/SN | 0 |  |  |  | DATE CAL: | 2023-12-15 | [9 empty] | Then 5-15-10 |  |  | FALSE | Greater of ±1°F or ±0.1% of Reading |  |  |  | 34 | 0 | "0" | 16 |  |  |  |  | INSPECTOR FIELD |  | KF
[9 empty] | °F |  | 2024-07-01 |  | Cal Date | 0 |  |  |  | DUE: | 2025-12-15 | [12 empty] | FALSE | <1000°F:±2.0°F,>/= 1000°F: ±0.4% Rdg. |  |  |  |  | CALIBRATION PROCEDURE |  |  | (other than WP4007, which is in the form)
[11 empty] | 2025-07-01 |  | Due Date | 0 |  |  |  | TRACEABILITY # | G295632 | [12 empty] | FALSE | Type T: ±1.8°F / Type J,K,N: ±4.0°F | [12 empty] | KF
[11 empty] | 56561-060578 |  | Traceability | 0 |  | 0 | 0 | [14 empty] | FALSE | ±1.0°F or 0.1% of reading |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
[18 empty] | STACKER | [16 empty] | Controlling Specifications: | AMS2750H, ASTM E220-19
[11 empty] | HP3458A |  |  |  |  | Fluke 1502A | Fluke 1502A | Fluke 1502A | Fluke 1502A | Std 2 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | [7 empty] | TOLERANCE SOURCE: | [6 empty] | LOCATION FIELD: | Production
[11 empty] | H.P. 34420A |  |  |  |  | PRT | PRT | PRT | PRT | Std 2: Std 3; 1: PRT; 2: PRT; 3: PRT |  |  | TRUE | [7 empty] | Specification |  |  |  | 0
[20 empty] | Std 2: Std 4 | IF BPS is TRUE, and all else is false, BPS, if BPS is true, and anything before it is true, ", BPS" | [13 empty] | Production |  |  | CONDITION FIELD
[20 empty] | Std 2: Std 5 | IF baerdgen is True, and all else is false, baerd gen, if Beardgen is true and anything else is true, ", baerdgen" | [6 empty] | K | " | [5 empty] | 0
[20 empty] | Std 6 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | [7 empty] | J | " | "J"
[11 empty] | Std | S/N | Date Cert | Date Due | [23 empty] | N | " | [9 empty] | New
|  | (INCLUDE "BLANKET" UNCERTAINTY STATEMENT) | [6 empty] | Meter 1 |  | Additel 286 | 6851022030037 | 2024-07-01 | 2025-07-01 | [5 empty] | Std 5 Std 6: Banks 1-5; Std 5 11: PRTPRTPRT | [7 empty] | T | "
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor." | [7 empty] | Std 2 | Fluke 1502A | Fluke 1502A | B55852 | 2025-01-20 | 2026-01-20 | [5 empty] | Std 5 Std 6: Banks 6-10 | [7 empty] | R | " |  | J
[9 empty] | Std 3 | PRT | PRT | 496142 | 2023-12-15 | 2025-12-15 |  | T/N |  |  |  | Std 5 Std 6: Banks11-15 | [7 empty] | S | "
Condition: | New | [7 empty] | Std 4 | [10 empty] | Std 5 Std 6: Banks 16-20 | [7 empty] | E | "
Results: | ACCEPTED | [7 empty] | Std 5 | [6 empty] | 56561-068751 |  |  |  | Std 5 Std 6: All 20 Banks; Std 5 11: PRTPRTPRT
| Calibration Standard | [7 empty] | Std 6 | [6 empty] | G295632 |  |  |  | Std 5 Std 6: Sensors
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142 | [13 empty] | Std 5 Std 6: PRT
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 2
Customer: | [15 empty] | 56561-068751 |  | G295632,
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF | [9 empty] | G295632
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A | [9 empty] | T/N
Temp: | 72 |  |  | Relative Humidity: |  | 25 | % | [8 empty] | 56561-061486
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19 | [9 empty] | WO-00790034
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 3
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 4
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 5
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 6
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 7
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 8
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 9
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures
SN |  |  | 0 | 0 | 0
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 10
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | 0 | 0 | 0
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 11
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | 0 | 0 | 0
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 12
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | 0 | 0 | 0
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 13
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | 0 | 0 | 0
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 14
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | 0 | 0 | 0
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 15
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | 0 | 0 | 0
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 16
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | 0 | 0 | 0
SN
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 17
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below |  | 199.1 | 448.95
Test Temperatures |  |  | -50 | 200 | 450 |  |  |  | 0.9 | 1.05
SN | 072513A | Beginning | -51.1 | 199.1 | 448.9 |  |  | Unaveraged CFs
|  | End | -51 | 199.1 | 449 |  |  | -51.05
|  | C.F. | 1.1 | 0.9 | 1.1 |  |  | 1.05
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 18
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN | [7 empty] | Unaveraged CFs
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 19
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN | 072513A |  | -51.1 | 199.1 | 448.9
|  | C.F. | 1.1 | 0.9 | 1.1
|  |  | -51 | 199.1 | 449
|  | C.F. | 1 | 0.9 | 1
SN: 072513A
C.F.
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 20
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading |  |  |  | w7 | ae7 | am7
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19 |  |  |  |  | au7
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450 |  |  | o7
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 21
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN | 072513A |  | -51.1 | 199.1 | 448.9
|  | C.F. | 1.1 | 0.9 | 1.1
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 22
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19 |  |  |  |  | Bank 6
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below |  |  | Temp | AU7
Test Temperatures |  |  | -50 | 200 | 450 |  |  |  |  | C.F | AV7
SN |  |  | -51 | 199.1 | 449 | [5 empty] | 7
|  | C.F. | 1 | 0.9 | 1 |  |  |  |  | temp | BC
[10 empty] | cf | BD
8
[10 empty] | temp | BK
[10 empty] | cf | BL
9
[10 empty] | temp | BS
[10 empty] | cf | BT
10
[10 empty] | temp | CA
[10 empty] | cf | CB
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 23
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 24
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 25
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 26
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 27
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 28
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 29
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 30
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 31
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 32
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 33
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 34
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 35
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 36
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 37
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 38
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 39
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 40
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "0" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 41
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  | -50 | 200 | 450
SN | 072513A | -51.1 | 199.1 | 448.9
|  | -51 | 199.1 | 449
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 42
Customer:
Item | Primary Standard Thermocouple |  |  | INSPECTOR: |  | KF
Type | "S" |  |  | Date Calibrated: |  | 2022-12-22
Gauge: | 20 |  |  | Date Due: |  | 2023-12-22
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | 32 | 200 | 400 | 600 | 800
SN | SP-1 |  | 32.7333 | 201.0775 | 400.2841 | 599.934 | 800.0139
|  | C.F. | -0.73333 | -1.0775 | -0.28409 | 0.065963 | -0.013947
|  |  | 1000 | 1200 | 1400 | 1600 | 1800
|  |  | 1000.1554 | 1200.839 | 1400.17 | 1599.7961 | 1799.6955
|  | C.F. | -0.15542 | -0.83904 | -0.16996 | 0.20393 | 0.30448
|  |  | 2000 | 2200 | 2400
|  |  | 2001.7175 | 2201.7619 | #VALUE!
|  | C.F. | -1.7175 | -1.7619 | #VALUE!
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | ST-7
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-05-16
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2028-05-16
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | 46756
Report: |  |  | PAGE | 43
Customer:
Item | Primary Standard Thermocouple |  |  | INSPECTOR: |  | KF
Type | "S" |  |  | Date Calibrated: |  | 2022-12-22
Gauge: | 20 |  |  | Date Due: |  | 2023-12-22
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | 32 | 200 | 400 | 600 | 800
SN | SP-2 |  | 32.7333 | 199.6775 | 399.6341 | 599.734 | 799.8139
|  | C.F. | -0.73333 | 0.32254 | 0.36591 | 0.26596 | 0.18605
|  |  | 1000 | 1200
| SP-2 |  | 999.6454 | 1199.689
|  | C.F. | 0.35458 | 0.31096
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | ST-7
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-05-16
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2028-05-16
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | 46756
Report: |  |  | PAGE | 44
Customer:
Item | Secondary Standard Thermocouple |  |  | INSPECTOR: |  | KF
Type | "T" Special Limits |  |  | Date Calibrated: |  | 2022-12-22
Gauge: | 20 |  |  | Date Due: |  | 2023-12-22
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 32 | 120
SN | TTS-1 |  | -18 | 64 | 152
|  | C.F. | -32 | -32 | -32
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | #REF!
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Report: |  |  | PAGE | 45
Customer:
Item | Thermocouple Wire |  |  | INSPECTOR: |  | KF
Type | "J" |  |  | Date Calibrated: |  | 2025-04-09
Gauge: | N/A |  |  | Date Due: |  | N/A
Temp: | 72 |  |  | Relative Humidity: |  | 25 | %
Location: | Production |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | AMS2750H, ASTM E220-19
Calibration/Inspection Procedure: |  |  |  | Tolerance Source: |  | Specification | See Below
Test Temperatures |  |  | -50 | 200 | 450
SN
C.F.
(INCLUDE "BLANKET" UNCERTAINTY STATEMENT)
Comments: | Found in tolerance. No action required. "C.F" = "Correction Factor."
Condition: | New
Results: | ACCEPTED
Calibration Standard
TEST INSTRUMENT: |  | Additel 286 SN:6851022030037 |  |  | TEST T/C: | PRT SN:496142
DATE CERTIFIED: |  | 2024-07-01 |  |  | CERTIFIED | 2023-12-15
DATE DUE: |  | 2025-07-01 |  |  | DUE | 2025-12-15
TRACEABILITY: |  | 56561-060578 |  |  | TRACEABILITY: |  | G295632
Note on BI5: IF unit is "Accepted as noted", Include "Note" in the additional comments
RAW mVDC input Cells:
BANK 1: E7-26
Bank 2: M7-26
Bank3: U7-26
Bank 4: AC7-26
Bank 5: AK7-26
adjusted temp result cells:
BANK1: G7-26
Bank 2: O7-26
Bank 3: W7-26
Bank 4: AE7-26
Bank 5: AM7-26
Correction Factor Cells
Bank 1: H7-26
Bank 2: P7-26
Bank 3: X7-26
Bank 4: AF7-26
Bank 5: AN7-26
TEST Temperature Cells:
Banks1-5
E1,M1,U1,AC1,AK1
Banks6-10
AS1,BA1,BI1,BQ1,BY1

## Dropdowns
Standards
H.P. 3458A SN:2823A08926
FLUKE 1502 SN:B55852
Additel 286 SN:6851022030037

## CERT, Indiv TCs
| Johnson Gage and Inspection, Inc. | [7 empty] | password: jgi | [6 empty] | Tables updated: | 2023-04-12 | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base
| 5920 W. 21st St. N., Wichita, KS 67205 | [16 empty] | TC E with PRT | -112 to -22 | °F | °F | 0.212 |  | TC J with PRT | -112 to -22 | °F | °F | 0.222 |  | TC K with PRT | -112 to -22 | °F | °F | 0.276 |  | TC N with PRT | -112 to -22 | °F | °F | 0.405 |  | TC R with PRT | 32 to 392 | °F | °F | 0.871 | TC R with S ref | 32 to 392 | °F | °F | 1.04 |  | TC S with PRT | 32 to 392 | °F | °F | 0.871 |  | TC S with S ref | 32 to 392 | °F | °F | 1.04 |  | TC T with PRT | -112 to -22 | °F | °F | 0.286
| Phone 316.943.7532; Fax 316.944.6256 | [16 empty] | TC E with PRT | -22 to 392 | °F | °F | 0.19 |  | TC J with PRT | -22 to 392 | °F | °F | 0.21 |  | TC K with PRT | -22 to 392 | °F | °F | 0.244 |  | TC N with PRT | -22 to 392 | °F | °F | 0.333 |  | TC R with PRT | 392 to 752 | °F | °F | 0.81 | TC R with S ref | 392 to 1200 | °F | °F | 1.04 |  | TC S with PRT | 392 to 752 | °F | °F | 0.81 |  | TC S with S ref | 392 to 1200 | °F | °F | 1.04 |  | TC T with PRT | -22 to 392 | °F | °F | 0.227
[18 empty] | TC E with PRT | 392 to 752 | °F | °F | 0.183 |  | TC J with PRT | 392 to 752 | °F | °F | 0.222 |  | TC K with PRT | 392 to 752 | °F | °F | 0.272 |  | TC N with PRT | 392 to 752 | °F | °F | 0.312 |  | TC R with S std | 32 to 392 | °F | °F | 1.34 | TC R with S ref | 1200 to 2000 | °F | °F | 1.2 |  | TC S with S std | 32 to 392 | °F | °F | 1.33 |  | TC S with S ref | 1200 to 2000 | °F | °F | 1.2 |  | TC T with PRT | 392 to 752 | °F | °F | 0.213
| DATASHEET | [7 empty] | Page Number | [8 empty] | TC E with S | 32 to 392 | °F | °F | 1.33 |  | TC J with S | 32 to 392 | °F | °F | 1.07 |  | TC K with S | 32 to 392 | °F | °F | 1.33 |  | TC N with S std | 32 to 392 | °F | °F | 1.34 |  | TC R with S std | 392 to 1200 | °F | °F | 1.6 | TC R with S ref | 2000 to 2500 | °F | °F | 3.2 |  | TC S with S std | 392 to 752 | °F | °F | 1.33 | TC S with S std 2000 to 2500 | TC S with S ref | 2000 to 2500 | °F | °F | 3.2 |  | TC T with S std | 32 to 392 | °F | °F | 1.33
Report Number: | 072513A |  |  |  | Page 1 of 1 |  |  | Form 4007 | 1 | 1 | [7 empty] | TC E with S | 392 to 1200 | °F | °F | 2 |  | TC J with S | 392 to 1200 | °F | °F | 1.5 |  | TC K with S std | 392 to 1200 | °F | °F | 1.5 |  | TC N with S std | 392 to 1200 | °F | °F | 1.6 |  | TC R with S std | 1200 to 2000 | °F | °F | 2 | [6 empty] | TC S with S std | 752 to 1200 | °F | °F | 1.6 | 3.2 | [6 empty] | TC T with S std | 392 to 752 | °F | °F | 1.54
Customer: | JOHNSON GAGE & INSPECTION, INC. | [6 empty] | Rev. F 10-20-20 | [9 empty] | TC E with S | 1200 to 1832 | °F | °F | 2.2 |  | TC J with S | 1200 to 2000 | °F | °F | 2.01 |  | TC K with S std | 1200 to 2000 | °F | °F | 2.01 | S 2200 to 2500 | TC N with S std | 1200 to 2000 | °F | °F | 2.1 |  | TC R with S std | 2000 to 2500 | °F | °F | 3.24 | [6 empty] | TC S with S std | 1200 to 2000 | °F | °F | 2 | [6 empty] | PRT | 0.286 | 0.227 | 0.213 | 0 | 0
| 5920 W. 21st ST. N. | [15 empty] | PRT | 0.212 | 0.19 | 0.183 | 0 | 0 |  | TC J with S | 2000 to 2192 | °F | °F | 2.7 |  | TC K with S std | 2000 to 2200 | °F | °F | 2.4 | 3.2 | TC N with S std | 2000 to 2372 | °F | °F | 2.9 | PRT | 0 | 0.871 | 0.81 | 0 | 0 | [5 empty] | PRT | 0 | 0.871 | 0.81 | 0 | 0 | [6 empty] | S | 0 | 1.33 | 1.54 | 0 | 0
| WICHITA, KS 67205 | [15 empty] | S | 0 | 1.33 | 2 | 0 | 0 | PRT | 0.222 | 0.21 | 0.222 | 0 | 0 | PRT | 0.276 | 0.244 | 0.272 | 0 | 0 | PRT | 0.405 | 0.333 | 0.312 | 0 | 0 | S std | 0 | 1.34 | 1.6 | 0 | 0 | [5 empty] | S std | 0 | 1.33 | 1.33 | 0 | 0 | [6 empty] | PRT rng 1 | 0.286 | 0 | 0 | 0 | 0
[17 empty] | PRT rng 1 | 0.212 | 0 | 0 | 0 | 0 | S | 0 | 1.07 | 1.5 | 0 | 0 | S | 0 | 1.33 | 1.5 | 0 | 0 | S | 0 | 1.34 | 1.6 | 0 | 0 | S ref | 0 | 1.04 | 1.04 | 0 | 0 | [5 empty] | S ref | 0 | 1.04 | 1.04 | 0 | 0 | [6 empty] | PRT rng 2 | 0 | 0.227 | 0 | 0 | 0
Item: | Thermocouple Wire |  |  |  |  | Inspector: | KF | [9 empty] | PRT rng 2 | 0 | 0.19 | 0 | 0 | 0 | PRT rng 1 | 0.222 | 0 | 0 | 0 | 0 | PRT rng 1 | 0.276 | 0 | 0 | 0 | 0 | PRT rng 1 | 0.405 | 0 | 0 | 0 | 0 | PRT rng 1 | 0 | 0.871 | 0 | 0 | 0 | [5 empty] | PRT rng 1 | 0 | 0.871 | 0 | 0 | 0 | [6 empty] | PRT rng 3 | 0 | 0 | 0.213 | 0 | 0
Type: | "J" |  |  |  |  | Date Calibrated: | 2025-04-09 | [9 empty] | PRT rng 3 | 0 | 0.183 | 0.183 | 0 | 0 | PRT rng 2 | 0 | 0.21 | 0 | 0 | 0 | PRT rng 2 | 0 | 0.244 | 0 | 0 | 0 | PRT rng 2 | 0 | 0.333 | 0 | 0 | 0 | PRT rng 2 | 0 | 0 | 0.81 | 0 | 0 | [5 empty] | PRT rng 2 | 0 | 0 | 0.81 | 0 | 0 | [6 empty] | S rng 1 | 0 | 1.33 | 0 | 0 | 0
Gauge: | N/A |  |  |  |  | Recalibration Date: | N/A | [9 empty] | S rng 1 | 0 | 1.33 | 0 | 0 | 0 | PRT rng 3 | 0 | 0 | 0.222 | 0 | 0 | PRT rng 3 | 0 | 0 | 0.272 | 0 | 0 | PRT rng 3 | 0 | 0 | 0.312 | 0 | 0 | S rng 1 | 0 | 1.34 | 0 | 0 | 0 | [5 empty] | S rng 1 | 0 | 1.33 | 0 | 0 | 0 | [6 empty] | S rng 2 | 0 | 0 | 1.54 | 0 | 0
Temperature: | 72 |  |  |  |  | Relative Humidity: | 25% | [9 empty] | S rng 2 | 0 | 0 | 2 | 0 | 0 | S rng 1 | 0 | 1.07 | 0 | 0 | 0 | S rng 1 | 0 | 1.33 | 0 | 0 | 0 | S rng 1 | 0 | 1.34 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 1.6 | 0 | 0 | [5 empty] | S rng 2 | 0 | 0 | 1.33 | 0 | 0
[17 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 1.5 | 0 | 0 | S rng 2 | 0 | 0 | 1.5 | 0 | 0 | S rng 2 | 0 | 0 | 1.6 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0
| Calibration Procedure(s): | WP4007 |  |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading | [15 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 4 | 0 | 0 | 0 | 0 | 0
| Controlling Specification(s): | AMS2750H, ASTM E220-19 |  |  |  | Tolerance Source: | Specification |  |  |  | raw uncerts | 0.222 | 0.21 | 0.222 | 0 | 0 | [6 empty] | S rng 4 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | Sref Rng 1 | 0 | 1.04 | 0 | 0 | 0 | [5 empty] | S rng 5 | 0 | 0 | 0 | 0 | 0
[6 empty] | Condition of Item: | New |  |  |  |  | PRT | PRT | PRT | 0 | 0 | [12 empty] | S rng 5 | 0 | 0 | 0 | 0 | 0 | [6 empty] | Sref Rng 2 | 0 | 0 | 1.04 | 0 | 0 | [5 empty] | Sref Rng 1 | 0 | 1.04 | 0 | 0 | 0
[41 empty] | Sref Rng 3 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 2 | 0 | 0 | 1.04 | 0 | 0
Uncertainty @k=2 |  | 0°F |  | 0.21°F |  | 0.22°F |  |  |  |  | Uncerts | 0 | 0.21 | 0.22 | [26 empty] | Sref Rng 4 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref rng 3 | 0 | 0 | 0 | 0 | 0
Test Temperatures |  | -50 |  | 200 |  | 450 |  |  |  |  | Test Temps | -50 | 200 | 450 | [37 empty] | Sref rng 4 | 0 | 0 | 0 | 0 | 0
SN: 072513A |  | -51.1 |  | 199.1 |  | 448.9 |  |  |  |  | E | 0.212 | 0.19 | 0.183 | 0 | 0
| C.F. | 1.1 |  | 0.9 |  | 1.1 |  |  |  |  | J | 0.222 | 0.21 | 0.222 | 0 | 0
[11 empty] | K | 0.276 | 0.244 | 0.272 | 0 | 0 | [48 empty] | Temp 1 | Temp 2 | T3 | T4 | Temp 5
|  | -51 |  | 199.1 |  | 449 |  |  |  |  | N | 0.405 | 0.333 | 0.312 | 0 | 0
|  | 1 |  | 0.9 |  | 1 |  |  |  |  | R | 0 | 0.871 | 0.81 | 0 | 0
[11 empty] | S | 0 | 0.871 | 0.81 | 0 | 0
[11 empty] | T | 0.286 | 0.227 | 0.213 | 0 | 0
The JGI quality system governing this calibration/inspection is based upon and maintained in compliance with quality specifications ISO/IEC 17025, ANSI/NCSL Z540-1 and ISO 10012.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection.
Comments:
Found in tolerance. No action required. "C.F" = "Correction Factor."
RESULTS: | ACCEPTED
TEST EQUIPMENT
Test Equipment |  | S/N |  | Date Certified |  | Date Due | Traceability Number
Additel 286 |  | 6851022030037 |  | 2024-07-01 |  | 2025-07-01
Fluke 1502A |  | B55852 |  | 2025-01-20 |  | 2026-01-20 | 56561-068751
PRT |  | 496142 |  | 2023-12-15 |  | 2025-12-15 | G295632
Johnson Gage and Inspection, Inc.
5920 W. 21st St. N., Wichita, KS 67205
Phone 316.943.7532; Fax 316.944.6256
| DATASHEET | [7 empty] | Page Number
Report Number: | 072513A |  |  |  | Page 2 of 1 |  |  | Form 4007 | 1 | 1
Customer: | JOHNSON GAGE & INSPECTION, INC. | [6 empty] | Rev. F 10-20-20
5920 W. 21st ST. N.
WICHITA, KS 67205
Item: | Thermocouple Wire |  |  |  |  | Inspector: | KF
Type: | "J" |  |  |  |  | Date Calibrated: | 2025-04-09
Gauge: | N/A |  |  |  |  | Recalibration Date: | N/A
Temperature: | 72 |  |  |  |  | Relative Humidity: | 25%
| Calibration Procedure(s): | WP4007 |  |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
| Controlling Specification(s): | AMS2750H, ASTM E220-19 |  |  |  | Tolerance Source: | Specification |  |  |  | raw uncerts | 0.222 | 0.21 | 0.222 | 0 | 0
[6 empty] | Condition of Item: | New |  |  |  |  | PRT | PRT | PRT | 0 | 0
Uncertainty @k=2 |  | 0.22°F |  | 0.21°F |  | 0.22°F |  |  |  |  | Uncerts | 0.22 | 0.21 | 0.22
Test Temperatures |  | -50 |  | 200 |  | 450 |  |  |  |  | Test Temps | -50 | 200 | 450
[11 empty] | E | 0.212 | 0.19 | 0.183 | 0 | 0
[11 empty] | J | 0.222 | 0.21 | 0.222 | 0 | 0
[11 empty] | K | 0.276 | 0.244 | 0.272 | 0 | 0
[11 empty] | N | 0.405 | 0.333 | 0.312 | 0 | 0
[11 empty] | R | 0 | 0.871 | 0 | 0 | 0
[11 empty] | S | 0 | 0.871 | 0.81 | 0 | 0
[11 empty] | T | 0.286 | 0.227 | 0.213 | 0 | 0
The JGI quality system governing this calibration/inspection is based upon and maintained in compliance with quality specifications ISO/IEC 17025, ANSI/NCSL Z540-1 and ISO 10012.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection.
Comments:
Found in tolerance. No action required. "C.F" = "Correction Factor."
RESULTS: | ACCEPTED
TEST EQUIPMENT
Test Equipment |  | S/N |  | Date Certified |  | Date Due | Traceability Number
Additel 286 |  | 6851022030037 |  | 2024-07-01 |  | 2025-07-01
Fluke 1502A |  | B55852 |  | 2025-01-20 |  | 2026-01-20 | 56561-068751
PRT |  | 496142 |  | 2023-12-15 |  | 2025-12-15 | G295632
Johnson Gage and Inspection, Inc.
5920 W. 21st St. N., Wichita, KS 67205
Phone 316.943.7532; Fax 316.944.6256
| DATASHEET | [7 empty] | Page Number
Report Number: | 072513A |  |  |  | Page 3 of 1 |  |  | Form 4007 | 1 | 1
Customer: | JOHNSON GAGE & INSPECTION, INC. | [6 empty] | Rev. F 10-20-20
5920 W. 21st ST. N.
WICHITA, KS 67205
Item: | Thermocouple Wire |  |  |  |  | Inspector: | KF
Type: | "J" |  |  |  |  | Date Calibrated: | 2025-04-09
Gauge: | N/A |  |  |  |  | Recalibration Date: | N/A
Temperature: | 72 |  |  |  |  | Relative Humidity: | 25%
| Calibration Procedure(s): | WP4007 |  |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
| Controlling Specification(s): | AMS2750H, ASTM E220-19 |  |  |  | Tolerance Source: | Specification |  |  |  | raw uncerts | 0.222 | 0.21 | 0.222 | 0 | 0
[6 empty] | Condition of Item: | New |  |  |  |  | PRT | PRT | PRT | 0 | 0
Uncertainty @k=2 |  | 0.22°F |  | 0.21°F |  | 0.22°F |  |  |  |  | Uncerts | 0.22 | 0.21 | 0.22
Test Temperatures |  | -50 |  | 200 |  | 450 |  |  |  |  | Test Temps | -50 | 200 | 450
[11 empty] | E | 0.212 | 0.19 | 0.183 | 0 | 0
[11 empty] | J | 0.222 | 0.21 | 0.222 | 0 | 0
[11 empty] | K | 0.276 | 0.244 | 0.272 | 0 | 0
[11 empty] | N | 0.405 | 0.333 | 0.312 | 0 | 0
[11 empty] | R | 0 | 0.871 | 0 | 0 | 0
[11 empty] | S | 0 | 0.871 | 0.81 | 0 | 0
[11 empty] | T | 0.286 | 0.227 | 0.213 | 0 | 0
The JGI quality system governing this calibration/inspection is based upon and maintained in compliance with quality specifications ISO/IEC 17025, ANSI/NCSL Z540-1 and ISO 10012.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection.
Comments:
Found in tolerance. No action required. "C.F" = "Correction Factor."
RESULTS: | ACCEPTED
TEST EQUIPMENT
Test Equipment |  | S/N |  | Date Certified |  | Date Due | Traceability Number
Additel 286 |  | 6851022030037 |  | 2024-07-01 |  | 2025-07-01
Fluke 1502A |  | B55852 |  | 2025-01-20 |  | 2026-01-20 | 56561-068751
PRT |  | 496142 |  | 2023-12-15 |  | 2025-12-15 | G295632
Johnson Gage and Inspection, Inc.
5920 W. 21st St. N., Wichita, KS 67205
Phone 316.943.7532; Fax 316.944.6256
| DATASHEET | [7 empty] | Page Number
Report Number: | 072513A |  |  |  | Page 4 of 1 |  |  | Form 4007 | 1 | 1
Customer: | JOHNSON GAGE & INSPECTION, INC. | [6 empty] | Rev. F 10-20-20
5920 W. 21st ST. N.
WICHITA, KS 67205
Item: | Thermocouple Wire |  |  |  |  | Inspector: | KF
Type: | "J" |  |  |  |  | Date Calibrated: | 2025-04-09
Gauge: | N/A |  |  |  |  | Recalibration Date: | N/A
Temperature: | 72 |  |  |  |  | Relative Humidity: | 25%
| Calibration Procedure(s): | WP4007 |  |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
| Controlling Specification(s): | AMS2750H, ASTM E220-19 |  |  |  | Tolerance Source: | Specification |  |  |  | raw uncerts | 0.222 | 0.21 | 0.222 | 0 | 0
[6 empty] | Condition of Item: | New |  |  |  |  | PRT | PRT | PRT | 0 | 0
Uncertainty @k=2 |  | 0.22°F |  | 0.21°F |  | 0.22°F |  |  |  |  | Uncerts | 0.22 | 0.21 | 0.22
Test Temperatures |  | -50 |  | 200 |  | 450 |  |  |  |  | Test Temps | -50 | 200 | 450
[11 empty] | E | 0.212 | 0.19 | 0.183 | 0 | 0
[11 empty] | J | 0.222 | 0.21 | 0.222 | 0 | 0
[11 empty] | K | 0.276 | 0.244 | 0.272 | 0 | 0
[11 empty] | N | 0.405 | 0.333 | 0.312 | 0 | 0
[11 empty] | R | 0 | 0.871 | 0 | 0 | 0
[11 empty] | S | 0 | 0.871 | 0.81 | 0 | 0
[11 empty] | T | 0.286 | 0.227 | 0.213 | 0 | 0
The JGI quality system governing this calibration/inspection is based upon and maintained in compliance with quality specifications ISO/IEC 17025, ANSI/NCSL Z540-1 and ISO 10012.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection.
Comments:
Found in tolerance. No action required. "C.F" = "Correction Factor."
RESULTS: | ACCEPTED
TEST EQUIPMENT
Test Equipment |  | S/N |  | Date Certified |  | Date Due | Traceability Number
Additel 286 |  | 6851022030037 |  | 2024-07-01 |  | 2025-07-01
Fluke 1502A |  | B55852 |  | 2025-01-20 |  | 2026-01-20 | 56561-068751
PRT |  | 496142 |  | 2023-12-15 |  | 2025-12-15 | G295632
| Johnson Gage and Inspection, Inc. | [14 empty] | Tables updated: | 2023-04-12 | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base
| 5920 W. 21st St. N., Wichita, KS 67205 | [16 empty] | TC E with PRT | -112 to -22 | °F | °F | 0.212 |  | TC J with PRT | -112 to -22 | °F | °F | 0.222 |  | TC K with PRT | -112 to -22 | °F | °F | 0.276 |  | TC N with PRT | -112 to -22 | °F | °F | 0.405 |  | TC R with PRT | 32 to 392 | °F | °F | 0.871 | TC R with S ref | 32 to 392 | °F | °F | 1.04 |  | TC S with PRT | 32 to 392 | °F | °F | 0.871 |  | TC S with S ref | 32 to 392 | °F | °F | 1.04 |  | TC T with PRT | -112 to -22 | °F | °F | 0.286
| Phone 316.943.7532; Fax 316.944.6256 | [16 empty] | TC E with PRT | -22 to 392 | °F | °F | 0.19 |  | TC J with PRT | -22 to 392 | °F | °F | 0.21 |  | TC K with PRT | -22 to 392 | °F | °F | 0.244 |  | TC N with PRT | -22 to 392 | °F | °F | 0.333 |  | TC R with PRT | 392 to 752 | °F | °F | 0.81 | TC R with S ref | 392 to 1200 | °F | °F | 1.04 |  | TC S with PRT | 392 to 752 | °F | °F | 0.81 |  | TC S with S ref | 392 to 1200 | °F | °F | 1.04 |  | TC T with PRT | -22 to 392 | °F | °F | 0.227
[18 empty] | TC E with PRT | 392 to 752 | °F | °F | 0.183 |  | TC J with PRT | 392 to 752 | °F | °F | 0.222 |  | TC K with PRT | 392 to 752 | °F | °F | 0.272 |  | TC N with PRT | 392 to 752 | °F | °F | 0.312 |  | TC R with S std | 32 to 392 | °F | °F | 1.34 | TC R with S ref | 1200 to 2000 | °F | °F | 1.2 |  | TC S with S std | 32 to 392 | °F | °F | 1.33 |  | TC S with S ref | 1200 to 2000 | °F | °F | 1.2 |  | TC T with PRT | 392 to 752 | °F | °F | 0.213
| DATASHEET | [7 empty] | Page Number | [8 empty] | TC E with S | 32 to 392 | °F | °F | 1.33 |  | TC J with S | 32 to 392 | °F | °F | 1.07 |  | TC K with S | 32 to 392 | °F | °F | 1.33 |  | TC N with S std | 32 to 392 | °F | °F | 1.34 |  | TC R with S std | 392 to 1200 | °F | °F | 1.6 | TC R with S ref | 2000 to 2500 | °F | °F | 3.2 |  | TC S with S std | 392 to 1200 | °F | °F | 1.33 |  | TC S with S ref | 2000 to 2500 | °F | °F | 3.2 |  | TC T with S std | 32 to 392 | °F | °F | 1.33
Report Number: | 072513A |  |  |  | Page 5 of 1 |  |  | Form 4007 | 1 | 1 | [7 empty] | TC E with S | 392 to 1200 | °F | °F | 2 |  | TC J with S | 392 to 1200 | °F | °F | 1.5 |  | TC K with S std | 392 to 1200 | °F | °F | 1.5 |  | TC N with S std | 392 to 1200 | °F | °F | 1.6 |  | TC R with S std | 1200 to 2000 | °F | °F | 2 | [6 empty] | TC S with S std | 1200 to 2000 | °F | °F | 1.6 | [7 empty] | TC T with S std | 392 to 752 | °F | °F | 1.54
Customer: | JOHNSON GAGE & INSPECTION, INC. | [6 empty] | Rev. F 10-20-20 | [9 empty] | TC E with S | 1200 to 1832 | °F | °F | 2.2 |  | TC J with S | 1200 to 2000 | °F | °F | 2.01 |  | TC K with S std | 1200 to 2000 | °F | °F | 2.01 |  | TC N with S std | 1200 to 2000 | °F | °F | 2.1 |  | TC R with S std | 2000 to 2500 | °F | °F | 3.24 | [6 empty] | TC S with S std | 2000 to 2500 | °F | °F | 2 | [6 empty] | PRT | 0 | 0 | 0 | 0 | 0
| 5920 W. 21st ST. N. | [15 empty] | PRT | 0 | 0 | 0 | 0 | 0 |  | TC J with S | 2000 to 2192 | °F | °F | 2.7 |  | TC K with S std | 2000 to 2500 | °F | °F | 2.4 |  | TC N with S std | 2000 to 2372 | °F | °F | 2.9 | PRT | 0 | 0 | 0 | 0 | 0 | [5 empty] | PRT | 0.871 | 0.871 | 0.871 | 0.871 | 0.871 | [6 empty] | S | 0 | 0 | 0 | 0 | 0
| WICHITA, KS 67205 | [15 empty] | S | 0 | 1.33 | 0 | 0 | 0 | PRT | 0 | 0 | 0 | 0 | 0 | PRT | 0 | 0 | 0 | 0 | 0 | PRT | 0 | 0 | 0 | 0 | 0 | S std | 0 | 0 | 0 | 0 | 0 | [5 empty] | S std | 0 | 0 | 0 | 0 | 0 | [6 empty] | PRT rng 1 | 0 | 0 | 0 | 0 | 0
[17 empty] | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | S | 0 | 0 | 0 | 0 | 0 | S | 0 | 0 | 0 | 0 | 0 | S | 0 | 0 | 0 | 0 | 0 | S ref | 0 | 0 | 0 | 0 | 0 | [5 empty] | S ref | 0 | 0 | 0 | 0 | 0 | [6 empty] | PRT rng 2 | 0 | 0 | 0 | 0 | 0
Item: | Thermocouple Wire |  |  |  |  | Inspector: | KF | [9 empty] | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | [5 empty] | PRT rng 1 | 0.871 | 0.871 | 0.871 | 0.871 | 0.871 | [6 empty] | PRT rng 3 | 0 | 0 | 0 | 0 | 0
Type: | "J" |  |  |  |  | Date Calibrated: | 2025-04-09 | [9 empty] | PRT rng 3 | 0 | 0 | 0 | 0 | 0 | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | [5 empty] | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | [6 empty] | S rng 1 | 0 | 0 | 0 | 0 | 0
Gauge: | N/A |  |  |  |  | Recalibration Date: | N/A | [9 empty] | S rng 1 | 0 | 1.33 | 0 | 0 | 0 | PRT rng 3 | 0 | 0 | 0 | 0 | 0 | PRT rng 3 | 0 | 0 | 0 | 0 | 0 | PRT rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 1 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 1 | 0 | 0 | 0 | 0 | 0 | [6 empty] | S rng 2 | 0 | 0 | 0 | 0 | 0
Temperature: | 72 |  |  |  |  | Relative Humidity: | 25% | [9 empty] | S rng 2 | 0 | 0 | 0 | 0 | 0 | S rng 1 | 0 | 0 | 0 | 0 | 0 | S rng 1 | 0 | 0 | 0 | 0 | 0 | S rng 1 | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 2 | 0 | 0 | 0 | 0 | 0
[17 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0
| Calibration Procedure(s): | WP4007 |  |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading | [15 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 4 | 0 | 0 | 0 | 0 | 0
| Controlling Specification(s): | AMS2750H, ASTM E220-19 |  |  |  | Tolerance Source: | Specification |  |  |  | raw uncerts | 0 | 0 | 0 | 0 | 0 | [6 empty] | S rng 4 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | Sref Rng 1 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | Condition of Item: | New |  |  |  |  | 0 | 0 | 0 | 0 | 0 | [24 empty] | Sref Rng 2 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 2 | 0 | 0 | 0 | 0 | 0
[41 empty] | Sref Rng 3 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 3 | 0 | 0 | 0 | 0 | 0
Uncertainty @k=2 | [10 empty] | Uncerts | [29 empty] | Sref Rng 4 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 4 | 0 | 0 | 0 | 0 | 0
Test Temperatures | [10 empty] | Test Temps
[11 empty] | E | 0 | 1.33 | 0 | 0 | 0
[11 empty] | J | 0 | 0 | 0 | 0 | 0
[11 empty] | K | 0 | 0 | 0 | 0 | 0
[11 empty] | N | 0 | 0 | 0 | 0 | 0
[11 empty] | R | 0 | 0 | 0 | 0 | 0
[11 empty] | S | 0 | 0 | 0 | 0 | 0
[11 empty] | T | 0 | 0 | 0 | 0 | 0
The JGI quality system governing this calibration/inspection is based upon and maintained in compliance with quality specifications ISO/IEC 17025, ANSI/NCSL Z540-1 and ISO 10012.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection.
Comments:
Found in tolerance. No action required. "C.F" = "Correction Factor."
RESULTS: | ACCEPTED
TEST EQUIPMENT
Test Equipment |  | S/N |  | Date Certified |  | Date Due | Traceability Number
Additel 286 |  | 6851022030037 |  | 2024-07-01 |  | 2025-07-01
Fluke 1502A |  | B55852 |  | 2025-01-20 |  | 2026-01-20 | 56561-068751
PRT |  | 496142 |  | 2023-12-15 |  | 2025-12-15 | G295632
Johnson Gage and Inspection, Inc.
5920 W. 21st St. N., Wichita, KS 67205
Phone 316.943.7532; Fax 316.944.6256
| DATASHEET | [7 empty] | Page Number
Report Number: | 072513A |  |  |  | Page 6 of 1 |  |  | Form 4007 | 1 | 1
Customer: | JOHNSON GAGE & INSPECTION, INC. | [6 empty] | Rev. F 10-20-20
5920 W. 21st ST. N.
WICHITA, KS 67205
Item: | Thermocouple Wire |  |  |  |  | Inspector: | KF
Type: | "J" |  |  |  |  | Date Calibrated: | 2025-04-09
Gauge: | N/A |  |  |  |  | Recalibration Date: | N/A
Temperature: | 72 |  |  |  |  | Relative Humidity: | 25%
| Calibration Procedure(s): | WP4007 |  |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
| Controlling Specification(s): | AMS2750H, ASTM E220-19 |  |  |  | Tolerance Source: | Specification |  |  |  | raw uncerts | 0 | 0 | 0 | 0 | 0
[6 empty] | Condition of Item: | New |  |  |  |  | 0 | 0 | 0 | 0 | 0
Uncertainty @k=2 | [10 empty] | Uncerts
Test Temperatures | [10 empty] | Test Temps
[11 empty] | E | 0 | 1.33 | 0 | 0 | 0
[11 empty] | J | 0 | 0 | 0 | 0 | 0
[11 empty] | K | 0 | 0 | 0 | 0 | 0
[11 empty] | N | 0 | 0 | 0 | 0 | 0
[11 empty] | R | 0 | 0 | 0 | 0 | 0
[11 empty] | S | 0 | 0 | 0 | 0 | 0
[11 empty] | T | 0 | 0 | 0 | 0 | 0
The JGI quality system governing this calibration/inspection is based upon and maintained in compliance with quality specifications ISO/IEC 17025, ANSI/NCSL Z540-1 and ISO 10012.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection.
Comments:
Found in tolerance. No action required. "C.F" = "Correction Factor."
RESULTS: | ACCEPTED
TEST EQUIPMENT
Test Equipment |  | S/N |  | Date Certified |  | Date Due | Traceability Number
Additel 286 |  | 6851022030037 |  | 2024-07-01 |  | 2025-07-01
Fluke 1502A |  | B55852 |  | 2025-01-20 |  | 2026-01-20 | 56561-068751
PRT |  | 496142 |  | 2023-12-15 |  | 2025-12-15 | G295632
Johnson Gage and Inspection, Inc.
5920 W. 21st St. N., Wichita, KS 67205
Phone 316.943.7532; Fax 316.944.6256
| DATASHEET | [7 empty] | Page Number
Report Number: | 072513A |  |  |  | Page 7 of 1 |  |  | Form 4007 | 1 | 1
Customer: | JOHNSON GAGE & INSPECTION, INC. | [6 empty] | Rev. F 10-20-20
5920 W. 21st ST. N.
WICHITA, KS 67205
Item: | Thermocouple Wire |  |  |  |  | Inspector: | KF
Type: | "J" |  |  |  |  | Date Calibrated: | 2025-04-09
Gauge: | N/A |  |  |  |  | Recalibration Date: | N/A
Temperature: | 72 |  |  |  |  | Relative Humidity: | 25%
| Calibration Procedure(s): | WP4007 |  |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
| Controlling Specification(s): | AMS2750H, ASTM E220-19 |  |  |  | Tolerance Source: | Specification |  |  |  | raw uncerts | 0 | 0 | 0 | 0 | 0
[6 empty] | Condition of Item: | New |  |  |  |  | 0 | 0 | 0 | 0 | 0
Uncertainty @k=2 | [10 empty] | Uncerts
Test Temperatures | [10 empty] | Test Temps
[11 empty] | E | 0 | 1.33 | 0 | 0 | 0
[11 empty] | J | 0 | 0 | 0 | 0 | 0
[11 empty] | K | 0 | 0 | 0 | 0 | 0
[11 empty] | N | 0 | 0 | 0 | 0 | 0
[11 empty] | R | 0 | 0 | 0 | 0 | 0
[11 empty] | S | 0 | 0 | 0 | 0 | 0
[11 empty] | T | 0 | 0 | 0 | 0 | 0
The JGI quality system governing this calibration/inspection is based upon and maintained in compliance with quality specifications ISO/IEC 17025, ANSI/NCSL Z540-1 and ISO 10012.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection.
Comments:
Found in tolerance. No action required. "C.F" = "Correction Factor."
RESULTS: | ACCEPTED
TEST EQUIPMENT
Test Equipment |  | S/N |  | Date Certified |  | Date Due | Traceability Number
Additel 286 |  | 6851022030037 |  | 2024-07-01 |  | 2025-07-01
Fluke 1502A |  | B55852 |  | 2025-01-20 |  | 2026-01-20 | 56561-068751
PRT |  | 496142 |  | 2023-12-15 |  | 2025-12-15 | G295632
Johnson Gage and Inspection, Inc.
5920 W. 21st St. N., Wichita, KS 67205
Phone 316.943.7532; Fax 316.944.6256
| DATASHEET | [7 empty] | Page Number
Report Number: | 072513A |  |  |  | Page 8 of 1 |  |  | Form 4007 | 1 | 1
Customer: | JOHNSON GAGE & INSPECTION, INC. | [6 empty] | Rev. F 10-20-20
5920 W. 21st ST. N.
WICHITA, KS 67205
Item: | Thermocouple Wire |  |  |  |  | Inspector: | KF
Type: | "J" |  |  |  |  | Date Calibrated: | 2025-04-09
Gauge: | N/A |  |  |  |  | Recalibration Date: | N/A
Temperature: | 72 |  |  |  |  | Relative Humidity: | 25%
| Calibration Procedure(s): | WP4007 |  |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
| Controlling Specification(s): | AMS2750H, ASTM E220-19 |  |  |  | Tolerance Source: | Specification |  |  |  | raw uncerts | 0 | 0 | 0 | 0 | 0
[6 empty] | Condition of Item: | New |  |  |  |  | 0 | 0 | 0 | 0 | 0
Uncertainty @k=2 | [10 empty] | Uncerts
Test Temperatures | [10 empty] | Test Temps
[11 empty] | E | 0 | 1.33 | 0 | 0 | 0
[11 empty] | J | 0 | 0 | 0 | 0 | 0
[11 empty] | K | 0 | 0 | 0 | 0 | 0
[11 empty] | N | 0 | 0 | 0 | 0 | 0
[11 empty] | R | 0 | 0 | 0 | 0 | 0
[11 empty] | S | 0 | 0 | 0 | 0 | 0
[11 empty] | T | 0 | 0 | 0 | 0 | 0
The JGI quality system governing this calibration/inspection is based upon and maintained in compliance with quality specifications ISO/IEC 17025, ANSI/NCSL Z540-1 and ISO 10012.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection.
Comments:
Found in tolerance. No action required. "C.F" = "Correction Factor."
RESULTS: | ACCEPTED
TEST EQUIPMENT
Test Equipment |  | S/N |  | Date Certified |  | Date Due | Traceability Number
Additel 286 |  | 6851022030037 |  | 2024-07-01 |  | 2025-07-01
Fluke 1502A |  | B55852 |  | 2025-01-20 |  | 2026-01-20 | 56561-068751
PRT |  | 496142 |  | 2023-12-15 |  | 2025-12-15 | G295632

## CERT, Wire Roll
| Johnson Gage and Inspection, Inc. | [7 empty] | password: jgi | [5 empty] | Tables updated: | 2021-06-03 | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base
| 5920 W. 21st St. N., Wichita, KS 67205 | [15 empty] | TC E with PRT | -112 to -22 | °F | °F | 0.19 |  | TC J with PRT | -112 to -22 | °F | °F | 0.2 |  | TC K with PRT | -112 to -22 | °F | °F | 0.22 |  | TC N with PRT | -112 to -22 | °F | °F | 0.28 |  | TC R with PRT | 32 to 392 | °F | °F | 0.81 | TC R with S ref | 32 to 392 | °F | °F | 1.02 |  | TC S with PRT | 32 to 392 | °F | °F | 0.81 | TC S with S ref | 32 to 392 | °F | °F | 1.01 |  | TC T with PRT | -112 to -22 | °F | °F | 0.23
| Phone 316.943.7532; Fax 316.944.6256 | [15 empty] | TC E with PRT | -22 to 392 | °F | °F | 0.19 |  | TC J with PRT | -22 to 392 | °F | °F | 0.21 |  | TC K with PRT | -22 to 392 | °F | °F | 0.24 |  | TC N with PRT | -22 to 392 | °F | °F | 0.26 |  | TC R with PRT | 392 to 752 | °F | °F | 0.81 | TC R with S ref | 392 to 1200 | °F | °F | 1.02 |  | TC S with PRT | 392 to 752 | °F | °F | 0.81 | TC S with S ref | 392 to 1200 | °F | °F | 1.01 |  | TC T with PRT | -22 to 392 | °F | °F | 0.2
[17 empty] | TC E with PRT | 392 to 752 | °F | °F | 0.18 |  | TC J with PRT | 392 to 752 | °F | °F | 0.21 |  | TC K with PRT | 392 to 752 | °F | °F | 0.24 |  | TC N with PRT | 392 to 752 | °F | °F | 0.26 |  | TC R with S std | 32 to 392 | °F | °F | 1.3 | TC R with S ref | 1200 to 2000 | °F | °F | 1.2 |  | TC S with S std | 32 to 392 | °F | °F | 1.3 | TC S with S ref | 1200 to 2000 | °F | °F | 1.2 |  | TC T with PRT | 392 to 752 | °F | °F | 0.2
| DATASHEET | [15 empty] | TC E with S | 32 to 392 | °F | °F | 1.3068 |  | TC J with S | 32 to 392 | °F | °F | 1.0166 |  | TC K with S | 32 to 392 | °F | °F | 1.31 |  | TC N with S std | 32 to 392 | °F | °F | 1.31 |  | TC R with S std | 392 to 1200 | °F | °F | 1.6 | TC R with S ref | 2000 to 2500 | °F | °F | 3.2 |  | TC S with S std | 392 to 1200 | °F | °F | 1.3 | TC S with S ref | 2000 to 2500 | °F | °F | 3.2 |  | TC T with S std | 32 to 392 | °F | °F | 1.31
Report Number: | 072513A |  |  |  | Page 1 of 1 |  |  | Form 4007 | [8 empty] | TC E with S | 392 to 1200 | °F | °F | 2 |  | TC J with S | 392 to 1200 | °F | °F | 1.5 |  | TC K with S std | 392 to 1200 | °F | °F | 1.5 |  | TC N with S std | 392 to 1200 | °F | °F | 1.6 |  | TC R with S std | 1200 to 2000 | °F | °F | 2 | [6 empty] | TC S with S std | 1200 to 2000 | °F | °F | 1.6 | [6 empty] | TC T with S std | 392 to 752 | °F | °F | 1.44
Customer: | JOHNSON GAGE & INSPECTION, INC. | [6 empty] | Rev. F 10-20-20 | [8 empty] | TC E with S | 1200 to 1832 | °F | °F | 2.2 |  | TC J with S | 1200 to 2000 | °F | °F | 2.0018 |  | TC K with S std | 1200 to 2000 | °F | °F | 2 |  | TC N with S std | 1200 to 2000 | °F | °F | 2.1 |  | TC R with S std | 2000 to 2500 | °F | °F | 3.23 | [6 empty] | TC S with S std | 2000 to 2500 | °F | °F | 2 | [5 empty] | PRT | 0.23 | 0.2 | 0.2 | 0 | 0
| 5920 W. 21st ST. N. | [14 empty] | PRT | 0.19 | 0.19 | 0.18 | 0 | 0 |  | TC J with S | 2000 to 2192 | °F | °F | 2.7 |  | TC K with S std | 2000 to 2500 | °F | °F | 3.2 |  | TC N with S std | 2000 to 2372 | °F | °F | 2.9 | PRT | 0 | 0.81 | 0.81 | 0 | 0 | [5 empty] | PRT | 0 | 0.81 | 0.81 | 0.81 | 3.2 | [5 empty] | S | 0 | 1.31 | 1.44 | 0 | 0
| WICHITA, KS 67205 | [14 empty] | S | 0 | 1.3068 | 2 | 0 | 0 | PRT | 0.2 | 0.21 | 0.21 | 0 | 0 | PRT | 0.22 | 0.24 | 0.24 | 0 | 0 | PRT | 0.28 | 0.26 | 0.26 | 0 | 0 | S std | 0 | 1.3 | 1.6 | 0 | 0 | [5 empty] | S std | 0 | 1.3 | 1.3 | 0 | 0 | [5 empty] | PRT rng 1 | 0.23 | 0 | 0 | 0 | 0
[16 empty] | PRT rng 1 | 0.19 | 0 | 0 | 0 | 0 | S | 0 | 1.0166 | 1.5 | 0 | 0 | S | 0 | 1.31 | 1.5 | 0 | 0 | S | 0 | 1.31 | 1.6 | 0 | 0 | S ref | 0 | 1.02 | 1.02 | 0 | 0 | [5 empty] | S ref | 0 | 1.01 | 1.01 | 0 | 0 | [5 empty] | PRT rng 2 | 0 | 0.2 | 0 | 0 | 0
Item: | Thermocouple Wire |  |  |  |  | Inspector: | KF | [8 empty] | PRT rng 2 | 0 | 0.19 | 0 | 0 | 0 | PRT rng 1 | 0.2 | 0 | 0 | 0 | 0 | PRT rng 1 | 0.22 | 0 | 0 | 0 | 0 | PRT rng 1 | 0.28 | 0 | 0 | 0 | 0 | PRT rng 1 | 0 | 0.81 | 0 | 0 | 0 | [5 empty] | PRT rng 1 | 0 | 0.81 | 0.81 | 0.81 | 0.81 | [5 empty] | PRT rng 3 | 0 | 0 | 0.2 | 0 | 0
Type: | "J" |  |  |  |  | Date Calibrated: | 2025-04-09 | [8 empty] | PRT rng 3 | 0 | 0.18 | 0.18 | 0 | 0 | PRT rng 2 | 0 | 0.21 | 0 | 0 | 0 | PRT rng 2 | 0 | 0.24 | 0 | 0 | 0 | PRT rng 2 | 0 | 0.26 | 0 | 0 | 0 | PRT rng 2 | 0 | 0 | 0.81 | 0 | 0 | [5 empty] | PRT rng 2 | 0 | 0 | 0.81 | 0 | 0 | [5 empty] | S rng 1 | 0 | 1.31 | 0 | 0 | 0
Gauge: | N/A |  |  |  |  | Recalibration Date: | N/A | [8 empty] | S rng 1 | 0 | 1.3068 | 0 | 0 | 0 | PRT rng 3 | 0 | 0 | 0.21 | 0 | 0 | PRT rng 3 | 0 | 0 | 0.24 | 0 | 0 | PRT rng 3 | 0 | 0 | 0.26 | 0 | 0 | S rng 1 | 0 | 1.3 | 0 | 0 | 0 | [5 empty] | S rng 1 | 0 | 1.3 | 0 | 0 | 0 | [5 empty] | S rng 2 | 0 | 0 | 1.44 | 0 | 0
Temperature: | 72 |  |  |  |  | Relative Humidity: | 25% | [8 empty] | S rng 2 | 0 | 0 | 2 | 0 | 0 | S rng 1 | 0 | 1.0166 | 0 | 0 | 0 | S rng 1 | 0 | 1.31 | 0 | 0 | 0 | S rng 1 | 0 | 1.31 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 1.6 | 0 | 0 | [5 empty] | S rng 2 | 0 | 0 | 1.3 | 0 | 0
[16 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 1.5 | 0 | 0 | S rng 2 | 0 | 0 | 1.5 | 0 | 0 | S rng 2 | 0 | 0 | 1.6 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0
| Calibration Procedure(s): | WP4007 |  |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading | [14 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 4 | 0 | 0 | 0 | 0 | 0
| Controlling Specification(s): | AMS2750H, ASTM E220-19 |  |  |  | Tolerance Source: | Specification |  |  | raw uncerts | 0.2 | 0.21 | 0.21 | 0 | 0 | [6 empty] | S rng 4 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | Sref Rng 1 | 0 | 1.02 | 0 | 0 | 0 | [5 empty] | Sref Rng 1 | 0 | 1.01 | 0 | 0 | 0
[6 empty] | Condition of Item: | New |  |  |  | PRT | PRT | PRT | 0 | 0 | [24 empty] | Sref Rng 2 | 0 | 0 | 1.02 | 0 | 0 | [5 empty] | Sref Rng 2 | 0 | 0 | 1.01 | 0 | 0
[40 empty] | Sref Rng 3 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 3 | 0 | 0 | 0 | 0 | 0
Uncertainty @k=2 |  | 0.2 |  | 0.21 |  | 0.21 |  |  |  | Uncerts | 0.2 | 0.21 | 0.21 | [26 empty] | Sref Rng 4 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 4 | 0 | 0 | 0 | 0 | 0
Test Temperatures |  | -50 |  | 200 |  | 450 |  |  |  | Test Temps | -50 | 200 | 450
SN: 072513A | Beginning | -51.1 |  | 199.1 |  | 448.9 |  |  |  | E | 0.19 | 0.19 | 0.18 | 0 | 0
| End | -51 |  | 199.1 |  | 449 |  |  |  | J | 0.2 | 0.21 | 0.21 | 0 | 0
| C.F. | 1.1 |  | 0.9 |  | 1.1 |  |  |  | K | 0.22 | 0.24 | 0.24 | 0 | 0 |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base | Item | Range | Meas Units | Uncert Units | Base |  | Item | Range | Meas Units | Uncert Units | Base
[10 empty] | N | 0.28 | 0.26 | 0.26 | 0 | 0 |  | TC E with PRT | -112 to -22 | °F | °F | 0.19 |  | TC J with PRT | -112 to -22 | °F | °F | 0.2 |  | TC K with PRT | -112 to -22 | °F | °F | 0.22 |  | TC N with PRT | -112 to -22 | °F | °F | 0.28 |  | TC R with PRT | 32 to 392 | °F | °F | 0.81 | TC R with S ref | 32 to 392 | °F | °F | 1.02 |  | TC S with PRT | 32 to 392 | °F | °F | 0.81 | TC S with S ref | 32 to 392 | °F | °F | 1.01 |  | TC T with PRT | -112 to -22 | °F | °F | 0.23
[10 empty] | R | 0 | 0.81 | 0.81 | 0 | 0 |  | TC E with PRT | -22 to 392 | °F | °F | 0.19 |  | TC J with PRT | -22 to 392 | °F | °F | 0.21 |  | TC K with PRT | -22 to 392 | °F | °F | 0.24 |  | TC N with PRT | -22 to 392 | °F | °F | 0.26 |  | TC R with PRT | 392 to 752 | °F | °F | 0.81 | TC R with S ref | 392 to 1200 | °F | °F | 1.02 |  | TC S with PRT | 392 to 752 | °F | °F | 0.81 | TC S with S ref | 392 to 1200 | °F | °F | 1.01 |  | TC T with PRT | -22 to 392 | °F | °F | 0.2
[10 empty] | S | 0 | 0.81 | 0.81 | 0 | 0 |  | TC E with PRT | 392 to 752 | °F | °F | 0.18 |  | TC J with PRT | 392 to 752 | °F | °F | 0.21 |  | TC K with PRT | 392 to 752 | °F | °F | 0.24 |  | TC N with PRT | 392 to 752 | °F | °F | 0.26 |  | TC R with S std | 32 to 392 | °F | °F | 1.3 | TC R with S ref | 1200 to 2000 | °F | °F | 1.2 |  | TC S with S std | 32 to 392 | °F | °F | 1.3 | TC S with S ref | 1200 to 2000 | °F | °F | 1.2 |  | TC T with PRT | 392 to 752 | °F | °F | 0.2
[10 empty] | T | 0.23 | 0.2 | 0.2 | 0 | 0 |  | TC E with S | 32 to 392 | °F | °F | 1.3068 |  | TC J with S | 32 to 392 | °F | °F | 1.0166 |  | TC K with S | 32 to 392 | °F | °F | 1.31 |  | TC N with S std | 32 to 392 | °F | °F | 1.31 |  | TC R with S std | 392 to 1200 | °F | °F | 1.6 | TC R with S ref | 2000 to 2500 | °F | °F | 3.2 |  | TC S with S std | 392 to 1200 | °F | °F | 1.3 | TC S with S ref | 2000 to 2500 | °F | °F | 3.2 |  | TC T with S std | 32 to 392 | °F | °F | 1.31
[17 empty] | TC E with S | 392 to 1200 | °F | °F | 2 |  | TC J with S | 392 to 1200 | °F | °F | 1.5 |  | TC K with S std | 392 to 1200 | °F | °F | 1.5 |  | TC N with S std | 392 to 1200 | °F | °F | 1.6 |  | TC R with S std | 1200 to 2000 | °F | °F | 2 | [6 empty] | TC S with S std | 1200 to 2000 | °F | °F | 1.6 | [6 empty] | TC T with S std | 392 to 752 | °F | °F | 1.44
[10 empty] | raw uncerts | 0 | 0 | 0 | 0 | 0 |  | TC E with S | 1200 to 1832 | °F | °F | 2.2 |  | TC J with S | 1200 to 2000 | °F | °F | 2.0018 |  | TC K with S std | 1200 to 2000 | °F | °F | 2 |  | TC N with S std | 1200 to 2000 | °F | °F | 2.1 |  | TC R with S std | 2000 to 2500 | °F | °F | 3.23 | [6 empty] | TC S with S std | 2000 to 2500 | °F | °F | 2 | [5 empty] | PRT | 0 | 0 | 0 | 0 | 0
[11 empty] | 0 | 0 | 0 | 0 | 0 | PRT | 0 | 0.19 | 0.18 | 0 | 0 |  | TC J with S | 2000 to 2192 | °F | °F | 2.7 |  | TC K with S std | 2000 to 2500 | °F | °F | 3.2 |  | TC N with S std | 2000 to 2372 | °F | °F | 2.9 |  | 0 | 0.81 | 0.81 | 0 | 0 | [5 empty] | PRT | 0 | 0 | 0 | 0 | 3.2 | [5 empty] | S | 0 | 0 | 0 | 0 | 0
[16 empty] | S | 0 | 0 | 0 | 0 | 0 | PRT | 0 | 0 | 0 | 0 | 0 | PRT | 0 | 0 | 0 | 0 | 0 | PRT | 0 | 0 | 0 | 0 | 0 | PRT | 0 | 0 | 0 | 0 | 0 | [5 empty] | S std | 0 | 0 | 0 | 0 | 0 | [5 empty] | PRT rng 1 | 0 | 0 | 0 | 0 | 0
[10 empty] | Uncerts | [5 empty] | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | S | 0 | 0 | 0 | 0 | 0 | S | 0 | 0 | 0 | 0 | 0 | S | 0 | 0 | 0 | 0 | 0 | S std | 0 | 0 | 0 | 0 | 0 | [5 empty] | S ref | 0 | 0 | 0 | 0 | 0 | [5 empty] | PRT rng 2 | 0 | 0 | 0 | 0 | 0
[10 empty] | Test Temps | [5 empty] | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | S ref | 0 | 0 | 0 | 0 | 0 | [5 empty] | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | [5 empty] | PRT rng 3 | 0 | 0 | 0 | 0 | 0
[10 empty] | E | 0 | 0 | 0 | 0 | 0 | PRT rng 3 | 0 | 0 | 0 | 0 | 0 | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | PRT rng 1 | 0 | 0 | 0 | 0 | 0 | [5 empty] | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 1 | 0 | 0 | 0 | 0 | 0
[10 empty] | J | 0 | 0 | 0 | 0 | 0 | S rng 1 | 0 | 0 | 0 | 0 | 0 | PRT rng 3 | 0 | 0 | 0 | 0 | 0 | PRT rng 3 | 0 | 0 | 0 | 0 | 0 | PRT rng 3 | 0 | 0 | 0 | 0 | 0 | PRT rng 2 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 1 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 2 | 0 | 0 | 0 | 0 | 0
The JGI quality system governing this calibration/inspection is based upon and maintained in compliance with quality specifications ISO/IEC 17025, ANSI/NCSL Z540-1 and ISO 10012. | [9 empty] | K | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 0 | 0 | 0 | S rng 1 | 0 | 0 | 0 | 0 | 0 | S rng 1 | 0 | 0 | 0 | 0 | 0 | S rng 1 | 0 | 0 | 0 | 0 | 0 | S rng 1 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 2 | 0 | 0 | 0 | 0 | 0
[10 empty] | N | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 0 | 0 | 0 | S rng 2 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection. | [9 empty] | R | 0 | 0 | 0 | 0 | 0 | [6 empty] | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | S rng 3 | 0 | 0 | 0 | 0 | 0 | [5 empty] | S rng 4 | 0 | 0 | 0 | 0 | 0
[10 empty] | S | 0 | 0 | 0 | 0 | 0 | [6 empty] | S rng 4 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | S rng 4 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 1 | 0 | 0 | 0 | 0 | 0
[10 empty] | T | 0 | 0 | 0 | 0 | 0 | [24 empty] | Sref Rng 1 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 2 | 0 | 0 | 0 | 0 | 0
[40 empty] | Sref Rng 2 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 3 | 0 | 0 | 0 | 0 | 0
[40 empty] | Sref Rng 3 | 0 | 0 | 0 | 0 | 0 | [5 empty] | Sref Rng 4 | 0 | 0 | 0 | 0 | 0
[40 empty] | Sref Rng 4 | 0 | 0 | 0 | 0 | 0
Comments: | Spool #G150221-1-4 1000'
Found in tolerance. No action required. "C.F" = "Correction Factor."
RESULTS: | ACCEPTED
TEST EQUIPMENT
Test Equipment |  | S/N |  | Date Certified |  | Date Due | Traceability Number
Additel 286 |  | 6851022030037 |  | 2024-07-01 |  | 2025-07-01 | 56561-060578
Fluke 1502A |  | B55852 |  | 2025-01-20 |  | 2026-01-20 | 56561-068751
PRT |  | 496142 |  | 2023-12-15 |  | 2025-12-15 | G295632
[10 empty] | raw uncerts | 0.2 | 0.21 | 0.21 | 0 | 0
[11 empty] | PRT | PRT | PRT | 0 | 0
[10 empty] | Uncerts | 0.2 | 0.21 | 0.21 | #NUM! | #NUM!
[10 empty] | Test Temps | -50 | 200 | 450
[10 empty] | E | 0.19 | 0.19 | 0.18 | 0 | 0
[10 empty] | J | 0.2 | 0.21 | 0.21 | 0 | 0
[10 empty] | K | 0.22 | 0.24 | 0.24 | 0 | 0
[10 empty] | N | 0.28 | 0.26 | 0.26 | 0 | 0
[10 empty] | R | 0 | 0.81 | 0 | 0 | 0
[10 empty] | S | 0 | 0.81 | 0.81 | 0 | 0
[10 empty] | T | 0.23 | 0.2 | 0.2 | 0 | 0

## CERT, Wireset
Wireset Number: | J100SET0525 | [7 empty] | password: jgi
Inspector Making Set: | JR
Date Made: | 2025-05-01
S/Ns (ex. - "J301-314"): | J101-J110
Number of Sensors: | 10
Johnson Gage and Inspection, Inc.
5920 W. 21st St. N., Wichita, KS 67205
Phone 316.943.7532; Fax 316.944.6256
EXPENDABLE WIRESET CERTIFICATE
Report Number: | J100SET0525 |  |  |  | Page 1 of 1 |  |  | Form 4007
Customer: | JOHNSON GAGE & INSPECTION, INC. | [6 empty] | Rev. F 10-20-20
Item: | THERMOCOUPLE SET |  |  |  |  | Inspector Making Set: | JR
Type: | "J" |  |  |  |  | Date Made: | 2025-05-01
Gauge: | N/A |  |  |  |  | Due for Retirement: | 2025-08-01
Serial Numbers: | J101-J110 |  |  |  |  | Number of Sensors: | 10
The above expendable wireset was made from wire roll 072513A. Test data for this wire roll is given below.
Test Date: | 2025-04-09 |  | Temperature: |  | 72 |  | Relative Humidity: | 25%
| Calibration Procedure(s): | WP4007 |  |  |  | Tolerance: | Greater of ±2.0°F or ±0.4% of Reading
| Controlling Specification(s): | AMS2750H, ASTM E220-19 |  |  |  | Tolerance Source: | Specification
[6 empty] | Condition of Item: | New
Uncertainty @k=2 |  | 0.2 |  | 0.21 |  | 0.21
Test Temperatures |  | -50 |  | 200 |  | 450
J101-J110 | Beginning | -51.1 |  | 199.1 |  | 448.9
| End | -51 |  | 199.1 |  | 449
| C.F. | 1.1 |  | 0.9 |  | 1.1
The JGI quality system governing this calibration/inspection is based upon and maintained in compliance with quality specifications ISO/IEC 17025, ANSI/NCSL Z540-1 and ISO 10012.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection.
Comments:
Found in tolerance. No action required. "C.F" = "Correction Factor."
RESULTS: | ACCEPTED
TEST EQUIPMENT
Test Equipment |  | S/N |  | Date Certified |  | Date Due | Traceability Number
Additel 286 |  | 6851022030037 |  | 2024-07-01 |  | 2025-07-01 | 56561-060578
Fluke 1502A |  | B55852 |  | 2025-01-20 |  | 2026-01-20 | 56561-068751
PRT |  | 496142 |  | 2023-12-15 |  | 2025-12-15 | G295632

## Customers
Company Name | Street | PO Box | City | State | Street Zip Code | PO Box Zip Code | Selected Customer
Company Name | Street | PO Box | City | State | Street Zip Code | PO Box Zip Code | JOHNSON GAGE & INSPECTION, INC. | Customer
3P PROCESSING, INC. | 1702 S. KNIGHT |  | WICHITA | KS | 67213 |  | 5920 W. 21st ST. N. | St. Address
AADFW, INC. | 1350 WESTPARK WAY |  | EULESS | TX | 76040 |  |  | PO Box
ABSOLUTE DIMENSIONS | 3838 W. MAY ST. |  | WICHITA | KS | 67213 |  | WICHITA | City
ACCURUS AEROSPACE WICHITA, LLC | 4011 E. 31 ST. SOUTH |  | WICHITA | KS | 67210 |  | KS | State
AERO METAL FORMS, INC. | 2649 ESTHNER COURT |  | WICHITA | KS | 67213 |  | 67205 | Street Zip
AEROSPACE SYSTEMS & COMPONENTS | 5201 E. 36TH ST. N. |  | WICHITA | KS | 67220-3206 |  |  | PO Box Zip
AERO-TECH ENGINEERING, INC. | 5555 N 119TH ST W |  | MAIZE | KS | 67101 |  | 3rd line concatenation
AIRTECH SUPPLY | 3058 HIGHWAY 290 |  | HOT SPRINGS | AR | 71913 |  | WICHITA, KS 67205
ALLTITE | 1600 E. MURDOCK ST. |  | WICHITA | KS | 67214
AMETEK ADVANCED INDUSTRIES, INC. | 4550 SOUTHEAST BOULEVARD |  | WICHITA | KS |  | 67216
AMETEK B & S AIRCRAFT PARTS & ACC., INC. | 1414 S. MOSLEY |  | WICHITA | KS | 67211-3387
APEX COMPOSITES - OKLAHOMA DIV. | 414 CHAMBER LOOP |  | ADA | OK | 74821
APEX ENGINEERING - WICHITA DIV. | 1804 WEST 2ND STREET |  | WICHITA | KS | 67203-5703
AVIATION CONSULTING & ENGINEERING SOL | 202 NORTH PARK |  | MAIZE | KS | 67101
B & C SPECIALTY PRODUCTS | 123 E. 4TH |  | NEWTON | KS | 67114
BG PRODUCTS, INC | P.O. BOX 1282 |  | WICHITA | KS | 67201
BODYCOTE THERMAL PROCESSING, INC.-McLEAN | 1019 McLEAN BLVD SOUTH |  | WICHITA | KS | 67213
BODYCOTE THERMAL PROCESSING, INC.-WEST | 1009 S. WEST STREET |  | WICHITA | KS | 67213
BRITTAIN MACHINE, INC. | 2621 S. SHERIDAN (door 12) | P.O. BOX 17227 | WICHITA | KS | 67217
BURNHAM COMPOSITE STRUCTURES, INC. | 6262 W. 34TH STREET SOUTH |  | WICHITA | KS | 67215
CAPPS MANUFACTURING INC. | 2121 S. EDWARDS |  | WICHITA | KS | 67213
CAPPS MANUFACTURING INC. - CUSTER | 2120 S. CUSTER |  | WICHITA | KS | 67213
CATALYTIC INDUSTRIAL GROUP | 713 N. 20TH STREET |  | INDEPENDENCE | KS | 67301
CELCO | 103 N. LAWRENCE |  | ANTHONY | KS | 67003
CENTER INDUSTRIES CORP. | 2505 S. CUSTER | P.O. BOX 17364 | WICHITA | KS |  | 67217-0364
CESSNA AIRCRAFT CO., DEPT. 352 | P.O. BOX 12917 |  | WICHITA | KS | 67277
CESSNA DEPT 074 | 5800 E. PAWNEE |  | WICHITA | KS | 67218
CHARLES ENGINEERING | 10400 S. 119TH ST. WEST | P.O. BOX 369 | CLEARWATER | KS | 67026
CHROME PLUS INTERNATIONAL, INC. - CUSTER FACILITY | 3939 W. 29TH ST. S. |  | WICHITA | KS | 67217
CHROME PLUS INTERNATIONAL, INC. - WICHITA FACILITY | 3939 W. 29TH ST. S. |  | WICHITA | KS | 67217
CLARK MANUFACTURING | 1936 NORTH A |  | WELLINGTON | KS |  | 67152
CNH AMERICA L.L.C. - CALIBRATION LAB | 3301 S. HOOVER RD. |  | WICHITA | KS | 67215
COOPER MACHINERY SERVICES | 1648 W. MAGNOLIA RD. |  | SALINA | KS | 67401
COX MACHINE - HARPER DIVISION | 949 N. HIGHWAY 14 |  | HARPER | KS | 67058
COX MACHINE - WICHITA DIVISION | 5338 W. 21ST N., SUITE 100 |  | WICHITA | KS | 67205
D-J AEROSPACE, INC. PHILIPPINES | #12 Aim High Corner Commitment Street | Subic Bay Industrial Par, Phase 1 | Subic Bay Freeport Zone | PHILIPPINES | 67010
D-J ENGINEERING, INC. | 219 W. SIXTH | P.O. BOX 278 | AUGUSTA | KS |  | 67010-0278
D-J EXTRUDING | 723 E. SPRING |  | CONWAY SPRINGS | KS | 67031
DUCOMMUN AEROSTRUCTURES, INC. | 3333 MAIN STREET |  | PARSONS | KS | 67357
EXACTA AEROSPACE, INC. | 4200 W. HARRY |  | WICHITA | KS | 67209-2730
EXACTA AEROSPACE, INC. SOUTH | 4261 S. WEST STREET |  | WICHITA | KS | 67217-3813
EXCEL INDUSTRIES | 200 S RIDGE RD | P.O. BOX 7000 | HESSTON | KS | 67062
FIBER DYNAMICS | 3730 MIDCO |  | WICHITA | KS | 67215
FIGEAC AERO | 9313 EAST 39TH STREET NORTH |  | WICHITA | KS | 67226
FREEMAN COMPANY | P.O. BOX 57 |  | YANKTON | SD | 57078
GKN PRECISION MACHINING, INC. | 429 N WEST ROAD |  | WELLINGTON | KS | 67152-8047
GLOBE ENGINEERING CO. | 1539 S. ST. PAUL |  | WICHITA | KS | 67277
GLOBE ENGINEERING CO. - MAIZE RD | 3505 S. MAIZE RD |  | WICHITA | KS | 67215
GLOBE ENGINEERING CO.-SHERIDAN LOC. | 2432 S. SHERIDAN |  | WICHITA | KS | 67217
GOODRICH CABIN SYSTEMS-PCI | 1644 S MAIZE RD | P.O. BOX 75278 | WICHITA | KS |  | 67275
HARLOW AEROSTRUCTURES LLC | 1501 S. MCLEAN BLVD | P.O. BOX 304 | WICHITA | KS |  | 67201-0304
HEARTLAND PRECISION FASTENERS INC. | 301 PRAIRIE VILLAGE DR. |  | NEW CENTURY | KS | 66031
JOHNSON CONTROLS | 404 S. HOLLAND |  | WICHITA | KS | 67209
JOHNSON CONTROLS-YORK |  | PO BOX 19014 | WICHITA | KS | 67204-9014
JOHNSON GAGE & INSPECTION, INC. | 5920 W. 21st ST. N. |  | WICHITA | KS | 67205
KAMAN COMPOSITES - WICHITA, INC. | 1650 S McCOMAS |  | WICHITA | KS | 67213
KANSAS COIL SPRING CO., INC. | 4310 SE BLVD |  | WICHITA | KS | 67210
LEADING TECHNOLOGY COMPOSITES | 2626 W. MAY |  | WICHITA | KS | 67213
LEE AEROSPACE, INC. | 9323 E 34th ST N |  | WICHITA | KS | 67226
LEE PRODUCTS ENGINEERING SERVICES, INC. | 2946 S. HYDRAULIC |  | WICHITA | KS | 67216
LEONARD'S METAL, INC. - WICHITA | 2621 ESTHNER CT. |  | WICHITA | KS | 67213
LTC INC. - PLANT 3 | 2700 S CUSTER |  | WICHITA | KS | 67217
MANUFACTURING DEVELOPMENT, INC. | 37515 WEST 15th SOUTH | P.O. BOX 550 | CHENEY | KS |  | 67025-0550
MAX AEROSTRUCTURES, LLC | 8219 W IRVING ST |  | WICHITA | KS | 67209
MAYNARD INC. | 7175 S. McGUIRE ST. |  | FAYETTEVILLE | AR | 72704-5237
MAYNARD INC. - WICHITA DIV. | 4422 W. WALKER ST. |  | WICHITA | KS | 67209-2657
McGINTY MACHINE CO., INC. | 222 N. HYDRAULIC |  | WICHITA | KS | 67214-4818
MERRITT PREFERRED COMPONENTS, INC. | PO BOX 1209 |  | KILGORE | TX | 75663
METAL FINISHING CO. | 1425 S. McLEAN BLVD. |  | WICHITA | KS | 67213
METAL FINISHING CO. - MURDOCK LOCATION | 701 E. MURDOCK |  | WICHITA | KS | 67214
METAL-FAB INC. | 3025 MAY AVE. |  | WICHITA | KS | 67213
MFCO HEAT TREAT | MFCO HEAT TREAT | MFCO HEAT TREAT | MFCO HEAT TREAT | MFCO HEAT TREAT | MFCO HEAT TREAT
MILLENNIUM CONCEPTS, INC. | 9050 MONROE CIRCLE |  | WICHITA | KS | 67209
NATIONAL INSTITUTE FOR AVIATION RESEARCH | 4004 N. WEBB ROAD |  | WICHITA | KS | 67226
NIAR - ADVANCED COATINGS LAB, RM M151 | 4004 N. WEBB RD |  | WICHITA | KS | 67226
NIAR - ENVIRONMENTAL TEST LAB | 1845 FAIRMOUNT |  | WICHITA | KS | 67260
NIAR ASTEC | 1229 E 85TH ST N |  | PARK CITY | KS | 67219
NIAR COMPOSITE LAB, RM-110 | 1845 FAIRMONT | 93 | WICHITA | KS | 67260-0093
NIAR COMPOSITE LAB, RM-118 | 1845 FAIRMONT, RM-118 | BOX 93 | WICHITA | KS | 67260-0093
NIAR-QUALITY CONTROL & STRAIN GAGE LAB | 1845 FAIRMOUNT |  | WICHITA | KS | 67260-0093
NORTH AMERICAN AVIATION, INC. | 7330 N. BROADWAY |  | WICHITA | KS | 67219
NWI KANSAS CITY LLC | 31800 W 196th ST |  | EDGERTON | KS | 66021
NWI KANSAS CITY LLC | 4020 E 138th ST |  | GRANDVIEW | MO | 64030
NWI WICHITA LLC | 3258 S HOOVER | P.O. BOX 12670 | WICHITA | KS |  | 67277
O & F MACHINE PRODUCTS INC. | 3020 WEST 20TH |  | JOPLIN | MO | 64802
OXWELL, INC. | 600 E 16th | P.O. BOX 588 | WELLINGTON | KS | 67152
PARK AEROSPACE TECHNOLOGIES CORP. | 486 N. OLIVER ROAD, BLDG. Z |  | NEWTON | KS | 67114
PCI - PERFORMANCE CONTRACTING, INC. | 16047 W. 110TH ST. |  | LENEXA | KS | 66219
PMC CORPORATION | 680 HAYWARD STREET |  | MANCHESTER | NH | 03103
PREMIER PROCESSING | 3002 WEST PAWNEE, STE 104 |  | WICHITA | KS | 67213-1806
PRIDE PLATING INC. | 2900 E HIGHWAY 10 |  | GROVE | OK | 74344
PROBATA METROLOGY & CALIBRATION SERVICES | 14600 METRO PLAZA BLVD. |  | EDMOND | OK | 73013
PRODUCT MANUFACTURING CORPORATION | 4425 W MAY ST., BLDG A |  | WICHITA | KS | 67209
RAPID PROCESSING SOLUTIONS | 1367 S. ANNA |  | WICHITA | KS | 67209
RV PRODUCTS DIVISION | 3050 N. ST. FRANCIS |  | WICHITA | KS | 67219
SHACKELFORD MACHINE | 116 S TRACY | P.O. BOX 670 | CLEARWATER | KS |  | 67026
SIEMENS ENERGY, INC. - E R WP SCM MF1 2 | 100 COMMERCE ST. |  | HUTCHINSON | KS | 67501
SIEMENS-WICHITA TOOLING CENTER | 1090 E 37TH ST N |  | WICHITA | KS | 67219
SIGMA-TEK, INC. | 1001 INDUSTRIAL |  | AUGUSTA | KS | 67010-9566
STANDRIDGE COLOR CORPORATION | 1011 INDUSTRIAL DR. |  | NEWTON | KS | 67114
TECT AEROSPACE PARK CITY | 5545 N. MILL HEIGHTS DRIVE |  | PARK CITY | KS | 67219
TECT AEROSPACE WELLINGTON | 1515 N. A ST. |  | WELLINGTON | KS | 67152
TEXTRON AVIATION INC | 2 CESSNA BLVD |  | WICHITA | KS | 67215
TEXTRON AVIATION - BEECHCRAFT FACILITY | 10511 E. CENTRAL |  | WICHITA | KS | 67206
THE BOEING CO.--IDS WICHITA | 4615 S. OLIVER |  | WICHITA | KS | 67210
THE BOEING COMPANY | 8900 FROST AVE. |  | BERKELEY | MO | 63134
TRIUMPH AEROSPACE SYSTEMS - WICHITA DIVISION
TRIUMPH FABRICATIONS-SHELBYVILLE
TRIUMPH STRUCTURES - EAST TEXAS, INC.
TRIUMPH STRUCTURES - KANSAS CITY
UNIVERSAL PRODUCTS | 521 INDUSTRIAL STREET |  | GODDARD | KS | 67052
VALCO MANUFACTURING CO., INC. | 925 BOREN AVE. |  | DUNCAN | OK | 73533
VEGAS FASTENER MANUFACTURING | 4315 W. OQUENDO RD. |  | LAS VEGAS | NV | 89118
VIEGA LLC | 2211 VIEGA AVE. |  | McPHERSON | KS | 67460
WEAVER MANUFACTURING, INC. | 1005 E 17th |  | WICHITA | KS | 67214
WEBSTER COMBUSTION TECHNOLOGY LLC | 619 INDUSTRIAL ROAD |  | WINFIELD | KS | 67156-9123
WESCO AIRCRAFT - ELECTRONIC PRODUCTS GROUP | 3851 N. WEBB RD. |  | WICHITA | KS | 67266
WICHITA AEROSPACE SPARES, INC. | 4626 W. MAY ST., BUILDING D |  | WICHITA | KS | 67209
WICHITA BRASS & ALUMINUM FOUNDRY, INC. | 412 E 29th ST N | P.O. BOX 4144 | WICHITA | KS |  | 67204-0144
YINGLING AVIATION | 2010 AIRPORT ROAD | P.O. BOX 9248 | WICHITA | KS | 67277-0248
YINGLING AVIATION PROP SHOP | 1440 S. RIDGE ROAD |  | WICHITA | KS | 67209

## formulas
NTS-10 |  | 850 | 900 | 950 |  | 1050 | 1100 | 1150
TEMP. | 800 |  |  |  | 1000 |  |  |  | 1200
Actual | 799.8139 |  |  |  | 999.6454 |  |  |  | 1199.689
C.F. | 0.18605 | 0.22818 | 0.27032 | 0.31245 | 0.35458 | 0.34367 | 0.33277 | 0.32187 | 0.31096
|  |  | 4 | 4 |  |  | 4 | 4
|  |  | 2 | 3 |  |  | 2 | 3
NTS-12 |  | 450 | 500 | 550 |  | 650 | 700 | 750
TEMP. | 400 |  |  |  | 600 |  |  |  | 800
Actual | 199.6775 |  |  |  | 399.6341 |  |  |  | 599.734
C.F. | 200.3225 | 200.3334 | 200.3442 | 200.3551 | 200.3659 | 200.3409 | 200.3159 | 200.2909 | 200.266
|  |  | 4 | 4 |  |  | 4 | 4
|  |  | 2 | 3 |  |  | 2 | 3
NTS-7 |  | 1250 | 1300 | 1350 |  | 1450 | 1500 | 1550 |  | 1650 | 1700 | 1750 |  | 1850 | 1900 | 1950 |  | 2050 | 2100 | 2150
TEMP. | 1200 |  |  |  | 1400 |  |  |  | 1600 |  |  |  | 1800 |  |  |  | 2000 |  |  |  | 2200
Actual | 32.7333 |  |  |  | 199.6775 |  |  |  | 399.6341 |  |  |  | 599.734 |  |  |  | 799.8139 |  |  |  | 999.6454
C.F. | 1167.2667 | 1175.5306 | 1183.7946 | 1192.0586 | 1200.3225 | 1200.3334 | 1200.3442 | 1200.3551 | 1200.3659 | 1200.3409 | 1200.3159 | 1200.2909 | 1200.266 | 1200.246 | 1200.226 | 1200.206 | 1200.1861 | 1200.2282 | 1200.2703 | 1200.3124 | 1200.3546
|  |  | 4 | 4 |  |  | 4 | 4 |  |  | 4 | 4 |  |  | 4 | 4 |  |  | 4 | 4
|  |  | 2 | 3 |  |  | 2 | 3 |  |  | 2 | 3 |  |  | 2 | 3 |  |  | 2 | 3
NTS-11 |  | 50 | 70 | 100 | 150 |  | 250 | 300 | 350
TEMP. | 32 |  |  |  |  | 200 |  |  |  | 400
Actual | 32.7318 |  | 70.7318 |  |  | 200.7318 |  |  |  | 32.7333
C.F. | -0.73183 | -0.73183 | -0.73183 | -0.73183 | -0.73183 | -0.73183 | 91.2678 | 183.2674 | 275.267 | 367.2667
[8 empty] | 4 | 4
[8 empty] | 2 | 3
TTS-1 |  | -40 | -30 | -20 | -10 | 0 | 10 | 20 |  | 50 | 70 | 90 | 110
TEMP. | -50 | [7 empty] | 32 |  |  |  |  | 120
Actual | -18 | [7 empty] | 64 |  |  |  |  | 152
C.F. | -32 | -32 | -32 | -32 | -32 | -32 | -32 | -32 | -32 | -32 | -32 | -32 | -32 | -32
|  |  | 8 | 8 | 8 | 8 | 8 | 8 |  | 44 | 44 | 44 | 44
|  |  | 2 | 3 | 4 | 5 | 6 | 7 |  | 9 | 19 | 29 | 39
SP-1 |  | 50 | 100 | 150 |  | 250 | 300 | 350 |  | 450 | 500 | 550 |  | 650 | 700 | 750 |  | 850 | 900 | 950 |  | 1050 | 1100 | 1150 |  | 1250 | 1300 | 1350 |  | 1450 | 1500 | 1550 |  | 1650 | 1700 | 1750 |  | 1850 | 1900 | 1950 |  | 2050 | 2100 | 2150 |  | 2250 | 2300 | 2350
TEMP | 32 |  |  |  | 200 |  |  |  | 400 |  |  |  | 600 |  |  |  | 800 |  |  |  | 1000 |  |  |  | 1200 |  |  |  | 1400 |  |  |  | 1600 |  |  |  | 1800 |  |  |  | 2000 |  |  |  | 2200 |  |  |  | 2400
ACTUAL | 32.7333 |  |  |  | 201.0775 |  |  |  | 400.2841 |  |  |  | 599.934 |  |  |  | 800.0139 |  |  |  | 1000.1554 |  |  |  | 1200.839 |  |  |  | 1400.17 |  |  |  | 1599.7961 |  |  |  | 1799.6955 |  |  |  | 2001.7175 |  |  |  | 2201.7619 |  |  |  | #VALUE!
C.F | -0.73333 | -0.7702 | -0.87262 | -0.97504 | -1.0775 | -0.87912 | -0.68078 | -0.48244 | -0.28409 | -0.19658 | -0.10907 | -0.021551 | 0.065963 | 0.045986 | 0.026008 | 0.0060309 | -0.013947 | -0.049316 | -0.084685 | -0.12005 | -0.15542 | -0.32633 | -0.49723 | -0.66813 | -0.83904 | -0.67177 | -0.5045 | -0.33723 | -0.16996 | -0.076486 | 0.016986 | 0.11046 | 0.20393 | 0.22907 | 0.2542 | 0.27934 | 0.30448 | -0.20102 | -0.70652 | -1.212 | -1.7175 | -1.7286 | -1.7397 | -1.7508 | -1.7619 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | -51.1 | 199.1 | 448.9
Test temps |  | -50 |  |  |  | 0 |  | 0 |  | 200 | [7 empty] | 450 | [15 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0
sns | TC type
072513A | J | 0 | -51.1 |  |  | 21 | 1 | 072513A | J |  |  | 199.1 |  | 21 | 1 | 072513A | J | 0 | 448.9 |  |  | 21 | 1 | 072513A | J | 0 | 32 | 0 |  | 21 | 1 | 072513A | J | 0 | 32 | 0 |  | 21 | 5 | 0 | J | 0 | 32 | 0 |  | 21 | 5 |  | J | 0 | 32 | 0 | 21 | 5 | J | 0 | 32 | 0 | 21 | 5 | J | 0 | 32 | 0 | 21 | 5 | J | 0 | 32 | 0 | 21 | 9 | J | 0 | 32 | 0 | 21 | 9 | J | 0 | 32 | 0 | 21 | 9 | J | 0 | 32 | 0 | 21 | 9 | J | 0 | 32 | 0 | 21 | 9 | J | 0 | 32 | 0 | 21 | 13 | J | 0 | 32 | 0 | 21 | 13 | J | 0 | 32 | 0 | 21 | 13 | J | 0 | 32 | 0 | 21 | 13 | J | 0 | 32 | 0 | 21 | 13 | J | 0 | 32 | 0
0 |  | 0 | -51 |  |  | 22 | 1 |  |  |  |  | 199.1 |  | 22 | 1 |  |  | 0 | 449
0 |  | 0 | 0 |  |  | 23 | 1 |  |  |  |  | 0 |  | 23 | 1 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 24 | 1 |  |  |  |  | 0 |  | 24 | 1 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 25 | 1 |  |  |  |  | 0 |  | 25 | 1 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 26 | 2 |  |  |  |  | 0 |  | 26 | 2 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 27 | 2 |  |  |  |  | 0 |  | 27 | 2 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 28 | 2 |  |  |  |  | 0 |  | 28 | 2 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 29 | 2 |  |  |  |  | 0 |  | 29 | 2 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 30 | 2 |  |  |  |  | 0 |  | 30 | 2 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 31 | 3 |  |  |  |  | 0 |  | 31 | 3 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 32 | 3 |  |  |  |  | 0 |  | 32 | 3 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 33 | 3 |  |  |  |  | 0 |  | 33 | 3 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 34 | 3 |  |  |  |  | 0 |  | 34 | 3 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 35 | 3 |  |  |  |  | 0 |  | 35 | 3 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 36 | 4 |  |  |  |  | 0 |  | 36 | 4 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 37 | 4 |  |  |  |  | 0 |  | 37 | 4 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 38 | 4 |  |  |  |  | 0 |  | 38 | 4 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 39 | 4 |  |  |  |  | 0 |  | 39 | 4 |  |  | 0 | 0
0 |  | 0 | 0 |  |  | 40 | 4 |  |  |  |  | 0 |  | 40 | 4 |  |  | 0 | 0
Results for UUT Calculator. 1-20, 20 possible UUTs per temperature, up to six TC Types per UUT
Standards selection |  | 1 | [6 empty] | 1 | [5 empty] | E | J | K | N | S | T
and correction |  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  |  | 1 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
calculator, 1-20 |  | SP1 |  | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | [6 empty] | E | J | K | N | S | T
|  | SP3 | -50 | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! |  | 2 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | [6 empty] | E | J | K | N | S | T
Type "S" |  | NTS10 |  | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! |  | 3 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
System Correction: |  | TTS1 |  | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | [6 empty] | E | J | K | N | S | T
0.7 |  | ST7 |  | -16623.5199 | -16623.5199 | -16623.5 |  |  |  | 4 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | -50 | -50 | -50 | -50 | [8 empty] | E | J | K | N | S | T
|  | STD USD | mvDC | TMP | [5 empty] | 5 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | -50 | -50 | [10 empty] | E | J | K | N | S | T
|  | Put dollar sign in front of each cell reference in the Temperatures tab in the formula | [7 empty] | 6 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ex: | $B$14 | [11 empty] | E | J | K | N | S | T
[5 empty] | Set Selections for S system Correction using formula in F99 |  |  |  |  | 7 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[15 empty] | E | J | K | N | S | T
[10 empty] | 8 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[15 empty] | E | J | K | N | S | T
[10 empty] | 9 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 2 | [12 empty] | E | J | K | N | S | T
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  |  | 10 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 34392.7234 | #VALUE! | #VALUE! |  | 13 | [6 empty] | E | J | K | N | S | T
|  | SP3 | 200 | 34392.7234 | 34401.638 | 34401.6 |  | 7 |  | 11 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 9984.6358 | 10184.9017 | 10184.9 |  | 3 | [6 empty] | E | J | K | N | S | T
|  | NTS10 |  | 9984.6358 | 9984.9467 | 9984.9 |  | 3 |  | 12 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 7205.826 | 5688.81 | 5688.8 |  | 3 | [6 empty] | E | J | K | N | S | T
|  | ST7 |  | 35202.9573 | 35202.9573 | 35203 |  |  |  | 13 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 200 | 200 | 200 | 200 | [8 empty] | E | J | K | N | S | T
|  | STD USD | mvDC | TMP | [5 empty] | 14 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 200 | 200 | [10 empty] | E | J | K | N | S | T
[10 empty] | 15 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 3 | [12 empty] | E | J | K | N | S | T
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  |  | 16 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 77383.6276 | #VALUE! | #VALUE! |  | 13 | [6 empty] | E | J | K | N | S | T
|  | SP3 | 450 | 77383.6276 | 77403.6856 | 77403.7 |  | 7 |  | 17 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 22465.4305 | 22665.6965 | 22665.7 |  | 3 | [6 empty] | E | J | K | N | S | T
|  | NTS10 |  | 22465.4305 | 22465.7415 | 22465.7 |  | 3 |  | 18 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 16213.1085 | 12799.8225 | 12799.8 |  | 3 | [6 empty] | E | J | K | N | S | T
|  | ST7 |  | 79206.654 | 79206.654 | 79206.7 |  |  |  | 19 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 450 | 450 | 450 | 450 | [8 empty] | E | J | K | N | S | T
|  | STD USD | mvDC | TMP | [5 empty] | 20 | 0 | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 450 | 450 |  |  |  |  | 2 | [5 empty] | E | J | K | N | S | T
[9 empty] | 1 | 1 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 4 | [6 empty] | 2 | 2 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 3 | 3 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 4 | 4 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 5 | 5 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 6 | 6 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 7 | 7 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 8 | 8 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 9 | 9 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 10 | 10 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 11 | 11 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 12 | 12 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 13 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 5 | [6 empty] | 14 | 14 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 15 | 15 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 16 | 16 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 17 | 17 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 18 | 18 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 19 | 19 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 20 | 20 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 3 | 21 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 2 | 22 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 3 | 23 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 4 | 24 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 25 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 6 | [6 empty] | 6 | 26 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 7 | 27 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 8 | 28 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 9 | 29 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 10 | 30 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 11 | 31 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 12 | 32 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 13 | 33 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 14 | 34 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 15 | 35 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 16 | 36 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 37 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 7 | [6 empty] | 18 | 38 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 19 | 39 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 20 | 40 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 4 | 41 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 2 | 42 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 3 | 43 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 4 | 44 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 5 | 45 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 6 | 46 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 7 | 47 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 8 | 48 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 49 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 8 | [6 empty] | 10 | 50 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 11 | 51 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 12 | 52 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 13 | 53 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 14 | 54 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 15 | 55 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 16 | 56 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 17 | 57 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 18 | 58 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 19 | 59 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 20 | 60 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 61 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 9 | [6 empty] | 2 | 62 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 3 | 63 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 4 | 64 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 5 | 65 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 6 | 66 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 7 | 67 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 8 | 68 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 9 | 69 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 10 | 70 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 11 | 71 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 12 | 72 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 73 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 10 | [6 empty] | 14 | 74 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 15 | 75 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 16 | 76 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 17 | 77 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 18 | 78 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 19 | 79 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 20 | 80 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 6 | 81 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 2 | 82 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 3 | 83 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 4 | 84 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 85 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 11 | [6 empty] | 6 | 86 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 7 | 87 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 8 | 88 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 9 | 89 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 10 | 90 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 11 | 91 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 12 | 92 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 13 | 93 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 14 | 94 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 15 | 95 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 16 | 96 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 97 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 12 | [6 empty] | 18 | 98 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 19 | 99 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 20 | 100 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 7 | 101 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 2 | 102 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 3 | 103 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 4 | 104 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 5 | 105 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 6 | 106 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 7 | 107 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 8 | 108 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 109 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 13 | [6 empty] | 10 | 110 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 11 | 111 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 12 | 112 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 13 | 113 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 14 | 114 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 15 | 115 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 16 | 116 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 17 | 117 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 18 | 118 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 19 | 119 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 20 | 120 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 8 | 121 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 14 | [6 empty] | 2 | 122 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 3 | 123 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 4 | 124 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 5 | 125 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 6 | 126 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 7 | 127 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 8 | 128 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 9 | 129 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 10 | 130 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 11 | 131 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 12 | 132 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 133 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 15 | [6 empty] | 14 | 134 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 15 | 135 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 16 | 136 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 17 | 137 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 18 | 138 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 19 | 139 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 20 | 140 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 9 | 141 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 2 | 142 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 3 | 143 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 4 | 144 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 145 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 16 | [6 empty] | 6 | 146 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 7 | 147 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 8 | 148 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 9 | 149 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 10 | 150 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 11 | 151 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 12 | 152 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 13 | 153 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 14 | 154 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 15 | 155 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 16 | 156 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 157 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 17 | [6 empty] | 18 | 158 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 19 | 159 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 20 | 160 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 10 | 161 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 2 | 162 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 3 | 163 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 4 | 164 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 5 | 165 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 6 | 166 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 7 | 167 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 8 | 168 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 169 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 18 | [6 empty] | 10 | 170 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 11 | 171 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 12 | 172 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 13 | 173 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 14 | 174 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 15 | 175 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 16 | 176 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 17 | 177 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 18 | 178 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 19 | 179 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 20 | 180 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 11 | 181 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 19 | [6 empty] | 2 | 182 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 3 | 183 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 4 | 184 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 5 | 185 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 6 | 186 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 7 | 187 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 8 | 188 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 9 | 189 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 10 | 190 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 11 | 191 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 12 | 192 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 193 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 20 | [6 empty] | 14 | 194 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Std | mvDC | Temp | Adjusted Temp (with c.f.) | Rounded |  |  | 15 | 195 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP1 |  | 32 | 31.2682 | 31.3 |  | 1 | 16 | 196 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | SP3 | 0 | 32 | 31.262 | 31.3 |  | 1 | 17 | 197 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS12 |  | 32 | 232.3225 | 232.3 |  | #N/A | 18 | 198 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | NTS10 |  | 32 | 32.1861 | 32.2 |  | #N/A | 19 | 199 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | TTS1 |  | 32 | 0 | 0 |  | 2 | 20 | 200 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST7 |  | 32.0403 | 32.0403 | 32 |  |  | 12 | 201 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | PRT | 0 | 0 | 0 | 0 |  |  | 2 | 202 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | STD USD | mvDC | TMP |  |  |  |  | 3 | 203 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  | 0 | #N/A |  |  |  |  | 4 | 204 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  |  |  | SECONDARY STANDARDS |  |  |  |  | 5 | 205 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | sensor(selectable between ST6&SP1 | mVDC | Temp | Temp (w any C.F. (if applicable) |  |  |  | 6 | 206 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 0 | 0 | 32 | 31.2682 |  |  | 1 | 7 | 207 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 0 | 0 | 32 | 31.2682 |  |  | 1 | 8 | 208 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | 0 | 0 | 32 | 31.2682 |  |  | 1 | 9 | 209 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Icebath | 31.778 | 5464.6598 | #VALUE! |  |  | 13 | 10 | 210 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST-7 | 0.5933 | 199.1475 | 198.074 |  |  | 1 | 11 | 211 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST-7 | 1.4742 | 399.2659 | 398.9778 |  |  | 2 | 12 | 212 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST-7 | 2.4642 | 599.826 | 599.8917 |  |  | 3 | 13 | 213 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST-7 | 3.5162 | 800.0861 | 800.0721 |  |  | 5 | 14 | 214 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST-7 | 4.611 | 1000.1046 | 999.9493 |  |  | 6 | 15 | 215 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | ST-7 | 5.7377 | 1198.961 | 1198.1283 |  |  | 6 | 16 | 216 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 217 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[5 empty] | 31.2682 |  |  |  | 18 | 218 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | this chooses corrected or |  |  | 31.2682 |  |  |  | 19 | 219 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | uncorrected temps based on |  |  | 31.2682 |  |  |  | 20 | 220 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | C336-345 |  |  | 5464.6598 |  |  |  | 13 | 221 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | (Which standard is being used). |  |  | 199.1475 |  |  |  | 2 | 222 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | This is for calculating results for |  |  | 399.2659 |  |  |  | 3 | 223 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | standards calibration. |  |  | 599.826 |  |  |  | 4 | 224 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | Either the SP1 or the ST7 may be |  |  | 800.0861 |  |  |  | 5 | 225 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
|  | used |  |  | 1000.1046 |  |  |  | 6 | 226 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[5 empty] | 1198.961 |  |  |  | 7 | 227 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 8 | 228 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 229 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 10 | 230 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 11 | 231 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 12 | 232 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 233 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 14 | 234 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 15 | 235 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 16 | 236 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 237 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 18 | 238 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 19 | 239 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 20 | 240 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 14 | 241 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 2 | 242 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 3 | 243 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 4 | 244 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 245 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 6 | 246 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 7 | 247 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 8 | 248 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 249 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 10 | 250 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 11 | 251 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 12 | 252 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 253 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 14 | 254 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 15 | 255 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 16 | 256 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 257 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 18 | 258 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 19 | 259 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 20 | 260 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 15 | 261 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 2 | 262 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 3 | 263 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 4 | 264 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 265 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 6 | 266 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 7 | 267 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 8 | 268 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 269 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 10 | 270 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 11 | 271 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 12 | 272 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 273 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 14 | 274 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 15 | 275 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 16 | 276 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 277 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 18 | 278 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 19 | 279 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 20 | 280 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 16 | 281 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 2 | 282 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 3 | 283 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 4 | 284 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 285 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 6 | 286 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 7 | 287 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 8 | 288 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 289 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 10 | 290 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 11 | 291 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 12 | 292 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 293 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 14 | 294 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 15 | 295 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 16 | 296 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 297 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 18 | 298 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 19 | 299 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 20 | 300 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 301 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 2 | 302 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 3 | 303 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 4 | 304 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 305 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 6 | 306 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 7 | 307 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 8 | 308 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 309 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 10 | 310 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 11 | 311 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 12 | 312 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 313 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 14 | 314 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 15 | 315 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 16 | 316 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 317 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 18 | 318 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 19 | 319 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 20 | 320 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 18 | 321 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 2 | 322 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 3 | 323 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 4 | 324 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 325 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 6 | 326 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 7 | 327 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 8 | 328 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 329 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 10 | 330 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 11 | 331 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 12 | 332 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 333 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 14 | 334 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 15 | 335 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 16 | 336 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 337 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 18 | 338 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 19 | 339 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 20 | 340 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 19 | 341 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 2 | 342 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 3 | 343 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 4 | 344 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 345 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 6 | 346 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 7 | 347 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 8 | 348 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 349 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 10 | 350 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 11 | 351 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 12 | 352 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 353 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 14 | 354 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 15 | 355 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 16 | 356 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 357 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 18 | 358 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 19 | 359 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 20 | 360 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 20 | 361 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 2 | 362 | J | 0 | 32 | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 3 | 363 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 4 | 364 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 5 | 365 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 6 | 366 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 7 | 367 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 8 | 368 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 9 | 369 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 10 | 370 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 11 | 371 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 12 | 372 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 13 | 373 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 14 | 374 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 15 | 375 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 16 | 376 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 17 | 377 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 18 | 378 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 19 | 379 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32
[9 empty] | 20 | 380 |  | 0 | #N/A | 0 | 32 | 32 | 32 | 32 | 32 | 32

## Update Standards
| MULTIMETERS |  |  |  |  | Index | Name | Delimeter | Id
Shorthand | Type | Serial Number | Calibration Date | Date Due | Traceability | 1 | SP | - | 3
3458 | H.P. 3458A | 2823A08926 | 2024-05-08 | 2025-05-08 | WO-00790034 | 2 | PRT2 | SN: | 1624273
1502 | FLUKE 1502 | B55852 | 2025-01-20 | 2026-01-20 | 56561-068751 | 3
ADT286 | Additel 286 | 6851022030037 | 2024-07-01 | 2025-07-01 | 56561-060578 | 4 | TTS | - | 1
[6 empty] | 5 | ST | - | 7
[6 empty] | 6 | PRT | SN: | 496142
[6 empty] | 7 | SP | - | 1
[6 empty] | 8 | SP | - | 3
THERMOCOUPLES
[8 empty] | Name |  | Calibration Date | Date Due | Traceability
[7 empty] | 14 | SP-3 |  | 2024-01-19 | 2025-01-19 | SP30124
[7 empty] | 15 | PRT2SN:1624273 |  | 2024-03-27 | 2025-03-27 | 56561-056584
[7 empty] | 16 |  |  | 2024-01-19 | 2025-01-19 | 56561-056584
[7 empty] | 17 | TTS-1 |  | 2024-01-19 | 2025-01-19 | G295632
[7 empty] | 18 | ST-7 |  | 2023-05-16 | 2028-05-16 | 46756
[7 empty] | 19 | PRT SN:496142 |  | 2023-12-15 | 2025-12-15 | G295632
[7 empty] | 20 | SP-1 |  | 2024-01-19 | 2025-01-19 | SP10124
[7 empty] | 21 | SP-3 |  | 2024-01-19 | 2025-01-19 | SP30124
[8 empty] | THE Set of Base metal Transfer Standards is grouped: |  |  |  | Set SST-1
Put the set SN instead of individual standard on certs?
[8 empty] | If so, Enter "Y" | N

## SP-1Calibration
| CAL DATE: | 2022-12-22 | DUE: | 2023-12-22 | ONLY ENTER A DUE DATE IF DIFFERENT FROM ST-7 DUE DATE
|  | mvDC | °F |  |  |  |  | 8/23/10 results
32 | icebath | 31.841 | 32 |  |  |  |  | 31.8
| SP-1 | 0.0001 | 32.7333 | 32.7333 |  |  |  | 200.3
400.1
200 | ST-7 | 0.5933 | 199.1475 |  | NOTE PRINT USING PAGE 42 |  |  | 600.1
| SP-1 | 0.5979 | 200.225 | 201.0775 |  |  |  | 800
1000
400 | ST-7 | 1.4742 | 399.2659 |  |  |  |  | 1199.8
| SP-1 | 1.4762 | 399.55 | 400.2841 |  |  |  | 1400.7
1600.6
600 | ST-7 | 2.4642 | 599.826 |  |  |  |  | 1800.7
| SP-1 | 2.4648 | 599.76 | 599.934 |  |  |  | 2001
2201.4
800 | ST-7 | 3.5162 | 800.0861 |  |  |  |  | 2201.7
| SP-1 | 3.5165 | 800.1 | 800.0139
[7 empty] | COMPARISON RDGS: |  | 1 | 2 | 3 | 4 | 5
1000 | ST-7 | 4.611 | 1000.1046 | [5 empty] | 4.5964 | 4.5952 | 4.5948
| SP-1 | 4.6113 | 1000.26 | 1000.1554 |  |  |  |  | 4.5933 | 4.592 | 4.5922
1200 | ST-7 | 5.7377 | 1198.961
| SP-1 | 5.74 | 1199.8 | 1200.839
[9 empty] | -0.3 | -0.3 | -0.8 | -0.1
1400 | ST-7 | 6.9019 | 1397.48 |  | 6.9268 | 1401.8
| SP-1 | 6.8989 | 1397.65 | 1400.17 | 6.9287 | 1402.6
1600 | ST-7 | 8.1222 | 1598.2896 |  | 8.111 | 1596.7
| SP-1 | 8.1146 | 1598.0857 | 1599.7961 | 8.1161 | 1598.3
1800 | ST-7 | 9.3802 | 1798.3045 |  | 9.3459 | 1793.15
| SP-1 | 9.37 | 1798 | 1799.6955 | 9.3492 | 1794.7
2000 | ST-7 | 10.6898 | 2000.0968 |  | 10.6331 | 1991.85 |  |  | 5.7215
| SP-1 | 10.6897 | 2001.8143 | 2001.7175 | 10.6344 | 1993.4 |  |  | 5.735
2200 | ST-7 | 12.006 | 2198.0952 |  | 11.9053 | 2183.471
| SP-1 | 12.003 | 2199.8571 | 2201.7619 | 11.9047 | 2185.117
2400 | ST-7 | NO | #VALUE!
| SP-1 | NO | #VALUE! | #VALUE!

## SP-3Calibration
| CAL DATE: | 2022-12-22 | DUE: | 2023-12-22 |  |  | mvDC | °F
[5 empty] | 32°F | Icebath | 31.778 | 32 | 32
[6 empty] | SP-2 | 0.0001 | 32.7333 | 32.7333
|  | mvDC | °F
-50°F | PRT (in °F) |  | 0 | -50 | 200°F | ST-7 | 0.5933 | 199.1475 | 200 | NOTE: PRINT CERT USING PAGE 43.
| TTS-1 |  | 32 | -18 |  | SP-2 | 0.5923 | 198.825 | 199.6775
32°F | PRT (in °F) |  | 0 | 32 | 400°F | ST-7 | 1.4742 | 399.2659 | 400
| TTS-1 |  | 32 | 64 |  | SP-2 | 1.4735 | 398.9 | 399.6341
120°F | PRT (in °F) |  | 0 | 120 | 600°F | ST-7 | 2.4642 | 599.826 | 600
| TTS-1 |  | 32 | 152 |  | SP-2 | 2.4638 | 599.56 | 599.734
32°F |  |  | 31.2682 | 32 | 800°F | ST-7 | 3.5162 | 800.0861 | 800
| NTS-11 |  | 32 | 32.7318 |  | SP-2 | 3.5154 | 799.9 | 799.8139
70°F |  |  | 31.2682 | 70 | 1000°F | ST-7 | 4.611 | 1000.1046 | 1000
| NTS-11 |  | 32 | 70.7318 |  | SP-2 | 4.6085 | 999.75 | 999.6454
200°F | 0 |  | 31.2682 | 200 | 1200°F | ST-7 | 5.7377 | 1198.961 | 1200
| NTS-11 |  | 32 | 200.7318 |  | SP-2 | 5.7339 | 1198.65 | 1199.689
Note on E3: NTS11 is on pg 42. NTS12 and 10 are pg 43 TTS1 is pg 44 tmp, RH, and standards/comments need to be filled out on certworksht.

## Type E
[7 empty] | Temp(F) | mV
mV | Temp(F) | [5 empty] | -454 | -9.835
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -453 | -9.834
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -452 | -9.833
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -451 | -9.832
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -450 | -9.83
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -449 | -9.829
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -448 | -9.827
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -447 | -9.825
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -446 | -9.823
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -445 | -9.821
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -444 | -9.819
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -443 | -9.817
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -442 | -9.814
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -441 | -9.812
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -440 | -9.809
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -439 | -9.806
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -438 | -9.803
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -437 | -9.8
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -436 | -9.797
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -435 | -9.793
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -434 | -9.79
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -433 | -9.786
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -432 | -9.782
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -431 | -9.779
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -430 | -9.775
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -429 | -9.771
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -428 | -9.766
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -427 | -9.762
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -426 | -9.758
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -425 | -9.753
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -424 | -9.749
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -423 | -9.744
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -422 | -9.739
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -421 | -9.734
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -420 | -9.729
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -419 | -9.724
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -418 | -9.718
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -417 | -9.713
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -416 | -9.707
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -415 | -9.702
7 | 230.0526 | 685 | 6.998 | 7.036 | 230 | 231 | -414 | -9.696
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -413 | -9.69
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -412 | -9.684
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -411 | -9.678
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -410 | -9.672
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -409 | -9.666
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -408 | -9.659
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -407 | -9.653
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -406 | -9.646
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -405 | -9.639
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -404 | -9.632
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -403 | -9.625
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -402 | -9.618
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -401 | -9.611
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -400 | -9.604
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -399 | -9.597
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -398 | -9.589
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -397 | -9.581
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -396 | -9.574
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -395 | -9.566
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -394 | -9.558
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -393 | -9.55
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -392 | -9.542
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -391 | -9.534
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -390 | -9.525
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -389 | -9.517
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -388 | -9.508
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -387 | -9.5
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -386 | -9.491
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -385 | -9.482
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -384 | -9.473
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -383 | -9.464
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -382 | -9.455
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -381 | -9.446
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -380 | -9.436
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -379 | -9.427
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -378 | -9.417
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -377 | -9.408
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -376 | -9.398
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -375 | -9.388
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -374 | -9.378
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -373 | -9.368
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -372 | -9.358
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -371 | -9.348
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -370 | -9.338
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -369 | -9.327
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -368 | -9.317
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -367 | -9.306
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -366 | -9.295
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -365 | -9.285
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -364 | -9.274
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -363 | -9.263
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -362 | -9.252
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -361 | -9.241
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -360 | -9.229
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -359 | -9.218
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -358 | -9.207
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -357 | -9.195
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -356 | -9.184
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -355 | -9.172
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -354 | -9.16
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -353 | -9.148
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -352 | -9.136
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -351 | -9.124
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -350 | -9.112
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -349 | -9.1
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -348 | -9.088
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -347 | -9.075
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -346 | -9.063
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -345 | -9.05
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -344 | -9.038
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -343 | -9.025
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -342 | -9.012
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -341 | -8.999
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -340 | -8.986
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -339 | -8.973
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -338 | -8.96
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -337 | -8.947
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -336 | -8.934
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -335 | -8.92
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -334 | -8.907
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -333 | -8.893
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -332 | -8.88
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -331 | -8.866
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -330 | -8.852
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -329 | -8.839
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -328 | -8.825
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -327 | -8.811
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -326 | -8.797
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -325 | -8.782
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -324 | -8.768
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -323 | -8.754
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -322 | -8.739
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -321 | -8.725
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -320 | -8.71
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -319 | -8.696
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -318 | -8.681
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -317 | -8.666
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -316 | -8.652
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -315 | -8.637
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -314 | -8.622
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -313 | -8.607
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -312 | -8.591
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -311 | -8.576
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -310 | -8.561
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -309 | -8.546
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -308 | -8.53
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -307 | -8.515
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -306 | -8.499
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -305 | -8.483
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -304 | -8.468
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -303 | -8.452
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -302 | -8.436
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -301 | -8.42
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -300 | -8.404
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -299 | -8.388
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -298 | -8.372
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -297 | -8.356
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -296 | -8.339
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -295 | -8.323
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -294 | -8.307
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -293 | -8.29
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -292 | -8.273
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -291 | -8.257
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -290 | -8.24
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -289 | -8.223
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -288 | -8.206
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -287 | -8.189
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -286 | -8.173
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -285 | -8.155
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -284 | -8.138
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -283 | -8.121
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -282 | -8.104
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -281 | -8.087
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -280 | -8.069
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -279 | -8.052
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -278 | -8.034
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -277 | -8.017
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -276 | -7.999
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -275 | -7.981
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -274 | -7.963
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -273 | -7.945
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -272 | -7.928
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -271 | -7.91
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -270 | -7.891
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -269 | -7.873
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -268 | -7.855
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -267 | -7.837
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -266 | -7.819
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -265 | -7.8
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -264 | -7.782
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -263 | -7.763
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -262 | -7.745
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -261 | -7.726
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -260 | -7.707
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -259 | -7.688
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -258 | -7.67
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -257 | -7.651
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -256 | -7.632
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -255 | -7.613
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -254 | -7.593
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -253 | -7.574
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -252 | -7.555
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -251 | -7.536
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -250 | -7.516
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -249 | -7.497
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -248 | -7.478
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -247 | -7.458
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -246 | -7.438
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -245 | -7.419
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -244 | -7.399
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -243 | -7.379
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -242 | -7.359
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -241 | -7.339
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -240 | -7.319
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -239 | -7.299
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -238 | -7.279
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -237 | -7.259
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -236 | -7.239
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -235 | -7.219
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -234 | -7.198
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -233 | -7.178
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -232 | -7.157
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -231 | -7.137
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -230 | -7.116
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -229 | -7.096
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -228 | -7.075
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -227 | -7.054
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -226 | -7.033
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -225 | -7.013
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -224 | -6.992
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -223 | -6.971
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -222 | -6.95
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -221 | -6.928
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -220 | -6.907
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -219 | -6.886
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -218 | -6.865
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -217 | -6.843
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -216 | -6.822
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -215 | -6.801
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -214 | -6.779
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -213 | -6.757
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -212 | -6.736
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -211 | -6.714
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -210 | -6.692
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -209 | -6.671
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -208 | -6.649
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -207 | -6.627
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -206 | -6.605
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -205 | -6.583
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -204 | -6.561
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -203 | -6.539
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -202 | -6.516
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -201 | -6.494
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -200 | -6.472
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -199 | -6.449
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -198 | -6.427
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -197 | -6.405
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -196 | -6.382
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -195 | -6.359
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -194 | -6.337
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -193 | -6.314
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -192 | -6.291
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -191 | -6.269
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -190 | -6.246
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -189 | -6.223
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -188 | -6.2
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -187 | -6.177
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -186 | -6.154
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -185 | -6.13
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -184 | -6.107
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -183 | -6.084
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -182 | -6.061
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -181 | -6.037
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -180 | -6.014
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -179 | -5.991
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -178 | -5.967
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -177 | -5.943
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -176 | -5.92
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -175 | -5.896
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -174 | -5.872
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -173 | -5.849
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -172 | -5.825
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -171 | -5.801
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -170 | -5.777
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -169 | -5.753
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -168 | -5.729
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -167 | -5.705
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -166 | -5.681
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -165 | -5.656
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -164 | -5.632
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -163 | -5.608
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -162 | -5.584
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -161 | -5.559
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -160 | -5.535
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -159 | -5.51
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -158 | -5.486
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -157 | -5.461
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -156 | -5.436
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -155 | -5.412
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -154 | -5.387
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -153 | -5.362
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -152 | -5.337
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -151 | -5.312
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -150 | -5.287
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -149 | -5.262
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -148 | -5.237
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -147 | -5.212
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -146 | -5.187
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -145 | -5.162
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -144 | -5.136
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -143 | -5.111
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -142 | -5.086
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -141 | -5.06
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -140 | -5.035
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -139 | -5.009
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -138 | -4.984
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -137 | -4.958
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -136 | -4.932
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -135 | -4.907
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -134 | -4.881
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -133 | -4.855
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -132 | -4.829
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -131 | -4.803
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -130 | -4.777
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -129 | -4.751
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -128 | -4.725
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -127 | -4.699
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -126 | -4.673
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -125 | -4.647
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -124 | -4.621
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -123 | -4.594
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -122 | -4.568
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -121 | -4.542
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -120 | -4.515
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -119 | -4.489
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -118 | -4.462
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -117 | -4.436
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -116 | -4.409
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -115 | -4.382
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -114 | -4.355
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -113 | -4.329
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -112 | -4.302
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -111 | -4.275
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -110 | -4.248
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -109 | -4.221
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -108 | -4.194
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -107 | -4.167
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -106 | -4.14
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -105 | -4.113
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -104 | -4.086
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -103 | -4.058
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -102 | -4.031
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -101 | -4.004
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -100 | -3.976
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -99 | -3.949
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -98 | -3.922
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -97 | -3.894
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -96 | -3.867
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -95 | -3.839
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -94 | -3.811
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -93 | -3.784
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -92 | -3.756
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -91 | -3.728
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -90 | -3.7
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -89 | -3.672
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -88 | -3.645
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -87 | -3.617
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -86 | -3.589
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -85 | -3.561
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -84 | -3.532
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -83 | -3.504
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -82 | -3.476
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -81 | -3.448
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -80 | -3.42
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -79 | -3.391
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -78 | -3.363
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -77 | -3.335
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -76 | -3.306
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -75 | -3.278
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -74 | -3.249
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -73 | -3.221
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -72 | -3.192
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -71 | -3.163
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -70 | -3.135
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -69 | -3.106
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -68 | -3.077
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -67 | -3.048
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -66 | -3.02
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -65 | -2.991
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -64 | -2.962
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -63 | -2.933
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -62 | -2.904
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -61 | -2.875
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -60 | -2.846
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -59 | -2.816
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -58 | -2.787
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -57 | -2.758
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -56 | -2.729
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -55 | -2.699
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -54 | -2.67
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -53 | -2.641
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -52 | -2.611
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -51 | -2.582
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -50 | -2.552
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -49 | -2.523
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -48 | -2.493
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -47 | -2.463
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -46 | -2.434
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -45 | -2.404
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -44 | -2.374
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -43 | -2.344
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -42 | -2.315
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -41 | -2.285
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -40 | -2.255
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -39 | -2.225
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -38 | -2.195
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -37 | -2.165
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -36 | -2.135
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -35 | -2.105
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -34 | -2.074
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -33 | -2.044
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -32 | -2.014
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -31 | -1.984
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -30 | -1.953
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -29 | -1.923
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -28 | -1.893
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -27 | -1.862
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -26 | -1.832
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -25 | -1.801
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -24 | -1.771
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -23 | -1.74
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -22 | -1.709
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -21 | -1.679
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -20 | -1.648
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -19 | -1.617
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -18 | -1.587
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -17 | -1.556
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -16 | -1.525
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -15 | -1.494
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -14 | -1.463
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -13 | -1.432
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -12 | -1.401
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -11 | -1.37
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -10 | -1.339
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -9 | -1.308
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -8 | -1.277
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -7 | -1.245
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -6 | -1.214
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -5 | -1.183
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -4 | -1.152
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -3 | -1.12
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -2 | -1.089
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | -1 | -1.057
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 0 | -1.026
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1 | -0.994
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 2 | -0.963
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 3 | -0.931
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 4 | -0.9
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 5 | -0.868
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 6 | -0.836
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 7 | -0.805
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 8 | -0.773
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 9 | -0.741
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 10 | -0.709
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 11 | -0.677
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 12 | -0.645
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 13 | -0.614
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 14 | -0.582
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 15 | -0.55
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 16 | -0.517
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 17 | -0.485
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 18 | -0.453
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 19 | -0.421
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 20 | -0.389
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 21 | -0.357
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 22 | -0.324
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 23 | -0.292
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 24 | -0.26
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 25 | -0.227
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 26 | -0.195
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 27 | -0.163
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 28 | -0.13
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 29 | -0.098
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 30 | -0.065
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 31 | -0.033
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 32 | 0
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 33 | 0.033
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 34 | 0.065
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 35 | 0.098
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 36 | 0.131
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 37 | 0.163
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 38 | 0.196
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 39 | 0.229
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 40 | 0.262
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 41 | 0.294
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 42 | 0.327
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 43 | 0.36
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 44 | 0.393
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 45 | 0.426
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 46 | 0.459
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 47 | 0.492
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 48 | 0.525
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 49 | 0.558
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 50 | 0.591
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 51 | 0.624
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 52 | 0.657
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 53 | 0.691
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 54 | 0.724
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 55 | 0.757
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 56 | 0.79
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 57 | 0.824
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 58 | 0.857
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 59 | 0.89
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 60 | 0.924
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 61 | 0.957
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 62 | 0.99
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 63 | 1.024
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 64 | 1.057
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 65 | 1.091
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 66 | 1.124
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 67 | 1.158
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 68 | 1.192
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 69 | 1.225
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 70 | 1.259
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 71 | 1.292
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 72 | 1.326
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 73 | 1.36
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 74 | 1.394
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 75 | 1.427
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 76 | 1.461
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 77 | 1.495
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 78 | 1.529
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 79 | 1.563
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 80 | 1.597
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 81 | 1.631
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 82 | 1.665
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 83 | 1.699
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 84 | 1.733
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 85 | 1.767
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 86 | 1.801
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 87 | 1.835
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 88 | 1.869
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 89 | 1.904
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 90 | 1.938
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 91 | 1.972
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 92 | 2.006
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 93 | 2.041
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 94 | 2.075
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 95 | 2.109
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 96 | 2.144
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 97 | 2.178
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 98 | 2.212
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 99 | 2.247
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 100 | 2.281
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 101 | 2.316
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 102 | 2.351
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 103 | 2.385
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 104 | 2.42
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 105 | 2.454
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 106 | 2.489
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 107 | 2.524
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 108 | 2.558
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 109 | 2.593
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 110 | 2.628
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 111 | 2.663
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 112 | 2.698
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 113 | 2.733
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 114 | 2.767
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 115 | 2.802
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 116 | 2.837
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 117 | 2.872
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 118 | 2.907
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 119 | 2.942
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 120 | 2.977
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 121 | 3.012
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 122 | 3.048
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 123 | 3.083
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 124 | 3.118
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 125 | 3.153
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 126 | 3.188
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 127 | 3.224
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 128 | 3.259
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 129 | 3.294
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 130 | 3.33
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 131 | 3.365
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 132 | 3.4
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 133 | 3.436
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 134 | 3.471
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 135 | 3.507
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 136 | 3.542
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 137 | 3.578
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 138 | 3.613
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 139 | 3.649
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 140 | 3.685
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 141 | 3.72
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 142 | 3.756
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 143 | 3.792
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 144 | 3.827
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 145 | 3.863
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 146 | 3.899
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 147 | 3.935
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 148 | 3.97
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 149 | 4.006
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 150 | 4.042
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 151 | 4.078
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 152 | 4.114
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 153 | 4.15
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 154 | 4.186
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 155 | 4.222
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 156 | 4.258
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 157 | 4.294
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 158 | 4.33
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 159 | 4.366
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 160 | 4.403
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 161 | 4.439
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 162 | 4.475
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 163 | 4.511
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 164 | 4.547
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 165 | 4.584
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 166 | 4.62
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 167 | 4.656
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 168 | 4.693
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 169 | 4.729
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 170 | 4.766
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 171 | 4.802
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 172 | 4.839
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 173 | 4.875
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 174 | 4.912
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 175 | 4.948
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 176 | 4.985
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 177 | 5.021
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 178 | 5.058
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 179 | 5.095
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 180 | 5.131
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 181 | 5.168
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 182 | 5.205
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 183 | 5.242
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 184 | 5.278
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 185 | 5.315
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 186 | 5.352
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 187 | 5.389
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 188 | 5.426
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 189 | 5.463
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 190 | 5.5
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 191 | 5.537
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 192 | 5.574
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 193 | 5.611
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 194 | 5.648
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 195 | 5.685
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 196 | 5.722
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 197 | 5.759
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 198 | 5.796
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 199 | 5.833
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 200 | 5.871
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 201 | 5.908
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 202 | 5.945
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 203 | 5.982
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 204 | 6.02
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 205 | 6.057
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 206 | 6.094
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 207 | 6.132
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 208 | 6.169
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 209 | 6.207
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 210 | 6.244
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 211 | 6.281
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 212 | 6.319
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 213 | 6.356
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 214 | 6.394
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 215 | 6.432
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 216 | 6.469
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 217 | 6.507
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 218 | 6.544
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 219 | 6.582
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 220 | 6.62
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 221 | 6.658
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 222 | 6.695
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 223 | 6.733
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 224 | 6.771
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 225 | 6.809
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 226 | 6.847
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 227 | 6.884
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 228 | 6.922
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 229 | 6.96
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 230 | 6.998
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 231 | 7.036
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 232 | 7.074
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 233 | 7.112
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 234 | 7.15
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 235 | 7.188
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 236 | 7.226
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 237 | 7.264
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 238 | 7.302
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 239 | 7.341
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 240 | 7.379
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 241 | 7.417
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 242 | 7.455
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 243 | 7.493
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 244 | 7.532
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 245 | 7.57
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 246 | 7.608
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 247 | 7.647
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 248 | 7.685
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 249 | 7.723
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 250 | 7.762
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 251 | 7.8
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 252 | 7.839
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 253 | 7.877
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 254 | 7.916
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 255 | 7.954
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 256 | 7.993
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 257 | 8.031
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 258 | 8.07
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 259 | 8.108
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 260 | 8.147
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 261 | 8.186
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 262 | 8.224
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 263 | 8.263
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 264 | 8.302
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 265 | 8.34
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 266 | 8.379
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 267 | 8.418
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 268 | 8.457
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 269 | 8.496
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 270 | 8.535
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 271 | 8.573
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 272 | 8.612
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 273 | 8.651
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 274 | 8.69
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 275 | 8.729
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 276 | 8.768
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 277 | 8.807
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 278 | 8.846
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 279 | 8.885
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 280 | 8.924
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 281 | 8.963
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 282 | 9.002
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 283 | 9.041
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 284 | 9.081
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 285 | 9.12
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 286 | 9.159
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 287 | 9.198
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 288 | 9.237
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 289 | 9.277
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 290 | 9.316
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 291 | 9.355
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 292 | 9.395
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 293 | 9.434
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 294 | 9.473
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 295 | 9.513
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 296 | 9.552
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 297 | 9.591
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 298 | 9.631
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 299 | 9.67
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 300 | 9.71
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 301 | 9.749
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 302 | 9.789
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 303 | 9.828
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 304 | 9.868
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 305 | 9.907
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 306 | 9.947
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 307 | 9.987
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 308 | 10.026
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 309 | 10.066
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 310 | 10.106
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 311 | 10.145
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 312 | 10.185
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 313 | 10.225
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 314 | 10.265
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 315 | 10.304
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 316 | 10.344
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 317 | 10.384
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 318 | 10.424
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 319 | 10.464
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 320 | 10.503
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 321 | 10.543
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 322 | 10.583
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 323 | 10.623
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 324 | 10.663
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 325 | 10.703
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 326 | 10.743
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 327 | 10.783
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 328 | 10.823
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 329 | 10.863
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 330 | 10.903
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 331 | 10.943
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 332 | 10.983
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 333 | 11.024
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 334 | 11.064
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 335 | 11.104
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 336 | 11.144
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 337 | 11.184
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 338 | 11.224
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 339 | 11.265
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 340 | 11.305
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 341 | 11.345
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 342 | 11.385
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 343 | 11.426
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 344 | 11.466
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 345 | 11.506
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 346 | 11.547
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 347 | 11.587
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 348 | 11.627
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 349 | 11.668
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 350 | 11.708
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 351 | 11.749
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 352 | 11.789
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 353 | 11.83
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 354 | 11.87
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 355 | 11.911
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 356 | 11.951
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 357 | 11.992
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 358 | 12.032
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 359 | 12.073
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 360 | 12.113
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 361 | 12.154
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 362 | 12.195
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 363 | 12.235
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 364 | 12.276
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 365 | 12.317
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 366 | 12.357
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 367 | 12.398
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 368 | 12.439
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 369 | 12.48
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 370 | 12.52
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 371 | 12.561
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 372 | 12.602
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 373 | 12.643
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 374 | 12.684
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 375 | 12.724
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 376 | 12.765
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 377 | 12.806
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 378 | 12.847
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 379 | 12.888
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 380 | 12.929
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 381 | 12.97
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 382 | 13.011
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 383 | 13.052
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 384 | 13.093
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 385 | 13.134
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 386 | 13.175
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 387 | 13.216
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 388 | 13.257
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 389 | 13.298
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 390 | 13.339
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 391 | 13.38
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 392 | 13.421
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 393 | 13.462
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 394 | 13.504
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 395 | 13.545
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 396 | 13.586
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 397 | 13.627
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 398 | 13.668
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 399 | 13.71
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 400 | 13.751
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 401 | 13.792
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 402 | 13.833
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 403 | 13.875
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 404 | 13.916
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 405 | 13.957
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 406 | 13.999
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 407 | 14.04
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 408 | 14.081
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 409 | 14.123
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 410 | 14.164
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 411 | 14.205
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 412 | 14.247
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 413 | 14.288
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 414 | 14.33
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 415 | 14.371
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 416 | 14.413
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 417 | 14.454
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 418 | 14.496
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 419 | 14.537
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 420 | 14.579
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 421 | 14.62
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 422 | 14.662
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 423 | 14.704
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 424 | 14.745
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 425 | 14.787
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 426 | 14.828
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 427 | 14.87
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 428 | 14.912
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 429 | 14.953
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 430 | 14.995
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 431 | 15.037
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 432 | 15.078
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 433 | 15.12
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 434 | 15.162
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 435 | 15.204
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 436 | 15.245
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 437 | 15.287
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 438 | 15.329
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 439 | 15.371
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 440 | 15.413
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 441 | 15.454
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 442 | 15.496
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 443 | 15.538
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 444 | 15.58
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 445 | 15.622
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 446 | 15.664
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 447 | 15.706
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 448 | 15.748
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 449 | 15.79
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 450 | 15.831
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 451 | 15.873
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 452 | 15.915
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 453 | 15.957
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 454 | 15.999
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 455 | 16.041
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 456 | 16.083
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 457 | 16.125
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 458 | 16.168
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 459 | 16.21
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 460 | 16.252
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 461 | 16.294
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 462 | 16.336
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 463 | 16.378
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 464 | 16.42
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 465 | 16.462
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 466 | 16.504
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 467 | 16.547
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 468 | 16.589
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 469 | 16.631
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 470 | 16.673
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 471 | 16.715
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 472 | 16.758
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 473 | 16.8
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 474 | 16.842
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 475 | 16.884
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 476 | 16.927
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 477 | 16.969
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 478 | 17.011
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 479 | 17.054
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 480 | 17.096
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 481 | 17.138
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 482 | 17.181
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 483 | 17.223
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 484 | 17.265
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 485 | 17.308
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 486 | 17.35
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 487 | 17.392
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 488 | 17.435
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 489 | 17.477
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 490 | 17.52
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 491 | 17.562
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 492 | 17.605
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 493 | 17.647
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 494 | 17.69
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 495 | 17.732
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 496 | 17.775
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 497 | 17.817
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 498 | 17.86
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 499 | 17.902
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 500 | 17.945
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 501 | 17.987
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 502 | 18.03
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 503 | 18.073
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 504 | 18.115
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 505 | 18.158
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 506 | 18.2
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 507 | 18.243
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 508 | 18.286
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 509 | 18.328
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 510 | 18.371
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 511 | 18.414
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 512 | 18.456
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 513 | 18.499
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 514 | 18.542
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 515 | 18.585
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 516 | 18.627
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 517 | 18.67
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 518 | 18.713
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 519 | 18.756
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 520 | 18.798
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 521 | 18.841
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 522 | 18.884
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 523 | 18.927
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 524 | 18.969
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 525 | 19.012
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 526 | 19.055
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 527 | 19.098
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 528 | 19.141
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 529 | 19.184
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 530 | 19.227
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 531 | 19.269
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 532 | 19.312
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 533 | 19.355
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 534 | 19.398
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 535 | 19.441
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 536 | 19.484
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 537 | 19.527
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 538 | 19.57
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 539 | 19.613
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 540 | 19.656
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 541 | 19.699
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 542 | 19.742
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 543 | 19.785
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 544 | 19.828
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 545 | 19.871
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 546 | 19.914
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 547 | 19.957
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 548 | 20
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 549 | 20.043
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 550 | 20.086
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 551 | 20.129
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 552 | 20.172
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 553 | 20.216
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 554 | 20.259
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 555 | 20.302
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 556 | 20.345
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 557 | 20.388
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 558 | 20.431
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 559 | 20.474
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 560 | 20.517
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 561 | 20.561
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 562 | 20.604
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 563 | 20.647
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 564 | 20.69
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 565 | 20.733
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 566 | 20.777
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 567 | 20.82
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 568 | 20.863
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 569 | 20.906
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 570 | 20.95
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 571 | 20.993
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 572 | 21.036
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 573 | 21.08
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 574 | 21.123
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 575 | 21.166
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 576 | 21.209
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 577 | 21.253
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 578 | 21.296
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 579 | 21.339
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 580 | 21.383
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 581 | 21.426
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 582 | 21.47
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 583 | 21.513
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 584 | 21.556
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 585 | 21.6
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 586 | 21.643
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 587 | 21.686
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 588 | 21.73
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 589 | 21.773
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 590 | 21.817
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 591 | 21.86
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 592 | 21.904
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 593 | 21.947
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 594 | 21.991
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 595 | 22.034
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 596 | 22.078
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 597 | 22.121
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 598 | 22.165
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 599 | 22.208
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 600 | 22.252
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 601 | 22.295
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 602 | 22.339
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 603 | 22.382
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 604 | 22.426
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 605 | 22.469
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 606 | 22.513
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 607 | 22.556
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 608 | 22.6
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 609 | 22.644
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 610 | 22.687
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 611 | 22.731
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 612 | 22.774
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 613 | 22.818
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 614 | 22.862
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 615 | 22.905
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 616 | 22.949
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 617 | 22.993
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 618 | 23.036
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 619 | 23.08
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 620 | 23.124
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 621 | 23.167
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 622 | 23.211
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 623 | 23.255
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 624 | 23.298
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 625 | 23.342
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 626 | 23.386
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 627 | 23.429
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 628 | 23.473
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 629 | 23.517
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 630 | 23.561
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 631 | 23.604
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 632 | 23.648
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 633 | 23.692
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 634 | 23.736
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 635 | 23.78
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 636 | 23.823
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 637 | 23.867
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 638 | 23.911
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 639 | 23.955
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 640 | 23.999
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 641 | 24.042
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 642 | 24.086
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 643 | 24.13
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 644 | 24.174
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 645 | 24.218
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 646 | 24.262
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 647 | 24.305
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 648 | 24.349
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 649 | 24.393
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 650 | 24.437
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 651 | 24.481
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 652 | 24.525
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 653 | 24.569
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 654 | 24.613
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 655 | 24.657
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 656 | 24.701
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 657 | 24.745
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 658 | 24.789
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 659 | 24.832
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 660 | 24.876
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 661 | 24.92
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 662 | 24.964
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 663 | 25.008
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 664 | 25.052
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 665 | 25.096
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 666 | 25.14
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 667 | 25.184
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 668 | 25.228
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 669 | 25.272
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 670 | 25.316
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 671 | 25.36
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 672 | 25.404
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 673 | 25.448
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 674 | 25.493
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 675 | 25.537
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 676 | 25.581
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 677 | 25.625
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 678 | 25.669
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 679 | 25.713
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 680 | 25.757
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 681 | 25.801
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 682 | 25.845
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 683 | 25.889
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 684 | 25.933
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 685 | 25.977
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 686 | 26.022
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 687 | 26.066
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 688 | 26.11
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 689 | 26.154
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 690 | 26.198
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 691 | 26.242
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 692 | 26.286
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 693 | 26.331
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 694 | 26.375
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 695 | 26.419
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 696 | 26.463
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 697 | 26.507
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 698 | 26.552
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 699 | 26.596
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 700 | 26.64
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 701 | 26.684
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 702 | 26.728
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 703 | 26.773
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 704 | 26.817
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 705 | 26.861
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 706 | 26.905
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 707 | 26.95
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 708 | 26.994
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 709 | 27.038
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 710 | 27.082
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 711 | 27.127
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 712 | 27.171
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 713 | 27.215
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 714 | 27.259
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 715 | 27.304
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 716 | 27.348
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 717 | 27.392
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 718 | 27.437
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 719 | 27.481
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 720 | 27.525
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 721 | 27.57
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 722 | 27.614
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 723 | 27.658
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 724 | 27.703
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 725 | 27.747
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 726 | 27.791
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 727 | 27.836
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 728 | 27.88
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 729 | 27.924
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 730 | 27.969
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 731 | 28.013
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 732 | 28.057
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 733 | 28.102
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 734 | 28.146
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 735 | 28.191
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 736 | 28.235
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 737 | 28.279
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 738 | 28.324
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 739 | 28.368
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 740 | 28.413
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 741 | 28.457
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 742 | 28.501
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 743 | 28.546
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 744 | 28.59
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 745 | 28.635
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 746 | 28.679
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 747 | 28.724
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 748 | 28.768
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 749 | 28.813
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 750 | 28.857
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 751 | 28.901
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 752 | 28.946
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 753 | 28.99
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 754 | 29.035
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 755 | 29.079
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 756 | 29.124
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 757 | 29.168
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 758 | 29.213
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 759 | 29.257
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 760 | 29.302
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 761 | 29.346
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 762 | 29.391
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 763 | 29.435
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 764 | 29.48
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 765 | 29.525
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 766 | 29.569
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 767 | 29.614
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 768 | 29.658
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 769 | 29.703
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 770 | 29.747
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 771 | 29.792
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 772 | 29.836
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 773 | 29.881
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 774 | 29.925
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 775 | 29.97
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 776 | 30.015
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 777 | 30.059
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 778 | 30.104
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 779 | 30.148
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 780 | 30.193
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 781 | 30.238
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 782 | 30.282
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 783 | 30.327
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 784 | 30.371
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 785 | 30.416
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 786 | 30.461
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 787 | 30.505
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 788 | 30.55
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 789 | 30.595
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 790 | 30.639
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 791 | 30.684
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 792 | 30.728
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 793 | 30.773
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 794 | 30.818
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 795 | 30.862
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 796 | 30.907
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 797 | 30.952
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 798 | 30.996
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 799 | 31.041
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 800 | 31.086
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 801 | 31.13
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 802 | 31.175
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 803 | 31.22
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 804 | 31.264
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 805 | 31.309
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 806 | 31.354
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 807 | 31.398
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 808 | 31.443
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 809 | 31.488
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 810 | 31.533
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 811 | 31.577
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 812 | 31.622
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 813 | 31.667
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 814 | 31.711
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 815 | 31.756
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 816 | 31.801
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 817 | 31.846
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 818 | 31.89
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 819 | 31.935
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 820 | 31.98
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 821 | 32.025
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 822 | 32.069
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 823 | 32.114
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 824 | 32.159
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 825 | 32.204
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 826 | 32.248
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 827 | 32.293
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 828 | 32.338
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 829 | 32.383
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 830 | 32.427
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 831 | 32.472
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 832 | 32.517
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 833 | 32.562
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 834 | 32.606
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 835 | 32.651
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 836 | 32.696
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 837 | 32.741
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 838 | 32.786
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 839 | 32.83
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 840 | 32.875
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 841 | 32.92
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 842 | 32.965
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 843 | 33.01
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 844 | 33.054
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 845 | 33.099
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 846 | 33.144
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 847 | 33.189
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 848 | 33.234
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 849 | 33.278
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 850 | 33.323
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 851 | 33.368
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 852 | 33.413
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 853 | 33.458
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 854 | 33.503
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 855 | 33.547
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 856 | 33.592
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 857 | 33.637
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 858 | 33.682
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 859 | 33.727
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 860 | 33.772
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 861 | 33.816
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 862 | 33.861
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 863 | 33.906
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 864 | 33.951
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 865 | 33.996
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 866 | 34.041
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 867 | 34.086
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 868 | 34.13
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 869 | 34.175
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 870 | 34.22
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 871 | 34.265
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 872 | 34.31
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 873 | 34.355
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 874 | 34.4
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 875 | 34.445
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 876 | 34.489
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 877 | 34.534
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 878 | 34.579
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 879 | 34.624
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 880 | 34.669
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 881 | 34.714
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 882 | 34.759
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 883 | 34.804
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 884 | 34.849
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 885 | 34.893
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 886 | 34.938
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 887 | 34.983
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 888 | 35.028
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 889 | 35.073
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 890 | 35.118
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 891 | 35.163
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 892 | 35.208
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 893 | 35.253
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 894 | 35.298
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 895 | 35.343
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 896 | 35.387
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 897 | 35.432
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 898 | 35.477
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 899 | 35.522
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 900 | 35.567
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 901 | 35.612
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 902 | 35.657
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 903 | 35.702
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 904 | 35.747
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 905 | 35.792
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 906 | 35.837
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 907 | 35.882
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 908 | 35.927
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 909 | 35.972
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 910 | 36.016
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 911 | 36.061
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 912 | 36.106
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 913 | 36.151
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 914 | 36.196
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 915 | 36.241
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 916 | 36.286
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 917 | 36.331
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 918 | 36.376
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 919 | 36.421
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 920 | 36.466
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 921 | 36.511
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 922 | 36.556
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 923 | 36.601
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 924 | 36.646
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 925 | 36.691
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 926 | 36.736
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 927 | 36.781
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 928 | 36.826
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 929 | 36.87
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 930 | 36.915
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 931 | 36.96
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 932 | 37.005
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 933 | 37.05
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 934 | 37.095
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 935 | 37.14
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 936 | 37.185
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 937 | 37.23
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 938 | 37.275
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 939 | 37.32
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 940 | 37.365
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 941 | 37.41
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 942 | 37.455
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 943 | 37.5
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 944 | 37.545
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 945 | 37.59
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 946 | 37.635
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 947 | 37.68
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 948 | 37.725
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 949 | 37.77
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 950 | 37.815
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 951 | 37.86
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 952 | 37.905
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 953 | 37.95
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 954 | 37.995
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 955 | 38.04
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 956 | 38.085
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 957 | 38.13
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 958 | 38.175
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 959 | 38.22
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 960 | 38.265
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 961 | 38.309
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 962 | 38.354
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 963 | 38.399
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 964 | 38.444
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 965 | 38.489
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 966 | 38.534
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 967 | 38.579
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 968 | 38.624
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 969 | 38.669
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 970 | 38.714
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 971 | 38.759
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 972 | 38.804
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 973 | 38.849
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 974 | 38.894
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 975 | 38.939
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 976 | 38.984
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 977 | 39.029
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 978 | 39.074
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 979 | 39.119
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 980 | 39.164
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 981 | 39.209
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 982 | 39.254
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 983 | 39.299
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 984 | 39.344
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 985 | 39.389
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 986 | 39.434
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 987 | 39.479
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 988 | 39.524
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 989 | 39.569
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 990 | 39.614
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 991 | 39.659
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 992 | 39.704
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 993 | 39.749
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 994 | 39.794
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 995 | 39.839
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 996 | 39.884
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 997 | 39.929
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 998 | 39.974
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 999 | 40.019
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1000 | 40.064
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1001 | 40.109
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1002 | 40.154
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1003 | 40.199
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1004 | 40.243
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1005 | 40.288
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1006 | 40.333
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1007 | 40.378
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1008 | 40.423
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1009 | 40.468
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1010 | 40.513
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1011 | 40.558
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1012 | 40.603
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1013 | 40.648
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1014 | 40.693
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1015 | 40.738
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1016 | 40.783
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1017 | 40.828
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1018 | 40.873
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1019 | 40.918
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1020 | 40.963
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1021 | 41.008
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1022 | 41.053
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1023 | 41.098
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1024 | 41.143
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1025 | 41.188
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1026 | 41.233
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1027 | 41.278
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1028 | 41.323
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1029 | 41.368
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1030 | 41.412
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1031 | 41.457
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1032 | 41.502
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1033 | 41.547
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1034 | 41.592
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1035 | 41.637
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1036 | 41.682
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1037 | 41.727
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1038 | 41.772
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1039 | 41.817
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1040 | 41.862
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1041 | 41.907
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1042 | 41.952
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1043 | 41.997
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1044 | 42.042
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1045 | 42.087
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1046 | 42.132
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1047 | 42.176
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1048 | 42.221
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1049 | 42.266
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1050 | 42.311
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1051 | 42.356
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1052 | 42.401
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1053 | 42.446
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1054 | 42.491
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1055 | 42.536
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1056 | 42.581
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1057 | 42.626
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1058 | 42.671
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1059 | 42.715
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1060 | 42.76
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1061 | 42.805
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1062 | 42.85
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1063 | 42.895
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1064 | 42.94
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1065 | 42.985
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1066 | 43.03
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1067 | 43.075
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1068 | 43.12
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1069 | 43.165
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1070 | 43.209
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1071 | 43.254
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1072 | 43.299
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1073 | 43.344
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1074 | 43.389
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1075 | 43.434
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1076 | 43.479
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1077 | 43.524
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1078 | 43.569
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1079 | 43.613
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1080 | 43.658
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1081 | 43.703
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1082 | 43.748
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1083 | 43.793
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1084 | 43.838
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1085 | 43.883
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1086 | 43.928
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1087 | 43.972
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1088 | 44.017
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1089 | 44.062
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1090 | 44.107
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1091 | 44.152
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1092 | 44.197
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1093 | 44.242
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1094 | 44.286
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1095 | 44.331
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1096 | 44.376
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1097 | 44.421
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1098 | 44.466
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1099 | 44.511
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1100 | 44.555
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1101 | 44.6
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1102 | 44.645
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1103 | 44.69
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1104 | 44.735
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1105 | 44.78
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1106 | 44.824
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1107 | 44.869
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1108 | 44.914
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1109 | 44.959
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1110 | 45.004
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1111 | 45.049
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1112 | 45.093
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1113 | 45.138
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1114 | 45.183
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1115 | 45.228
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1116 | 45.273
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1117 | 45.317
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1118 | 45.362
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1119 | 45.407
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1120 | 45.452
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1121 | 45.497
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1122 | 45.541
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1123 | 45.586
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1124 | 45.631
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1125 | 45.676
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1126 | 45.72
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1127 | 45.765
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1128 | 45.81
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1129 | 45.855
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1130 | 45.9
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1131 | 45.944
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1132 | 45.989
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1133 | 46.034
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1134 | 46.079
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1135 | 46.123
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1136 | 46.168
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1137 | 46.213
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1138 | 46.258
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1139 | 46.302
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1140 | 46.347
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1141 | 46.392
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1142 | 46.437
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1143 | 46.481
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1144 | 46.526
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1145 | 46.571
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1146 | 46.616
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1147 | 46.66
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1148 | 46.705
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1149 | 46.75
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1150 | 46.794
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1151 | 46.839
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1152 | 46.884
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1153 | 46.929
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1154 | 46.973
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1155 | 47.018
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1156 | 47.063
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1157 | 47.107
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1158 | 47.152
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1159 | 47.197
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1160 | 47.241
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1161 | 47.286
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1162 | 47.331
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1163 | 47.375
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1164 | 47.42
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1165 | 47.465
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1166 | 47.509
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1167 | 47.554
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1168 | 47.599
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1169 | 47.643
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1170 | 47.688
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1171 | 47.733
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1172 | 47.777
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1173 | 47.822
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1174 | 47.867
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1175 | 47.911
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1176 | 47.956
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1177 | 48.001
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1178 | 48.045
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1179 | 48.09
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1180 | 48.135
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1181 | 48.179
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1182 | 48.224
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1183 | 48.268
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1184 | 48.313
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1185 | 48.358
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1186 | 48.402
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1187 | 48.447
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1188 | 48.492
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1189 | 48.536
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1190 | 48.581
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1191 | 48.625
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1192 | 48.67
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1193 | 48.715
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1194 | 48.759
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1195 | 48.804
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1196 | 48.848
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1197 | 48.893
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1198 | 48.937
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1199 | 48.982
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1200 | 49.027
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1201 | 49.071
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1202 | 49.116
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1203 | 49.16
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1204 | 49.205
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1205 | 49.249
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1206 | 49.294
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1207 | 49.338
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1208 | 49.383
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1209 | 49.428
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1210 | 49.472
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1211 | 49.517
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1212 | 49.561
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1213 | 49.606
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1214 | 49.65
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1215 | 49.695
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1216 | 49.739
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1217 | 49.784
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1218 | 49.828
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1219 | 49.873
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1220 | 49.917
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1221 | 49.962
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1222 | 50.006
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1223 | 50.051
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1224 | 50.095
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1225 | 50.14
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1226 | 50.184
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1227 | 50.229
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1228 | 50.273
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1229 | 50.318
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1230 | 50.362
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1231 | 50.407
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1232 | 50.451
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1233 | 50.495
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1234 | 50.54
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1235 | 50.584
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1236 | 50.629
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1237 | 50.673
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1238 | 50.718
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1239 | 50.762
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1240 | 50.807
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1241 | 50.851
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1242 | 50.895
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1243 | 50.94
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1244 | 50.984
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1245 | 51.029
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1246 | 51.073
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1247 | 51.118
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1248 | 51.162
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1249 | 51.206
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1250 | 51.251
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1251 | 51.295
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1252 | 51.34
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1253 | 51.384
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1254 | 51.428
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1255 | 51.473
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1256 | 51.517
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1257 | 51.561
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1258 | 51.606
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1259 | 51.65
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1260 | 51.695
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1261 | 51.739
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1262 | 51.783
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1263 | 51.828
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1264 | 51.872
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1265 | 51.916
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1266 | 51.961
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1267 | 52.005
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1268 | 52.049
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1269 | 52.094
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1270 | 52.138
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1271 | 52.182
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1272 | 52.227
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1273 | 52.271
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1274 | 52.315
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1275 | 52.36
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1276 | 52.404
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1277 | 52.448
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1278 | 52.493
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1279 | 52.537
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1280 | 52.581
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1281 | 52.625
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1282 | 52.67
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1283 | 52.714
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1284 | 52.758
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1285 | 52.803
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1286 | 52.847
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1287 | 52.891
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1288 | 52.935
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1289 | 52.98
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1290 | 53.024
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1291 | 53.068
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1292 | 53.112
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1293 | 53.157
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1294 | 53.201
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1295 | 53.245
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1296 | 53.289
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1297 | 53.334
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1298 | 53.378
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1299 | 53.422
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1300 | 53.466
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1301 | 53.51
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1302 | 53.555
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1303 | 53.599
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1304 | 53.643
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1305 | 53.687
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1306 | 53.732
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1307 | 53.776
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1308 | 53.82
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1309 | 53.864
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1310 | 53.908
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1311 | 53.952
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1312 | 53.997
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1313 | 54.041
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1314 | 54.085
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1315 | 54.129
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1316 | 54.173
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1317 | 54.218
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1318 | 54.262
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1319 | 54.306
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1320 | 54.35
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1321 | 54.394
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1322 | 54.438
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1323 | 54.482
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1324 | 54.527
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1325 | 54.571
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1326 | 54.615
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1327 | 54.659
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1328 | 54.703
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1329 | 54.747
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1330 | 54.791
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1331 | 54.835
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1332 | 54.879
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1333 | 54.924
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1334 | 54.968
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1335 | 55.012
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1336 | 55.056
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1337 | 55.1
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1338 | 55.144
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1339 | 55.188
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1340 | 55.232
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1341 | 55.276
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1342 | 55.32
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1343 | 55.364
0 | 32 | 486 | -0.033 | 0 | 31 | 32 | 1344 | 55.408
[7 empty] | 1345 | 55.453
[7 empty] | 1346 | 55.497
[7 empty] | 1347 | 55.541
[7 empty] | 1348 | 55.585
[7 empty] | 1349 | 55.629
[7 empty] | 1350 | 55.673
[7 empty] | 1351 | 55.717
[7 empty] | 1352 | 55.761
[7 empty] | 1353 | 55.805
[7 empty] | 1354 | 55.849
[7 empty] | 1355 | 55.893
[7 empty] | 1356 | 55.937
[7 empty] | 1357 | 55.981
[7 empty] | 1358 | 56.025
[7 empty] | 1359 | 56.069
[7 empty] | 1360 | 56.113
[7 empty] | 1361 | 56.157
[7 empty] | 1362 | 56.201
[7 empty] | 1363 | 56.245
[7 empty] | 1364 | 56.289
[7 empty] | 1365 | 56.333
[7 empty] | 1366 | 56.377
[7 empty] | 1367 | 56.421
[7 empty] | 1368 | 56.465
[7 empty] | 1369 | 56.509
[7 empty] | 1370 | 56.553
[7 empty] | 1371 | 56.597
[7 empty] | 1372 | 56.641
[7 empty] | 1373 | 56.685
[7 empty] | 1374 | 56.729
[7 empty] | 1375 | 56.773
[7 empty] | 1376 | 56.816
[7 empty] | 1377 | 56.86
[7 empty] | 1378 | 56.904
[7 empty] | 1379 | 56.948
[7 empty] | 1380 | 56.992
[7 empty] | 1381 | 57.036
[7 empty] | 1382 | 57.08
[7 empty] | 1383 | 57.124
[7 empty] | 1384 | 57.168
[7 empty] | 1385 | 57.212
[7 empty] | 1386 | 57.256
[7 empty] | 1387 | 57.3
[7 empty] | 1388 | 57.344
[7 empty] | 1389 | 57.387
[7 empty] | 1390 | 57.431
[7 empty] | 1391 | 57.475
[7 empty] | 1392 | 57.519
[7 empty] | 1393 | 57.563
[7 empty] | 1394 | 57.607
[7 empty] | 1395 | 57.651
[7 empty] | 1396 | 57.695
[7 empty] | 1397 | 57.738
[7 empty] | 1398 | 57.782
[7 empty] | 1399 | 57.826
[7 empty] | 1400 | 57.87
[7 empty] | 1401 | 57.914
[7 empty] | 1402 | 57.958
[7 empty] | 1403 | 58.002
[7 empty] | 1404 | 58.045
[7 empty] | 1405 | 58.089
[7 empty] | 1406 | 58.133
[7 empty] | 1407 | 58.177
[7 empty] | 1408 | 58.221
[7 empty] | 1409 | 58.265
[7 empty] | 1410 | 58.308
[7 empty] | 1411 | 58.352
[7 empty] | 1412 | 58.396
[7 empty] | 1413 | 58.44
[7 empty] | 1414 | 58.484
[7 empty] | 1415 | 58.527
[7 empty] | 1416 | 58.571
[7 empty] | 1417 | 58.615
[7 empty] | 1418 | 58.659
[7 empty] | 1419 | 58.702
[7 empty] | 1420 | 58.746
[7 empty] | 1421 | 58.79
[7 empty] | 1422 | 58.834
[7 empty] | 1423 | 58.878
[7 empty] | 1424 | 58.921
[7 empty] | 1425 | 58.965
[7 empty] | 1426 | 59.009
[7 empty] | 1427 | 59.053
[7 empty] | 1428 | 59.096
[7 empty] | 1429 | 59.14
[7 empty] | 1430 | 59.184
[7 empty] | 1431 | 59.228
[7 empty] | 1432 | 59.271
[7 empty] | 1433 | 59.315
[7 empty] | 1434 | 59.359
[7 empty] | 1435 | 59.402
[7 empty] | 1436 | 59.446
[7 empty] | 1437 | 59.49
[7 empty] | 1438 | 59.534
[7 empty] | 1439 | 59.577
[7 empty] | 1440 | 59.621
[7 empty] | 1441 | 59.665
[7 empty] | 1442 | 59.708
[7 empty] | 1443 | 59.752
[7 empty] | 1444 | 59.796
[7 empty] | 1445 | 59.839
[7 empty] | 1446 | 59.883
[7 empty] | 1447 | 59.927
[7 empty] | 1448 | 59.97
[7 empty] | 1449 | 60.014
[7 empty] | 1450 | 60.058
[7 empty] | 1451 | 60.101
[7 empty] | 1452 | 60.145
[7 empty] | 1453 | 60.189
[7 empty] | 1454 | 60.232
[7 empty] | 1455 | 60.276
[7 empty] | 1456 | 60.32
[7 empty] | 1457 | 60.363
[7 empty] | 1458 | 60.407
[7 empty] | 1459 | 60.451
[7 empty] | 1460 | 60.494
[7 empty] | 1461 | 60.538
[7 empty] | 1462 | 60.581
[7 empty] | 1463 | 60.625
[7 empty] | 1464 | 60.669
[7 empty] | 1465 | 60.712
[7 empty] | 1466 | 60.756
[7 empty] | 1467 | 60.799
[7 empty] | 1468 | 60.843
[7 empty] | 1469 | 60.887
[7 empty] | 1470 | 60.93
[7 empty] | 1471 | 60.974
[7 empty] | 1472 | 61.017
[7 empty] | 1473 | 61.061
[7 empty] | 1474 | 61.105
[7 empty] | 1475 | 61.148
[7 empty] | 1476 | 61.192
[7 empty] | 1477 | 61.235
[7 empty] | 1478 | 61.279
[7 empty] | 1479 | 61.322
[7 empty] | 1480 | 61.366
[7 empty] | 1481 | 61.409
[7 empty] | 1482 | 61.453
[7 empty] | 1483 | 61.496
[7 empty] | 1484 | 61.54
[7 empty] | 1485 | 61.583
[7 empty] | 1486 | 61.627
[7 empty] | 1487 | 61.671
[7 empty] | 1488 | 61.714
[7 empty] | 1489 | 61.758
[7 empty] | 1490 | 61.801
[7 empty] | 1491 | 61.845
[7 empty] | 1492 | 61.888
[7 empty] | 1493 | 61.932
[7 empty] | 1494 | 61.975
[7 empty] | 1495 | 62.018
[7 empty] | 1496 | 62.062
[7 empty] | 1497 | 62.105
[7 empty] | 1498 | 62.149
[7 empty] | 1499 | 62.192
[7 empty] | 1500 | 62.236
[7 empty] | 1501 | 62.279
[7 empty] | 1502 | 62.323
[7 empty] | 1503 | 62.366
[7 empty] | 1504 | 62.41
[7 empty] | 1505 | 62.453
[7 empty] | 1506 | 62.496
[7 empty] | 1507 | 62.54
[7 empty] | 1508 | 62.583
[7 empty] | 1509 | 62.627
[7 empty] | 1510 | 62.67
[7 empty] | 1511 | 62.714
[7 empty] | 1512 | 62.757
[7 empty] | 1513 | 62.8
[7 empty] | 1514 | 62.844
[7 empty] | 1515 | 62.887
[7 empty] | 1516 | 62.931
[7 empty] | 1517 | 62.974
[7 empty] | 1518 | 63.017
[7 empty] | 1519 | 63.061
[7 empty] | 1520 | 63.104
[7 empty] | 1521 | 63.148
[7 empty] | 1522 | 63.191
[7 empty] | 1523 | 63.234
[7 empty] | 1524 | 63.278
[7 empty] | 1525 | 63.321
[7 empty] | 1526 | 63.364
[7 empty] | 1527 | 63.408
[7 empty] | 1528 | 63.451
[7 empty] | 1529 | 63.494
[7 empty] | 1530 | 63.538
[7 empty] | 1531 | 63.581
[7 empty] | 1532 | 63.624
[7 empty] | 1533 | 63.668
[7 empty] | 1534 | 63.711
[7 empty] | 1535 | 63.754
[7 empty] | 1536 | 63.798
[7 empty] | 1537 | 63.841
[7 empty] | 1538 | 63.884
[7 empty] | 1539 | 63.927
[7 empty] | 1540 | 63.971
[7 empty] | 1541 | 64.014
[7 empty] | 1542 | 64.057
[7 empty] | 1543 | 64.101
[7 empty] | 1544 | 64.144
[7 empty] | 1545 | 64.187
[7 empty] | 1546 | 64.23
[7 empty] | 1547 | 64.274
[7 empty] | 1548 | 64.317
[7 empty] | 1549 | 64.36
[7 empty] | 1550 | 64.403
[7 empty] | 1551 | 64.447
[7 empty] | 1552 | 64.49
[7 empty] | 1553 | 64.533
[7 empty] | 1554 | 64.576
[7 empty] | 1555 | 64.619
[7 empty] | 1556 | 64.663
[7 empty] | 1557 | 64.706
[7 empty] | 1558 | 64.749
[7 empty] | 1559 | 64.792
[7 empty] | 1560 | 64.835
[7 empty] | 1561 | 64.879
[7 empty] | 1562 | 64.922
[7 empty] | 1563 | 64.965
[7 empty] | 1564 | 65.008
[7 empty] | 1565 | 65.051
[7 empty] | 1566 | 65.094
[7 empty] | 1567 | 65.138
[7 empty] | 1568 | 65.181
[7 empty] | 1569 | 65.224
[7 empty] | 1570 | 65.267
[7 empty] | 1571 | 65.31
[7 empty] | 1572 | 65.353
[7 empty] | 1573 | 65.396
[7 empty] | 1574 | 65.44
[7 empty] | 1575 | 65.483
[7 empty] | 1576 | 65.526
[7 empty] | 1577 | 65.569
[7 empty] | 1578 | 65.612
[7 empty] | 1579 | 65.655
[7 empty] | 1580 | 65.698
[7 empty] | 1581 | 65.741
[7 empty] | 1582 | 65.784
[7 empty] | 1583 | 65.827
[7 empty] | 1584 | 65.871
[7 empty] | 1585 | 65.914
[7 empty] | 1586 | 65.957
[7 empty] | 1587 | 66
[7 empty] | 1588 | 66.043
[7 empty] | 1589 | 66.086
[7 empty] | 1590 | 66.129
[7 empty] | 1591 | 66.172
[7 empty] | 1592 | 66.215
[7 empty] | 1593 | 66.258
[7 empty] | 1594 | 66.301
[7 empty] | 1595 | 66.344
[7 empty] | 1596 | 66.387
[7 empty] | 1597 | 66.43
[7 empty] | 1598 | 66.473
[7 empty] | 1599 | 66.516
[7 empty] | 1600 | 66.559
[7 empty] | 1601 | 66.602
[7 empty] | 1602 | 66.645
[7 empty] | 1603 | 66.688
[7 empty] | 1604 | 66.731
[7 empty] | 1605 | 66.774
[7 empty] | 1606 | 66.817
[7 empty] | 1607 | 66.86
[7 empty] | 1608 | 66.903
[7 empty] | 1609 | 66.946
[7 empty] | 1610 | 66.989
[7 empty] | 1611 | 67.031
[7 empty] | 1612 | 67.074
[7 empty] | 1613 | 67.117
[7 empty] | 1614 | 67.16
[7 empty] | 1615 | 67.203
[7 empty] | 1616 | 67.246
[7 empty] | 1617 | 67.289
[7 empty] | 1618 | 67.332
[7 empty] | 1619 | 67.375
[7 empty] | 1620 | 67.418
[7 empty] | 1621 | 67.46
[7 empty] | 1622 | 67.503
[7 empty] | 1623 | 67.546
[7 empty] | 1624 | 67.589
[7 empty] | 1625 | 67.632
[7 empty] | 1626 | 67.675
[7 empty] | 1627 | 67.718
[7 empty] | 1628 | 67.76
[7 empty] | 1629 | 67.803
[7 empty] | 1630 | 67.846
[7 empty] | 1631 | 67.889
[7 empty] | 1632 | 67.932
[7 empty] | 1633 | 67.974
[7 empty] | 1634 | 68.017
[7 empty] | 1635 | 68.06
[7 empty] | 1636 | 68.103
[7 empty] | 1637 | 68.146
[7 empty] | 1638 | 68.188
[7 empty] | 1639 | 68.231
[7 empty] | 1640 | 68.274
[7 empty] | 1641 | 68.317
[7 empty] | 1642 | 68.359
[7 empty] | 1643 | 68.402
[7 empty] | 1644 | 68.445
[7 empty] | 1645 | 68.488
[7 empty] | 1646 | 68.53
[7 empty] | 1647 | 68.573
[7 empty] | 1648 | 68.616
[7 empty] | 1649 | 68.659
[7 empty] | 1650 | 68.701
[7 empty] | 1651 | 68.744
[7 empty] | 1652 | 68.787
[7 empty] | 1653 | 68.829
[7 empty] | 1654 | 68.872
[7 empty] | 1655 | 68.915
[7 empty] | 1656 | 68.957
[7 empty] | 1657 | 69
[7 empty] | 1658 | 69.043
[7 empty] | 1659 | 69.085
[7 empty] | 1660 | 69.128
[7 empty] | 1661 | 69.171
[7 empty] | 1662 | 69.213
[7 empty] | 1663 | 69.256
[7 empty] | 1664 | 69.298
[7 empty] | 1665 | 69.341
[7 empty] | 1666 | 69.384
[7 empty] | 1667 | 69.426
[7 empty] | 1668 | 69.469
[7 empty] | 1669 | 69.511
[7 empty] | 1670 | 69.554
[7 empty] | 1671 | 69.597
[7 empty] | 1672 | 69.639
[7 empty] | 1673 | 69.682
[7 empty] | 1674 | 69.724
[7 empty] | 1675 | 69.767
[7 empty] | 1676 | 69.809
[7 empty] | 1677 | 69.852
[7 empty] | 1678 | 69.894
[7 empty] | 1679 | 69.937
[7 empty] | 1680 | 69.979
[7 empty] | 1681 | 70.022
[7 empty] | 1682 | 70.064
[7 empty] | 1683 | 70.107
[7 empty] | 1684 | 70.149
[7 empty] | 1685 | 70.192
[7 empty] | 1686 | 70.234
[7 empty] | 1687 | 70.277
[7 empty] | 1688 | 70.319
[7 empty] | 1689 | 70.362
[7 empty] | 1690 | 70.404
[7 empty] | 1691 | 70.447
[7 empty] | 1692 | 70.489
[7 empty] | 1693 | 70.531
[7 empty] | 1694 | 70.574
[7 empty] | 1695 | 70.616
[7 empty] | 1696 | 70.659
[7 empty] | 1697 | 70.701
[7 empty] | 1698 | 70.744
[7 empty] | 1699 | 70.786
[7 empty] | 1700 | 70.828
[7 empty] | 1701 | 70.871
[7 empty] | 1702 | 70.913
[7 empty] | 1703 | 70.955
[7 empty] | 1704 | 70.998
[7 empty] | 1705 | 71.04
[7 empty] | 1706 | 71.082
[7 empty] | 1707 | 71.125
[7 empty] | 1708 | 71.167
[7 empty] | 1709 | 71.209
[7 empty] | 1710 | 71.252
[7 empty] | 1711 | 71.294
[7 empty] | 1712 | 71.336
[7 empty] | 1713 | 71.379
[7 empty] | 1714 | 71.421
[7 empty] | 1715 | 71.463
[7 empty] | 1716 | 71.506
[7 empty] | 1717 | 71.548
[7 empty] | 1718 | 71.59
[7 empty] | 1719 | 71.632
[7 empty] | 1720 | 71.675
[7 empty] | 1721 | 71.717
[7 empty] | 1722 | 71.759
[7 empty] | 1723 | 71.801
[7 empty] | 1724 | 71.844
[7 empty] | 1725 | 71.886
[7 empty] | 1726 | 71.928
[7 empty] | 1727 | 71.97
[7 empty] | 1728 | 72.012
[7 empty] | 1729 | 72.055
[7 empty] | 1730 | 72.097
[7 empty] | 1731 | 72.139
[7 empty] | 1732 | 72.181
[7 empty] | 1733 | 72.223
[7 empty] | 1734 | 72.266
[7 empty] | 1735 | 72.308
[7 empty] | 1736 | 72.35
[7 empty] | 1737 | 72.392
[7 empty] | 1738 | 72.434
[7 empty] | 1739 | 72.476
[7 empty] | 1740 | 72.518
[7 empty] | 1741 | 72.561
[7 empty] | 1742 | 72.603
[7 empty] | 1743 | 72.645
[7 empty] | 1744 | 72.687
[7 empty] | 1745 | 72.729
[7 empty] | 1746 | 72.771
[7 empty] | 1747 | 72.813
[7 empty] | 1748 | 72.855
[7 empty] | 1749 | 72.897
[7 empty] | 1750 | 72.939
[7 empty] | 1751 | 72.981
[7 empty] | 1752 | 73.023
[7 empty] | 1753 | 73.066
[7 empty] | 1754 | 73.108
[7 empty] | 1755 | 73.15
[7 empty] | 1756 | 73.192
[7 empty] | 1757 | 73.234
[7 empty] | 1758 | 73.276
[7 empty] | 1759 | 73.318
[7 empty] | 1760 | 73.36
[7 empty] | 1761 | 73.402
[7 empty] | 1762 | 73.444
[7 empty] | 1763 | 73.486
[7 empty] | 1764 | 73.528
[7 empty] | 1765 | 73.57
[7 empty] | 1766 | 73.612
[7 empty] | 1767 | 73.654
[7 empty] | 1768 | 73.696
[7 empty] | 1769 | 73.738
[7 empty] | 1770 | 73.78
[7 empty] | 1771 | 73.821
[7 empty] | 1772 | 73.863
[7 empty] | 1773 | 73.905
[7 empty] | 1774 | 73.947
[7 empty] | 1775 | 73.989
[7 empty] | 1776 | 74.031
[7 empty] | 1777 | 74.073
[7 empty] | 1778 | 74.115
[7 empty] | 1779 | 74.157
[7 empty] | 1780 | 74.199
[7 empty] | 1781 | 74.241
[7 empty] | 1782 | 74.283
[7 empty] | 1783 | 74.324
[7 empty] | 1784 | 74.366
[7 empty] | 1785 | 74.408
[7 empty] | 1786 | 74.45
[7 empty] | 1787 | 74.492
[7 empty] | 1788 | 74.534
[7 empty] | 1789 | 74.576
[7 empty] | 1790 | 74.618
[7 empty] | 1791 | 74.659
[7 empty] | 1792 | 74.701
[7 empty] | 1793 | 74.743
[7 empty] | 1794 | 74.785
[7 empty] | 1795 | 74.827
[7 empty] | 1796 | 74.869
[7 empty] | 1797 | 74.91
[7 empty] | 1798 | 74.952
[7 empty] | 1799 | 74.994
[7 empty] | 1800 | 75.036
[7 empty] | 1801 | 75.078
[7 empty] | 1802 | 75.12
[7 empty] | 1803 | 75.161
[7 empty] | 1804 | 75.203
[7 empty] | 1805 | 75.245
[7 empty] | 1806 | 75.287
[7 empty] | 1807 | 75.329
[7 empty] | 1808 | 75.37
[7 empty] | 1809 | 75.412
[7 empty] | 1810 | 75.454
[7 empty] | 1811 | 75.496
[7 empty] | 1812 | 75.538
[7 empty] | 1813 | 75.579
[7 empty] | 1814 | 75.621
[7 empty] | 1815 | 75.663
[7 empty] | 1816 | 75.705
[7 empty] | 1817 | 75.746
[7 empty] | 1818 | 75.788
[7 empty] | 1819 | 75.83
[7 empty] | 1820 | 75.872
[7 empty] | 1821 | 75.913
[7 empty] | 1822 | 75.955
[7 empty] | 1823 | 75.997
[7 empty] | 1824 | 76.039
[7 empty] | 1825 | 76.081
[7 empty] | 1826 | 76.122
[7 empty] | 1827 | 76.164
[7 empty] | 1828 | 76.206
[7 empty] | 1829 | 76.248
[7 empty] | 1830 | 76.289
[7 empty] | 1831 | 76.331
[7 empty] | 1832 | 76.373
[7 empty] | 1833 | 75.997
[7 empty] | 1834 | 76.039
[7 empty] | 1835 | 76.081
[7 empty] | 1836 | 76.122
[7 empty] | 1837 | 76.164
[7 empty] | 1838 | 76.206
[7 empty] | 1839 | 76.248
[7 empty] | 1840 | 76.289
[7 empty] | 1841 | 76.331
[7 empty] | 1842 | 76.373

## Type J
[7 empty] | Temp(F) | mV
mV | Temp(F) | [5 empty] | -346 | -8.095
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -345 | -8.085
23.32 | 800 | 1146 | 23.289 | 23.32 | 799 | 800 | -344 | -8.074
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -343 | -8.063
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -342 | -8.052
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -341 | -8.041
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -340 | -8.03
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -339 | -8.019
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -338 | -8.008
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -337 | -7.996
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -336 | -7.985
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -335 | -7.973
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -334 | -7.962
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -333 | -7.95
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -332 | -7.938
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -331 | -7.927
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -330 | -7.915
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -329 | -7.903
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -328 | -7.89
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -327 | -7.878
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -326 | -7.866
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -325 | -7.854
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -324 | -7.841
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -323 | -7.829
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -322 | -7.816
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -321 | -7.804
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -320 | -7.791
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -319 | -7.778
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -318 | -7.765
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -317 | -7.752
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -316 | -7.739
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -315 | -7.726
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -314 | -7.713
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -313 | -7.699
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -312 | -7.686
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -311 | -7.672
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -310 | -7.659
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -309 | -7.645
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -308 | -7.632
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -307 | -7.618
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -306 | -7.604
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -305 | -7.59
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -304 | -7.576
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -303 | -7.562
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -302 | -7.548
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -301 | -7.534
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -300 | -7.519
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -299 | -7.505
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -298 | -7.491
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -297 | -7.476
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -296 | -7.462
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -295 | -7.447
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -294 | -7.432
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -293 | -7.417
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -292 | -7.403
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -291 | -7.388
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -290 | -7.373
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -289 | -7.357
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -288 | -7.342
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -287 | -7.327
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -286 | -7.312
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -285 | -7.296
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -284 | -7.281
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -283 | -7.265
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -282 | -7.25
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -281 | -7.234
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -280 | -7.219
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -279 | -7.203
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -278 | -7.187
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -277 | -7.171
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -276 | -7.155
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -275 | -7.139
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -274 | -7.123
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -273 | -7.107
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -272 | -7.09
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -271 | -7.074
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -270 | -7.058
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -269 | -7.041
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -268 | -7.025
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -267 | -7.008
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -266 | -6.991
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -265 | -6.975
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -264 | -6.958
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -263 | -6.941
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -262 | -6.924
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -261 | -6.907
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -260 | -6.89
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -259 | -6.873
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -258 | -6.856
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -257 | -6.839
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -256 | -6.821
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -255 | -6.804
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -254 | -6.787
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -253 | -6.769
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -252 | -6.752
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -251 | -6.734
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -250 | -6.716
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -249 | -6.699
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -248 | -6.681
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -247 | -6.663
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -246 | -6.645
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -245 | -6.627
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -244 | -6.609
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -243 | -6.591
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -242 | -6.573
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -241 | -6.555
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -240 | -6.536
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -239 | -6.518
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -238 | -6.5
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -237 | -6.481
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -236 | -6.463
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -235 | -6.444
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -234 | -6.426
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -233 | -6.407
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -232 | -6.388
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -231 | -6.37
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -230 | -6.351
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -229 | -6.332
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -228 | -6.313
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -227 | -6.294
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -226 | -6.275
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -225 | -6.256
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -224 | -6.236
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -223 | -6.217
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -222 | -6.198
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -221 | -6.179
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -220 | -6.159
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -219 | -6.14
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -218 | -6.12
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -217 | -6.101
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -216 | -6.081
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -215 | -6.061
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -214 | -6.042
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -213 | -6.022
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -212 | -6.002
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -211 | -5.982
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -210 | -5.962
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -209 | -5.942
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -208 | -5.922
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -207 | -5.902
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -206 | -5.882
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -205 | -5.862
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -204 | -5.842
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -203 | -5.821
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -202 | -5.801
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -201 | -5.781
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -200 | -5.76
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -199 | -5.74
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -198 | -5.719
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -197 | -5.699
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -196 | -5.678
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -195 | -5.657
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -194 | -5.637
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -193 | -5.616
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -192 | -5.595
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -191 | -5.574
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -190 | -5.553
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -189 | -5.532
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -188 | -5.511
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -187 | -5.49
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -186 | -5.469
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -185 | -5.448
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -184 | -5.426
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -183 | -5.405
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -182 | -5.384
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -181 | -5.363
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -180 | -5.341
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -179 | -5.32
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -178 | -5.298
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -177 | -5.277
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -176 | -5.255
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -175 | -5.233
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -174 | -5.212
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -173 | -5.19
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -172 | -5.168
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -171 | -5.146
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -170 | -5.125
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -169 | -5.103
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -168 | -5.081
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -167 | -5.059
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -166 | -5.037
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -165 | -5.015
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -164 | -4.992
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -163 | -4.97
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -162 | -4.948
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -161 | -4.926
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -160 | -4.903
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -159 | -4.881
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -158 | -4.859
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -157 | -4.836
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -156 | -4.814
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -155 | -4.791
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -154 | -4.769
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -153 | -4.746
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -152 | -4.724
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -151 | -4.701
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -150 | -4.678
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -149 | -4.655
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -148 | -4.633
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -147 | -4.61
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -146 | -4.587
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -145 | -4.564
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -144 | -4.541
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -143 | -4.518
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -142 | -4.495
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -141 | -4.472
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -140 | -4.449
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -139 | -4.425
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -138 | -4.402
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -137 | -4.379
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -136 | -4.356
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -135 | -4.332
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -134 | -4.309
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -133 | -4.286
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -132 | -4.262
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -131 | -4.239
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -130 | -4.215
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -129 | -4.192
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -128 | -4.168
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -127 | -4.144
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -126 | -4.121
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -125 | -4.097
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -124 | -4.073
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -123 | -4.05
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -122 | -4.026
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -121 | -4.002
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -120 | -3.978
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -119 | -3.954
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -118 | -3.93
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -117 | -3.906
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -116 | -3.882
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -115 | -3.858
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -114 | -3.834
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -113 | -3.81
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -112 | -3.786
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -111 | -3.761
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -110 | -3.737
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -109 | -3.713
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -108 | -3.688
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -107 | -3.664
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -106 | -3.64
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -105 | -3.615
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -104 | -3.591
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -103 | -3.566
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -102 | -3.542
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -101 | -3.517
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -100 | -3.493
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -99 | -3.468
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -98 | -3.443
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -97 | -3.419
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -96 | -3.394
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -95 | -3.369
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -94 | -3.344
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -93 | -3.32
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -92 | -3.295
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -91 | -3.27
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -90 | -3.245
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -89 | -3.22
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -88 | -3.195
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -87 | -3.17
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -86 | -3.145
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -85 | -3.12
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -84 | -3.095
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -83 | -3.07
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -82 | -3.044
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -81 | -3.019
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -80 | -2.994
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -79 | -2.969
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -78 | -2.943
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -77 | -2.918
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -76 | -2.893
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -75 | -2.867
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -74 | -2.842
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -73 | -2.817
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -72 | -2.791
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -71 | -2.766
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -70 | -2.74
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -69 | -2.714
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -68 | -2.689
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -67 | -2.663
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -66 | -2.638
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -65 | -2.612
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -64 | -2.586
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -63 | -2.56
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -62 | -2.535
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -61 | -2.509
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -60 | -2.483
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -59 | -2.457
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -58 | -2.431
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -57 | -2.405
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -56 | -2.379
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -55 | -2.353
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -54 | -2.327
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -53 | -2.301
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -52 | -2.275
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -51 | -2.249
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -50 | -2.223
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -49 | -2.197
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -48 | -2.171
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -47 | -2.145
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -46 | -2.118
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -45 | -2.092
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -44 | -2.066
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -43 | -2.04
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -42 | -2.013
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -41 | -1.987
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -40 | -1.961
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -39 | -1.934
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -38 | -1.908
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -37 | -1.881
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -36 | -1.855
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -35 | -1.828
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -34 | -1.802
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -33 | -1.775
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -32 | -1.749
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -31 | -1.722
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -30 | -1.695
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -29 | -1.669
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -28 | -1.642
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -27 | -1.615
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -26 | -1.589
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -25 | -1.562
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -24 | -1.535
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -23 | -1.508
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -22 | -1.482
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -21 | -1.455
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -20 | -1.428
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -19 | -1.401
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -18 | -1.374
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -17 | -1.347
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -16 | -1.32
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -15 | -1.293
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -14 | -1.266
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -13 | -1.239
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -12 | -1.212
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -11 | -1.185
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -10 | -1.158
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -9 | -1.131
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -8 | -1.104
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -7 | -1.076
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -6 | -1.049
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -5 | -1.022
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -4 | -0.995
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -3 | -0.967
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -2 | -0.94
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | -1 | -0.913
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 0 | -0.886
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1 | -0.858
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2 | -0.831
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 3 | -0.803
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 4 | -0.776
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 5 | -0.749
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 6 | -0.721
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 7 | -0.694
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 8 | -0.666
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 9 | -0.639
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 10 | -0.611
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 11 | -0.583
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 12 | -0.556
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 13 | -0.528
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 14 | -0.501
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 15 | -0.473
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 16 | -0.445
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 17 | -0.418
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 18 | -0.39
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 19 | -0.362
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 20 | -0.334
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 21 | -0.307
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 22 | -0.279
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 23 | -0.251
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 24 | -0.223
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 25 | -0.195
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 26 | -0.168
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 27 | -0.14
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 28 | -0.112
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 29 | -0.084
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 30 | -0.056
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 31 | -0.028
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 32 | 0
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 33 | 0.028
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 34 | 0.056
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 35 | 0.084
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 36 | 0.112
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 37 | 0.14
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 38 | 0.168
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 39 | 0.196
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 40 | 0.225
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 41 | 0.253
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 42 | 0.281
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 43 | 0.309
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 44 | 0.337
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 45 | 0.365
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 46 | 0.394
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 47 | 0.422
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 48 | 0.45
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 49 | 0.478
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 50 | 0.507
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 51 | 0.535
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 52 | 0.563
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 53 | 0.592
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 54 | 0.62
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 55 | 0.649
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 56 | 0.677
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 57 | 0.705
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 58 | 0.734
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 59 | 0.762
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 60 | 0.791
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 61 | 0.819
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 62 | 0.848
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 63 | 0.876
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 64 | 0.905
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 65 | 0.933
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 66 | 0.962
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 67 | 0.991
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 68 | 1.019
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 69 | 1.048
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 70 | 1.076
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 71 | 1.105
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 72 | 1.134
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 73 | 1.162
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 74 | 1.191
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 75 | 1.22
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 76 | 1.249
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 77 | 1.277
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 78 | 1.306
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 79 | 1.335
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 80 | 1.364
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 81 | 1.392
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 82 | 1.421
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 83 | 1.45
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 84 | 1.479
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 85 | 1.508
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 86 | 1.537
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 87 | 1.566
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 88 | 1.594
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 89 | 1.623
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 90 | 1.652
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 91 | 1.681
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 92 | 1.71
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 93 | 1.739
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 94 | 1.768
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 95 | 1.797
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 96 | 1.826
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 97 | 1.855
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 98 | 1.884
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 99 | 1.913
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 100 | 1.942
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 101 | 1.972
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 102 | 2.001
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 103 | 2.03
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 104 | 2.059
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 105 | 2.088
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 106 | 2.117
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 107 | 2.146
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 108 | 2.175
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 109 | 2.205
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 110 | 2.234
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 111 | 2.263
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 112 | 2.292
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 113 | 2.322
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 114 | 2.351
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 115 | 2.38
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 116 | 2.409
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 117 | 2.439
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 118 | 2.468
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 119 | 2.497
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 120 | 2.527
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 121 | 2.556
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 122 | 2.585
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 123 | 2.615
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 124 | 2.644
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 125 | 2.673
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 126 | 2.703
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 127 | 2.732
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 128 | 2.762
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 129 | 2.791
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 130 | 2.821
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 131 | 2.85
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 132 | 2.88
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 133 | 2.909
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 134 | 2.938
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 135 | 2.968
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 136 | 2.997
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 137 | 3.027
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 138 | 3.057
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 139 | 3.086
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 140 | 3.116
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 141 | 3.145
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 142 | 3.175
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 143 | 3.204
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 144 | 3.234
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 145 | 3.264
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 146 | 3.293
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 147 | 3.323
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 148 | 3.353
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 149 | 3.382
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 150 | 3.412
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 151 | 3.442
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 152 | 3.471
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 153 | 3.501
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 154 | 3.531
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 155 | 3.56
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 156 | 3.59
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 157 | 3.62
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 158 | 3.65
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 159 | 3.679
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 160 | 3.709
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 161 | 3.739
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 162 | 3.769
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 163 | 3.798
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 164 | 3.828
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 165 | 3.858
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 166 | 3.888
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 167 | 3.918
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 168 | 3.948
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 169 | 3.977
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 170 | 4.007
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 171 | 4.037
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 172 | 4.067
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 173 | 4.097
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 174 | 4.127
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 175 | 4.157
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 176 | 4.187
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 177 | 4.217
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 178 | 4.246
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 179 | 4.276
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 180 | 4.306
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 181 | 4.336
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 182 | 4.366
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 183 | 4.396
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 184 | 4.426
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 185 | 4.456
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 186 | 4.486
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 187 | 4.516
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 188 | 4.546
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 189 | 4.576
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 190 | 4.606
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 191 | 4.636
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 192 | 4.666
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 193 | 4.696
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 194 | 4.726
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 195 | 4.757
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 196 | 4.787
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 197 | 4.817
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 198 | 4.847
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 199 | 4.877
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 200 | 4.907
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 201 | 4.937
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 202 | 4.967
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 203 | 4.997
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 204 | 5.028
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 205 | 5.058
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 206 | 5.088
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 207 | 5.118
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 208 | 5.148
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 209 | 5.178
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 210 | 5.209
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 211 | 5.239
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 212 | 5.269
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 213 | 5.299
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 214 | 5.329
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 215 | 5.36
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 216 | 5.39
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 217 | 5.42
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 218 | 5.45
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 219 | 5.48
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 220 | 5.511
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 221 | 5.541
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 222 | 5.571
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 223 | 5.602
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 224 | 5.632
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 225 | 5.662
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 226 | 5.692
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 227 | 5.723
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 228 | 5.753
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 229 | 5.783
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 230 | 5.814
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 231 | 5.844
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 232 | 5.874
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 233 | 5.905
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 234 | 5.935
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 235 | 5.965
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 236 | 5.996
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 237 | 6.026
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 238 | 6.056
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 239 | 6.087
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 240 | 6.117
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 241 | 6.147
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 242 | 6.178
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 243 | 6.208
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 244 | 6.239
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 245 | 6.269
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 246 | 6.299
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 247 | 6.33
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 248 | 6.36
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 249 | 6.391
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 250 | 6.421
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 251 | 6.452
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 252 | 6.482
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 253 | 6.512
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 254 | 6.543
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 255 | 6.573
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 256 | 6.604
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 257 | 6.634
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 258 | 6.665
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 259 | 6.695
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 260 | 6.726
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 261 | 6.756
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 262 | 6.787
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 263 | 6.817
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 264 | 6.848
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 265 | 6.878
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 266 | 6.909
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 267 | 6.939
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 268 | 6.97
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 269 | 7
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 270 | 7.031
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 271 | 7.061
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 272 | 7.092
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 273 | 7.122
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 274 | 7.153
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 275 | 7.184
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 276 | 7.214
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 277 | 7.245
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 278 | 7.275
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 279 | 7.306
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 280 | 7.336
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 281 | 7.367
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 282 | 7.398
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 283 | 7.428
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 284 | 7.459
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 285 | 7.489
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 286 | 7.52
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 287 | 7.55
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 288 | 7.581
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 289 | 7.612
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 290 | 7.642
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 291 | 7.673
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 292 | 7.704
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 293 | 7.734
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 294 | 7.765
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 295 | 7.795
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 296 | 7.826
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 297 | 7.857
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 298 | 7.887
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 299 | 7.918
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 300 | 7.949
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 301 | 7.979
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 302 | 8.01
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 303 | 8.041
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 304 | 8.071
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 305 | 8.102
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 306 | 8.133
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 307 | 8.163
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 308 | 8.194
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 309 | 8.225
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 310 | 8.255
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 311 | 8.286
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 312 | 8.317
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 313 | 8.347
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 314 | 8.378
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 315 | 8.409
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 316 | 8.439
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 317 | 8.47
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 318 | 8.501
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 319 | 8.532
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 320 | 8.562
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 321 | 8.593
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 322 | 8.624
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 323 | 8.654
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 324 | 8.685
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 325 | 8.716
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 326 | 8.747
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 327 | 8.777
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 328 | 8.808
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 329 | 8.839
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 330 | 8.869
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 331 | 8.9
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 332 | 8.931
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 333 | 8.962
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 334 | 8.992
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 335 | 9.023
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 336 | 9.054
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 337 | 9.085
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 338 | 9.115
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 339 | 9.146
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 340 | 9.177
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 341 | 9.208
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 342 | 9.238
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 343 | 9.269
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 344 | 9.3
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 345 | 9.331
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 346 | 9.362
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 347 | 9.392
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 348 | 9.423
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 349 | 9.454
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 350 | 9.485
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 351 | 9.515
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 352 | 9.546
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 353 | 9.577
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 354 | 9.608
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 355 | 9.639
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 356 | 9.669
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 357 | 9.7
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 358 | 9.731
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 359 | 9.762
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 360 | 9.793
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 361 | 9.823
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 362 | 9.854
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 363 | 9.885
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 364 | 9.916
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 365 | 9.947
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 366 | 9.977
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 367 | 10.008
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 368 | 10.039
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 369 | 10.07
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 370 | 10.101
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 371 | 10.131
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 372 | 10.162
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 373 | 10.193
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 374 | 10.224
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 375 | 10.255
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 376 | 10.285
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 377 | 10.316
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 378 | 10.347
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 379 | 10.378
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 380 | 10.409
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 381 | 10.44
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 382 | 10.47
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 383 | 10.501
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 384 | 10.532
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 385 | 10.563
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 386 | 10.594
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 387 | 10.625
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 388 | 10.655
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 389 | 10.686
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 390 | 10.717
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 391 | 10.748
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 392 | 10.779
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 393 | 10.81
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 394 | 10.84
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 395 | 10.871
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 396 | 10.902
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 397 | 10.933
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 398 | 10.964
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 399 | 10.995
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 400 | 11.025
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 401 | 11.056
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 402 | 11.087
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 403 | 11.118
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 404 | 11.149
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 405 | 11.18
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 406 | 11.211
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 407 | 11.241
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 408 | 11.272
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 409 | 11.303
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 410 | 11.334
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 411 | 11.365
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 412 | 11.396
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 413 | 11.426
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 414 | 11.457
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 415 | 11.488
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 416 | 11.519
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 417 | 11.55
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 418 | 11.581
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 419 | 11.612
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 420 | 11.642
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 421 | 11.673
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 422 | 11.704
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 423 | 11.735
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 424 | 11.766
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 425 | 11.797
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 426 | 11.828
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 427 | 11.858
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 428 | 11.889
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 429 | 11.92
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 430 | 11.951
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 431 | 11.982
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 432 | 12.013
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 433 | 12.044
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 434 | 12.074
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 435 | 12.105
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 436 | 12.136
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 437 | 12.167
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 438 | 12.198
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 439 | 12.229
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 440 | 12.26
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 441 | 12.29
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 442 | 12.321
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 443 | 12.352
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 444 | 12.383
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 445 | 12.414
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 446 | 12.445
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 447 | 12.476
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 448 | 12.506
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 449 | 12.537
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 450 | 12.568
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 451 | 12.599
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 452 | 12.63
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 453 | 12.661
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 454 | 12.691
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 455 | 12.722
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 456 | 12.753
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 457 | 12.784
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 458 | 12.815
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 459 | 12.846
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 460 | 12.877
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 461 | 12.907
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 462 | 12.938
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 463 | 12.969
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 464 | 13
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 465 | 13.031
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 466 | 13.062
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 467 | 13.093
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 468 | 13.123
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 469 | 13.154
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 470 | 13.185
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 471 | 13.216
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 472 | 13.247
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 473 | 13.278
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 474 | 13.308
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 475 | 13.339
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 476 | 13.37
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 477 | 13.401
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 478 | 13.432
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 479 | 13.463
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 480 | 13.494
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 481 | 13.524
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 482 | 13.555
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 483 | 13.586
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 484 | 13.617
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 485 | 13.648
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 486 | 13.679
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 487 | 13.709
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 488 | 13.74
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 489 | 13.771
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 490 | 13.802
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 491 | 13.833
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 492 | 13.864
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 493 | 13.894
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 494 | 13.925
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 495 | 13.956
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 496 | 13.987
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 497 | 14.018
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 498 | 14.049
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 499 | 14.079
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 500 | 14.11
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 501 | 14.141
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 502 | 14.172
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 503 | 14.203
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 504 | 14.233
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 505 | 14.264
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 506 | 14.295
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 507 | 14.326
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 508 | 14.357
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 509 | 14.388
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 510 | 14.418
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 511 | 14.449
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 512 | 14.48
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 513 | 14.511
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 514 | 14.542
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 515 | 14.573
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 516 | 14.603
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 517 | 14.634
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 518 | 14.665
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 519 | 14.696
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 520 | 14.727
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 521 | 14.757
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 522 | 14.788
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 523 | 14.819
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 524 | 14.85
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 525 | 14.881
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 526 | 14.911
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 527 | 14.942
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 528 | 14.973
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 529 | 15.004
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 530 | 15.035
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 531 | 15.065
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 532 | 15.096
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 533 | 15.127
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 534 | 15.158
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 535 | 15.189
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 536 | 15.219
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 537 | 15.25
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 538 | 15.281
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 539 | 15.312
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 540 | 15.343
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 541 | 15.373
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 542 | 15.404
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 543 | 15.435
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 544 | 15.466
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 545 | 15.496
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 546 | 15.527
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 547 | 15.558
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 548 | 15.589
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 549 | 15.62
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 550 | 15.65
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 551 | 15.681
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 552 | 15.712
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 553 | 15.743
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 554 | 15.773
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 555 | 15.804
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 556 | 15.835
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 557 | 15.866
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 558 | 15.897
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 559 | 15.927
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 560 | 15.958
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 561 | 15.989
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 562 | 16.02
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 563 | 16.05
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 564 | 16.081
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 565 | 16.112
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 566 | 16.143
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 567 | 16.173
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 568 | 16.204
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 569 | 16.235
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 570 | 16.266
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 571 | 16.296
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 572 | 16.327
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 573 | 16.358
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 574 | 16.389
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 575 | 16.419
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 576 | 16.45
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 577 | 16.481
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 578 | 16.512
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 579 | 16.542
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 580 | 16.573
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 581 | 16.604
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 582 | 16.635
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 583 | 16.665
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 584 | 16.696
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 585 | 16.727
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 586 | 16.758
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 587 | 16.788
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 588 | 16.819
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 589 | 16.85
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 590 | 16.881
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 591 | 16.911
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 592 | 16.942
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 593 | 16.973
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 594 | 17.003
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 595 | 17.034
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 596 | 17.065
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 597 | 17.096
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 598 | 17.126
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 599 | 17.157
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 600 | 17.188
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 601 | 17.219
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 602 | 17.249
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 603 | 17.28
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 604 | 17.311
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 605 | 17.341
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 606 | 17.372
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 607 | 17.403
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 608 | 17.434
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 609 | 17.464
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 610 | 17.495
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 611 | 17.526
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 612 | 17.556
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 613 | 17.587
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 614 | 17.618
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 615 | 17.649
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 616 | 17.679
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 617 | 17.71
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 618 | 17.741
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 619 | 17.771
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 620 | 17.802
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 621 | 17.833
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 622 | 17.863
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 623 | 17.894
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 624 | 17.925
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 625 | 17.955
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 626 | 17.986
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 627 | 18.017
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 628 | 18.048
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 629 | 18.078
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 630 | 18.109
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 631 | 18.14
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 632 | 18.17
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 633 | 18.201
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 634 | 18.232
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 635 | 18.262
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 636 | 18.293
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 637 | 18.324
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 638 | 18.354
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 639 | 18.385
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 640 | 18.416
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 641 | 18.446
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 642 | 18.477
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 643 | 18.508
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 644 | 18.538
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 645 | 18.569
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 646 | 18.6
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 647 | 18.63
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 648 | 18.661
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 649 | 18.692
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 650 | 18.722
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 651 | 18.753
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 652 | 18.784
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 653 | 18.814
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 654 | 18.845
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 655 | 18.876
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 656 | 18.906
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 657 | 18.937
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 658 | 18.968
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 659 | 18.998
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 660 | 19.029
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 661 | 19.06
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 662 | 19.09
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 663 | 19.121
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 664 | 19.152
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 665 | 19.182
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 666 | 19.213
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 667 | 19.244
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 668 | 19.274
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 669 | 19.305
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 670 | 19.336
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 671 | 19.366
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 672 | 19.397
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 673 | 19.428
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 674 | 19.458
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 675 | 19.489
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 676 | 19.52
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 677 | 19.55
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 678 | 19.581
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 679 | 19.612
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 680 | 19.642
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 681 | 19.673
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 682 | 19.704
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 683 | 19.734
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 684 | 19.765
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 685 | 19.795
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 686 | 19.826
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 687 | 19.857
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 688 | 19.887
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 689 | 19.918
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 690 | 19.949
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 691 | 19.979
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 692 | 20.01
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 693 | 20.041
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 694 | 20.071
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 695 | 20.102
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 696 | 20.132
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 697 | 20.163
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 698 | 20.194
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 699 | 20.224
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 700 | 20.255
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 701 | 20.286
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 702 | 20.316
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 703 | 20.347
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 704 | 20.378
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 705 | 20.408
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 706 | 20.439
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 707 | 20.469
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 708 | 20.5
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 709 | 20.531
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 710 | 20.561
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 711 | 20.592
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 712 | 20.623
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 713 | 20.653
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 714 | 20.684
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 715 | 20.715
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 716 | 20.745
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 717 | 20.776
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 718 | 20.806
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 719 | 20.837
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 720 | 20.868
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 721 | 20.898
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 722 | 20.929
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 723 | 20.96
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 724 | 20.99
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 725 | 21.021
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 726 | 21.052
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 727 | 21.082
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 728 | 21.113
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 729 | 21.143
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 730 | 21.174
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 731 | 21.205
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 732 | 21.235
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 733 | 21.266
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 734 | 21.297
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 735 | 21.327
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 736 | 21.358
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 737 | 21.389
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 738 | 21.419
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 739 | 21.45
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 740 | 21.48
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 741 | 21.511
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 742 | 21.542
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 743 | 21.572
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 744 | 21.603
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 745 | 21.634
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 746 | 21.664
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 747 | 21.695
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 748 | 21.726
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 749 | 21.756
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 750 | 21.787
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 751 | 21.817
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 752 | 21.848
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 753 | 21.879
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 754 | 21.909
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 755 | 21.94
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 756 | 21.971
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 757 | 22.001
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 758 | 22.032
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 759 | 22.063
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 760 | 22.093
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 761 | 22.124
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 762 | 22.154
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 763 | 22.185
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 764 | 22.216
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 765 | 22.246
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 766 | 22.277
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 767 | 22.308
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 768 | 22.338
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 769 | 22.369
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 770 | 22.4
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 771 | 22.43
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 772 | 22.461
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 773 | 22.492
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 774 | 22.522
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 775 | 22.553
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 776 | 22.584
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 777 | 22.614
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 778 | 22.645
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 779 | 22.676
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 780 | 22.706
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 781 | 22.737
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 782 | 22.768
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 783 | 22.798
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 784 | 22.829
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 785 | 22.86
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 786 | 22.89
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 787 | 22.921
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 788 | 22.952
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 789 | 22.982
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 790 | 23.013
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 791 | 23.044
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 792 | 23.074
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 793 | 23.105
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 794 | 23.136
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 795 | 23.166
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 796 | 23.197
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 797 | 23.228
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 798 | 23.258
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 799 | 23.289
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 800 | 23.32
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 801 | 23.35
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 802 | 23.381
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 803 | 23.412
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 804 | 23.442
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 805 | 23.473
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 806 | 23.504
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 807 | 23.535
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 808 | 23.565
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 809 | 23.596
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 810 | 23.627
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 811 | 23.657
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 812 | 23.688
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 813 | 23.719
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 814 | 23.749
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 815 | 23.78
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 816 | 23.811
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 817 | 23.842
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 818 | 23.872
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 819 | 23.903
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 820 | 23.934
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 821 | 23.964
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 822 | 23.995
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 823 | 24.026
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 824 | 24.057
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 825 | 24.087
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 826 | 24.118
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 827 | 24.149
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 828 | 24.18
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 829 | 24.21
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 830 | 24.241
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 831 | 24.272
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 832 | 24.303
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 833 | 24.333
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 834 | 24.364
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 835 | 24.395
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 836 | 24.426
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 837 | 24.456
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 838 | 24.487
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 839 | 24.518
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 840 | 24.549
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 841 | 24.579
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 842 | 24.61
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 843 | 24.641
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 844 | 24.672
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 845 | 24.702
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 846 | 24.733
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 847 | 24.764
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 848 | 24.795
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 849 | 24.826
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 850 | 24.856
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 851 | 24.887
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 852 | 24.918
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 853 | 24.949
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 854 | 24.979
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 855 | 25.01
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 856 | 25.041
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 857 | 25.072
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 858 | 25.103
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 859 | 25.134
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 860 | 25.164
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 861 | 25.195
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 862 | 25.226
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 863 | 25.257
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 864 | 25.288
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 865 | 25.318
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 866 | 25.349
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 867 | 25.38
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 868 | 25.411
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 869 | 25.442
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 870 | 25.473
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 871 | 25.504
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 872 | 25.534
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 873 | 25.565
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 874 | 25.596
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 875 | 25.627
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 876 | 25.658
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 877 | 25.689
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 878 | 25.72
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 879 | 25.75
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 880 | 25.781
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 881 | 25.812
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 882 | 25.843
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 883 | 25.874
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 884 | 25.905
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 885 | 25.936
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 886 | 25.967
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 887 | 25.998
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 888 | 26.029
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 889 | 26.059
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 890 | 26.09
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 891 | 26.121
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 892 | 26.152
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 893 | 26.183
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 894 | 26.214
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 895 | 26.245
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 896 | 26.276
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 897 | 26.307
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 898 | 26.338
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 899 | 26.369
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 900 | 26.4
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 901 | 26.431
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 902 | 26.462
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 903 | 26.493
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 904 | 26.524
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 905 | 26.555
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 906 | 26.586
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 907 | 26.617
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 908 | 26.648
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 909 | 26.679
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 910 | 26.71
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 911 | 26.741
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 912 | 26.772
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 913 | 26.803
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 914 | 26.834
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 915 | 26.865
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 916 | 26.896
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 917 | 26.927
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 918 | 26.958
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 919 | 26.989
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 920 | 27.02
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 921 | 27.051
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 922 | 27.082
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 923 | 27.113
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 924 | 27.144
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 925 | 27.175
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 926 | 27.206
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 927 | 27.237
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 928 | 27.268
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 929 | 27.299
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 930 | 27.33
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 931 | 27.362
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 932 | 27.393
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 933 | 27.424
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 934 | 27.455
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 935 | 27.486
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 936 | 27.517
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 937 | 27.548
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 938 | 27.579
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 939 | 27.61
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 940 | 27.642
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 941 | 27.673
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 942 | 27.704
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 943 | 27.735
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 944 | 27.766
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 945 | 27.797
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 946 | 27.829
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 947 | 27.86
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 948 | 27.891
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 949 | 27.922
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 950 | 27.953
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 951 | 27.985
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 952 | 28.016
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 953 | 28.047
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 954 | 28.078
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 955 | 28.109
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 956 | 28.141
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 957 | 28.172
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 958 | 28.203
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 959 | 28.234
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 960 | 28.266
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 961 | 28.297
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 962 | 28.328
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 963 | 28.359
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 964 | 28.391
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 965 | 28.422
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 966 | 28.453
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 967 | 28.485
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 968 | 28.516
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 969 | 28.547
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 970 | 28.579
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 971 | 28.61
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 972 | 28.641
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 973 | 28.672
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 974 | 28.704
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 975 | 28.735
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 976 | 28.767
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 977 | 28.798
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 978 | 28.829
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 979 | 28.861
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 980 | 28.892
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 981 | 28.923
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 982 | 28.955
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 983 | 28.986
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 984 | 29.018
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 985 | 29.049
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 986 | 29.08
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 987 | 29.112
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 988 | 29.143
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 989 | 29.175
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 990 | 29.206
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 991 | 29.238
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 992 | 29.269
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 993 | 29.301
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 994 | 29.332
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 995 | 29.363
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 996 | 29.395
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 997 | 29.426
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 998 | 29.458
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 999 | 29.489
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1000 | 29.521
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1001 | 29.552
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1002 | 29.584
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1003 | 29.616
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1004 | 29.647
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1005 | 29.679
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1006 | 29.71
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1007 | 29.742
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1008 | 29.773
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1009 | 29.805
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1010 | 29.836
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1011 | 29.868
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1012 | 29.9
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1013 | 29.931
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1014 | 29.963
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1015 | 29.995
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1016 | 30.026
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1017 | 30.058
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1018 | 30.089
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1019 | 30.121
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1020 | 30.153
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1021 | 30.184
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1022 | 30.216
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1023 | 30.248
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1024 | 30.279
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1025 | 30.311
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1026 | 30.343
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1027 | 30.375
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1028 | 30.406
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1029 | 30.438
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1030 | 30.47
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1031 | 30.502
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1032 | 30.533
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1033 | 30.565
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1034 | 30.597
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1035 | 30.629
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1036 | 30.66
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1037 | 30.692
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1038 | 30.724
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1039 | 30.756
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1040 | 30.788
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1041 | 30.819
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1042 | 30.851
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1043 | 30.883
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1044 | 30.915
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1045 | 30.947
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1046 | 30.979
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1047 | 31.011
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1048 | 31.043
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1049 | 31.074
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1050 | 31.106
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1051 | 31.138
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1052 | 31.17
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1053 | 31.202
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1054 | 31.234
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1055 | 31.266
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1056 | 31.298
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1057 | 31.33
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1058 | 31.362
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1059 | 31.394
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1060 | 31.426
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1061 | 31.458
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1062 | 31.49
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1063 | 31.522
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1064 | 31.554
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1065 | 31.586
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1066 | 31.618
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1067 | 31.65
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1068 | 31.682
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1069 | 31.714
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1070 | 31.746
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1071 | 31.778
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1072 | 31.811
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1073 | 31.843
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1074 | 31.875
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1075 | 31.907
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1076 | 31.939
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1077 | 31.971
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1078 | 32.003
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1079 | 32.035
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1080 | 32.068
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1081 | 32.1
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1082 | 32.132
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1083 | 32.164
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1084 | 32.196
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1085 | 32.229
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1086 | 32.261
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1087 | 32.293
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1088 | 32.325
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1089 | 32.358
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1090 | 32.39
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1091 | 32.422
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1092 | 32.455
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1093 | 32.487
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1094 | 32.519
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1095 | 32.551
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1096 | 32.584
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1097 | 32.616
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1098 | 32.648
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1099 | 32.681
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1100 | 32.713
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1101 | 32.746
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1102 | 32.778
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1103 | 32.81
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1104 | 32.843
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1105 | 32.875
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1106 | 32.908
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1107 | 32.94
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1108 | 32.973
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1109 | 33.005
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1110 | 33.037
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1111 | 33.07
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1112 | 33.102
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1113 | 33.135
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1114 | 33.167
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1115 | 33.2
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1116 | 33.232
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1117 | 33.265
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1118 | 33.298
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1119 | 33.33
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1120 | 33.363
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1121 | 33.395
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1122 | 33.428
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1123 | 33.46
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1124 | 33.493
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1125 | 33.526
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1126 | 33.558
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1127 | 33.591
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1128 | 33.624
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1129 | 33.656
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1130 | 33.689
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1131 | 33.722
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1132 | 33.754
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1133 | 33.787
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1134 | 33.82
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1135 | 33.853
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1136 | 33.885
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1137 | 33.918
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1138 | 33.951
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1139 | 33.984
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1140 | 34.016
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1141 | 34.049
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1142 | 34.082
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1143 | 34.115
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1144 | 34.148
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1145 | 34.18
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1146 | 34.213
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1147 | 34.246
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1148 | 34.279
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1149 | 34.312
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1150 | 34.345
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1151 | 34.378
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1152 | 34.411
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1153 | 34.444
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1154 | 34.476
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1155 | 34.509
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1156 | 34.542
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1157 | 34.575
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1158 | 34.608
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1159 | 34.641
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1160 | 34.674
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1161 | 34.707
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1162 | 34.74
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1163 | 34.773
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1164 | 34.806
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1165 | 34.84
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1166 | 34.873
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1167 | 34.906
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1168 | 34.939
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1169 | 34.972
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1170 | 35.005
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1171 | 35.038
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1172 | 35.071
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1173 | 35.104
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1174 | 35.138
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1175 | 35.171
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1176 | 35.204
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1177 | 35.237
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1178 | 35.27
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1179 | 35.304
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1180 | 35.337
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1181 | 35.37
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1182 | 35.403
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1183 | 35.437
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1184 | 35.47
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1185 | 35.503
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1186 | 35.536
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1187 | 35.57
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1188 | 35.603
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1189 | 35.636
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1190 | 35.67
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1191 | 35.703
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1192 | 35.736
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1193 | 35.77
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1194 | 35.803
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1195 | 35.837
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1196 | 35.87
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1197 | 35.903
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1198 | 35.937
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1199 | 35.97
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1200 | 36.004
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1201 | 36.037
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1202 | 36.071
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1203 | 36.104
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1204 | 36.138
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1205 | 36.171
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1206 | 36.205
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1207 | 36.238
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1208 | 36.272
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1209 | 36.305
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1210 | 36.339
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1211 | 36.373
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1212 | 36.406
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1213 | 36.44
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1214 | 36.473
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1215 | 36.507
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1216 | 36.541
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1217 | 36.574
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1218 | 36.608
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1219 | 36.642
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1220 | 36.675
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1221 | 36.709
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1222 | 36.743
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1223 | 36.777
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1224 | 36.81
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1225 | 36.844
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1226 | 36.878
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1227 | 36.912
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1228 | 36.945
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1229 | 36.979
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1230 | 37.013
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1231 | 37.047
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1232 | 37.081
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1233 | 37.114
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1234 | 37.148
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1235 | 37.182
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1236 | 37.216
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1237 | 37.25
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1238 | 37.284
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1239 | 37.318
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1240 | 37.352
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1241 | 37.386
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1242 | 37.42
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1243 | 37.454
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1244 | 37.488
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1245 | 37.522
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1246 | 37.556
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1247 | 37.59
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1248 | 37.624
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1249 | 37.658
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1250 | 37.692
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1251 | 37.726
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1252 | 37.76
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1253 | 37.794
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1254 | 37.828
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1255 | 37.862
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1256 | 37.896
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1257 | 37.93
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1258 | 37.964
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1259 | 37.999
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1260 | 38.033
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1261 | 38.067
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1262 | 38.101
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1263 | 38.135
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1264 | 38.169
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1265 | 38.204
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1266 | 38.238
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1267 | 38.272
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1268 | 38.306
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1269 | 38.341
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1270 | 38.375
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1271 | 38.409
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1272 | 38.444
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1273 | 38.478
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1274 | 38.512
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1275 | 38.546
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1276 | 38.581
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1277 | 38.615
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1278 | 38.65
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1279 | 38.684
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1280 | 38.718
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1281 | 38.753
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1282 | 38.787
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1283 | 38.822
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1284 | 38.856
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1285 | 38.89
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1286 | 38.925
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1287 | 38.959
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1288 | 38.994
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1289 | 39.028
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1290 | 39.063
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1291 | 39.097
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1292 | 39.132
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1293 | 39.166
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1294 | 39.201
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1295 | 39.235
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1296 | 39.27
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1297 | 39.305
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1298 | 39.339
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1299 | 39.374
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1300 | 39.408
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1301 | 39.443
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1302 | 39.478
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1303 | 39.512
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1304 | 39.547
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1305 | 39.582
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1306 | 39.616
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1307 | 39.651
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1308 | 39.686
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1309 | 39.72
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1310 | 39.755
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1311 | 39.79
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1312 | 39.825
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1313 | 39.859
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1314 | 39.894
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1315 | 39.929
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1316 | 39.964
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1317 | 39.998
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1318 | 40.033
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1319 | 40.068
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1320 | 40.103
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1321 | 40.138
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1322 | 40.173
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1323 | 40.207
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1324 | 40.242
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1325 | 40.277
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1326 | 40.312
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1327 | 40.347
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1328 | 40.382
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1329 | 40.417
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1330 | 40.452
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1331 | 40.487
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1332 | 40.522
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1333 | 40.556
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1334 | 40.591
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1335 | 40.626
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1336 | 40.661
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1337 | 40.696
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1338 | 40.731
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1339 | 40.766
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1340 | 40.801
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1341 | 40.836
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1342 | 40.872
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1343 | 40.907
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1344 | 40.942
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1345 | 40.977
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1346 | 41.012
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1347 | 41.047
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1348 | 41.082
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1349 | 41.117
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1350 | 41.152
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1351 | 41.187
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1352 | 41.222
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1353 | 41.258
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1354 | 41.293
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1355 | 41.328
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1356 | 41.363
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1357 | 41.398
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1358 | 41.433
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1359 | 41.469
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1360 | 41.504
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1361 | 41.539
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1362 | 41.574
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1363 | 41.61
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1364 | 41.645
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1365 | 41.68
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1366 | 41.715
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1367 | 41.751
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1368 | 41.786
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1369 | 41.821
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1370 | 41.856
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1371 | 41.892
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1372 | 41.927
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1373 | 41.962
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1374 | 41.998
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1375 | 42.033
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1376 | 42.068
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1377 | 42.104
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1378 | 42.139
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1379 | 42.174
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1380 | 42.21
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1381 | 42.245
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1382 | 42.281
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1383 | 42.316
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1384 | 42.351
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1385 | 42.387
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1386 | 42.422
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1387 | 42.458
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1388 | 42.493
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1389 | 42.528
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1390 | 42.564
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1391 | 42.599
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1392 | 42.635
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1393 | 42.67
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1394 | 42.706
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1395 | 42.741
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1396 | 42.777
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1397 | 42.812
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1398 | 42.848
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1399 | 42.883
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1400 | 42.919
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1401 | 42.954
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1402 | 42.99
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1403 | 43.025
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1404 | 43.061
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1405 | 43.096
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1406 | 43.132
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1407 | 43.167
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1408 | 43.203
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1409 | 43.239
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1410 | 43.274
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1411 | 43.31
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1412 | 43.346
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1413 | 43.381
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1414 | 43.417
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1415 | 43.452
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1416 | 43.488
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1417 | 43.524
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1418 | 43.559
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1419 | 43.595
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1420 | 43.631
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1421 | 43.667
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1422 | 43.702
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1423 | 43.738
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1424 | 43.774
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1425 | 43.809
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1426 | 43.845
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1427 | 43.881
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1428 | 43.917
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1429 | 43.953
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1430 | 43.988
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1431 | 44.024
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1432 | 44.06
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1433 | 44.096
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1434 | 44.131
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1435 | 44.167
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1436 | 44.203
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1437 | 44.239
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1438 | 44.275
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1439 | 44.31
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1440 | 44.346
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1441 | 44.382
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1442 | 44.418
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1443 | 44.454
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1444 | 44.49
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1445 | 44.525
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1446 | 44.561
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1447 | 44.597
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1448 | 44.633
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1449 | 44.669
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1450 | 44.705
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1451 | 44.741
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1452 | 44.777
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1453 | 44.812
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1454 | 44.848
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1455 | 44.884
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1456 | 44.92
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1457 | 44.956
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1458 | 44.992
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1459 | 45.028
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1460 | 45.064
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1461 | 45.099
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1462 | 45.135
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1463 | 45.171
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1464 | 45.207
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1465 | 45.243
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1466 | 45.279
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1467 | 45.315
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1468 | 45.351
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1469 | 45.387
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1470 | 45.423
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1471 | 45.458
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1472 | 45.494
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1473 | 45.53
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1474 | 45.566
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1475 | 45.602
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1476 | 45.638
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1477 | 45.674
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1478 | 45.71
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1479 | 45.746
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1480 | 45.782
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1481 | 45.818
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1482 | 45.853
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1483 | 45.889
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1484 | 45.925
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1485 | 45.961
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1486 | 45.997
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1487 | 46.033
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1488 | 46.069
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1489 | 46.105
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1490 | 46.141
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1491 | 46.177
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1492 | 46.212
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1493 | 46.248
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1494 | 46.284
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1495 | 46.32
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1496 | 46.356
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1497 | 46.392
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1498 | 46.428
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1499 | 46.464
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1500 | 46.5
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1501 | 46.535
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1502 | 46.571
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1503 | 46.607
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1504 | 46.643
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1505 | 46.679
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1506 | 46.715
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1507 | 46.751
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1508 | 46.786
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1509 | 46.822
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1510 | 46.858
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1511 | 46.894
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1512 | 46.93
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1513 | 46.966
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1514 | 47.001
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1515 | 47.037
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1516 | 47.073
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1517 | 47.109
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1518 | 47.145
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1519 | 47.181
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1520 | 47.216
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1521 | 47.252
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1522 | 47.288
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1523 | 47.324
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1524 | 47.359
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1525 | 47.395
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1526 | 47.431
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1527 | 47.467
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1528 | 47.503
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1529 | 47.538
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1530 | 47.574
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1531 | 47.61
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1532 | 47.646
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1533 | 47.681
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1534 | 47.717
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1535 | 47.753
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1536 | 47.788
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1537 | 47.824
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1538 | 47.86
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1539 | 47.896
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1540 | 47.931
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1541 | 47.967
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1542 | 48.003
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1543 | 48.038
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1544 | 48.074
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1545 | 48.11
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1546 | 48.145
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1547 | 48.181
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1548 | 48.217
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1549 | 48.252
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1550 | 48.288
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1551 | 48.324
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1552 | 48.359
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1553 | 48.395
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1554 | 48.43
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1555 | 48.466
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1556 | 48.502
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1557 | 48.537
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1558 | 48.573
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1559 | 48.608
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1560 | 48.644
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1561 | 48.679
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1562 | 48.715
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1563 | 48.75
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1564 | 48.786
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1565 | 48.822
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1566 | 48.857
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1567 | 48.893
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1568 | 48.928
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1569 | 48.964
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1570 | 48.999
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1571 | 49.034
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1572 | 49.07
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1573 | 49.105
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1574 | 49.141
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1575 | 49.176
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1576 | 49.212
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1577 | 49.247
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1578 | 49.283
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1579 | 49.318
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1580 | 49.353
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1581 | 49.389
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1582 | 49.424
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1583 | 49.46
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1584 | 49.495
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1585 | 49.53
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1586 | 49.566
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1587 | 49.601
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1588 | 49.636
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1589 | 49.672
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1590 | 49.707
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1591 | 49.742
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1592 | 49.778
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1593 | 49.813
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1594 | 49.848
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1595 | 49.883
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1596 | 49.919
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1597 | 49.954
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1598 | 49.989
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1599 | 50.024
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1600 | 50.06
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1601 | 50.095
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1602 | 50.13
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1603 | 50.165
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1604 | 50.2
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1605 | 50.235
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1606 | 50.271
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1607 | 50.306
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1608 | 50.341
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1609 | 50.376
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1610 | 50.411
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1611 | 50.446
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1612 | 50.481
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1613 | 50.517
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1614 | 50.552
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1615 | 50.587
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1616 | 50.622
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1617 | 50.657
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1618 | 50.692
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1619 | 50.727
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1620 | 50.762
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1621 | 50.797
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1622 | 50.832
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1623 | 50.867
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1624 | 50.902
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1625 | 50.937
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1626 | 50.972
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1627 | 51.007
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1628 | 51.042
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1629 | 51.077
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1630 | 51.112
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1631 | 51.147
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1632 | 51.181
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1633 | 51.216
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1634 | 51.251
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1635 | 51.286
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1636 | 51.321
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1637 | 51.356
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1638 | 51.391
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1639 | 51.425
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1640 | 51.46
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1641 | 51.495
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1642 | 51.53
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1643 | 51.565
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1644 | 51.599
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1645 | 51.634
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1646 | 51.669
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1647 | 51.704
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1648 | 51.738
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1649 | 51.773
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1650 | 51.808
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1651 | 51.843
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1652 | 51.877
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1653 | 51.912
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1654 | 51.947
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1655 | 51.981
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1656 | 52.016
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1657 | 52.051
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1658 | 52.085
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1659 | 52.12
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1660 | 52.154
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1661 | 52.189
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1662 | 52.224
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1663 | 52.258
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1664 | 52.293
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1665 | 52.327
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1666 | 52.362
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1667 | 52.396
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1668 | 52.431
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1669 | 52.465
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1670 | 52.5
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1671 | 52.534
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1672 | 52.569
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1673 | 52.603
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1674 | 52.638
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1675 | 52.672
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1676 | 52.707
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1677 | 52.741
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1678 | 52.776
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1679 | 52.81
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1680 | 52.844
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1681 | 52.879
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1682 | 52.913
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1683 | 52.947
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1684 | 52.982
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1685 | 53.016
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1686 | 53.05
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1687 | 53.085
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1688 | 53.119
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1689 | 53.153
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1690 | 53.188
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1691 | 53.222
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1692 | 53.256
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1693 | 53.29
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1694 | 53.325
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1695 | 53.359
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1696 | 53.393
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1697 | 53.427
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1698 | 53.462
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1699 | 53.496
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1700 | 53.53
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1701 | 53.564
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1702 | 53.598
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1703 | 53.632
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1704 | 53.667
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1705 | 53.701
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1706 | 53.735
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1707 | 53.769
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1708 | 53.803
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1709 | 53.837
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1710 | 53.871
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1711 | 53.905
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1712 | 53.939
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1713 | 53.973
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1714 | 54.007
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1715 | 54.041
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1716 | 54.075
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1717 | 54.109
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1718 | 54.143
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1719 | 54.177
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1720 | 54.211
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1721 | 54.245
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1722 | 54.279
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1723 | 54.313
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1724 | 54.347
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1725 | 54.381
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1726 | 54.415
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1727 | 54.449
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1728 | 54.483
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1729 | 54.516
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1730 | 54.55
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1731 | 54.584
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1732 | 54.618
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1733 | 54.652
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1734 | 54.686
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1735 | 54.719
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1736 | 54.753
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1737 | 54.787
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1738 | 54.821
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1739 | 54.855
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1740 | 54.888
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1741 | 54.922
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1742 | 54.956
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1743 | 54.99
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1744 | 55.023
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1745 | 55.057
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1746 | 55.091
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1747 | 55.124
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1748 | 55.158
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1749 | 55.192
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1750 | 55.225
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1751 | 55.259
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1752 | 55.293
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1753 | 55.326
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1754 | 55.36
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1755 | 55.393
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1756 | 55.427
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1757 | 55.461
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1758 | 55.494
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1759 | 55.528
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1760 | 55.561
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1761 | 55.595
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1762 | 55.628
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1763 | 55.662
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1764 | 55.695
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1765 | 55.729
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1766 | 55.762
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1767 | 55.796
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1768 | 55.829
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1769 | 55.863
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1770 | 55.896
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1771 | 55.93
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1772 | 55.963
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1773 | 55.997
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1774 | 56.03
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1775 | 56.063
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1776 | 56.097
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1777 | 56.13
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1778 | 56.164
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1779 | 56.197
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1780 | 56.23
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1781 | 56.264
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1782 | 56.297
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1783 | 56.33
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1784 | 56.364
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1785 | 56.397
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1786 | 56.43
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1787 | 56.464
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1788 | 56.497
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1789 | 56.53
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1790 | 56.564
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1791 | 56.597
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1792 | 56.63
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1793 | 56.663
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1794 | 56.697
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1795 | 56.73
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1796 | 56.763
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1797 | 56.796
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1798 | 56.829
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1799 | 56.863
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1800 | 56.896
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1801 | 56.929
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1802 | 56.962
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1803 | 56.995
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1804 | 57.028
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1805 | 57.062
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1806 | 57.095
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1807 | 57.128
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1808 | 57.161
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1809 | 57.194
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1810 | 57.227
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1811 | 57.26
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1812 | 57.293
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1813 | 57.326
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1814 | 57.36
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1815 | 57.393
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1816 | 57.426
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1817 | 57.459
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1818 | 57.492
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1819 | 57.525
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1820 | 57.558
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1821 | 57.591
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1822 | 57.624
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1823 | 57.657
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1824 | 57.69
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1825 | 57.723
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1826 | 57.756
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1827 | 57.789
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1828 | 57.822
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1829 | 57.855
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1830 | 57.888
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1831 | 57.92
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1832 | 57.953
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1833 | 57.986
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1834 | 58.019
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1835 | 58.052
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1836 | 58.085
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1837 | 58.118
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1838 | 58.151
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1839 | 58.184
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1840 | 58.217
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1841 | 58.249
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1842 | 58.282
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1843 | 58.315
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1844 | 58.348
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1845 | 58.381
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1846 | 58.414
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1847 | 58.446
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1848 | 58.479
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1849 | 58.512
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1850 | 58.545
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1851 | 58.578
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1852 | 58.61
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1853 | 58.643
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1854 | 58.676
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1855 | 58.709
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1856 | 58.741
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1857 | 58.774
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1858 | 58.807
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1859 | 58.84
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1860 | 58.872
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1861 | 58.905
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1862 | 58.938
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1863 | 58.971
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1864 | 59.003
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1865 | 59.036
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1866 | 59.069
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1867 | 59.101
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1868 | 59.134
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1869 | 59.167
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1870 | 59.199
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1871 | 59.232
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1872 | 59.265
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1873 | 59.297
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1874 | 59.33
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1875 | 59.363
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1876 | 59.395
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1877 | 59.428
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1878 | 59.46
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1879 | 59.493
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1880 | 59.526
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1881 | 59.558
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1882 | 59.591
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1883 | 59.623
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1884 | 59.656
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1885 | 59.689
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1886 | 59.721
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1887 | 59.754
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1888 | 59.786
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1889 | 59.819
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1890 | 59.851
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1891 | 59.884
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1892 | 59.916
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1893 | 59.949
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1894 | 59.982
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1895 | 60.014
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1896 | 60.047
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1897 | 60.079
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1898 | 60.112
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1899 | 60.144
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1900 | 60.177
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1901 | 60.209
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1902 | 60.242
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1903 | 60.274
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1904 | 60.307
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1905 | 60.339
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1906 | 60.371
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1907 | 60.404
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1908 | 60.436
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1909 | 60.469
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1910 | 60.501
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1911 | 60.534
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1912 | 60.566
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1913 | 60.599
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1914 | 60.631
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1915 | 60.663
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1916 | 60.696
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1917 | 60.728
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1918 | 60.761
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1919 | 60.793
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1920 | 60.826
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1921 | 60.858
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1922 | 60.89
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1923 | 60.923
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1924 | 60.955
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1925 | 60.987
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1926 | 61.02
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1927 | 61.052
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1928 | 61.085
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1929 | 61.117
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1930 | 61.149
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1931 | 61.182
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1932 | 61.214
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1933 | 61.246
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1934 | 61.279
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1935 | 61.311
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1936 | 61.343
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1937 | 61.376
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1938 | 61.408
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1939 | 61.44
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1940 | 61.473
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1941 | 61.505
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1942 | 61.537
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1943 | 61.57
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1944 | 61.602
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1945 | 61.634
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1946 | 61.667
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1947 | 61.699
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1948 | 61.731
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1949 | 61.763
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1950 | 61.796
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1951 | 61.828
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1952 | 61.86
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1953 | 61.893
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1954 | 61.925
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1955 | 61.957
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1956 | 61.989
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1957 | 62.022
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1958 | 62.054
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1959 | 62.086
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1960 | 62.118
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1961 | 62.151
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1962 | 62.183
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1963 | 62.215
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1964 | 62.247
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1965 | 62.28
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1966 | 62.312
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1967 | 62.344
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1968 | 62.376
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1969 | 62.409
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1970 | 62.441
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1971 | 62.473
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1972 | 62.505
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1973 | 62.537
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1974 | 62.57
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1975 | 62.602
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1976 | 62.634
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1977 | 62.666
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1978 | 62.699
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1979 | 62.731
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1980 | 62.763
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1981 | 62.795
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1982 | 62.827
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1983 | 62.86
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1984 | 62.892
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1985 | 62.924
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1986 | 62.956
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1987 | 62.988
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1988 | 63.02
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1989 | 63.053
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1990 | 63.085
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1991 | 63.117
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1992 | 63.149
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1993 | 63.181
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1994 | 63.214
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1995 | 63.246
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1996 | 63.278
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1997 | 63.31
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1998 | 63.342
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 1999 | 63.374
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2000 | 63.406
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2001 | 63.439
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2002 | 63.471
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2003 | 63.503
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2004 | 63.535
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2005 | 63.567
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2006 | 63.599
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2007 | 63.632
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2008 | 63.664
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2009 | 63.696
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2010 | 63.728
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2011 | 63.76
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2012 | 63.792
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2013 | 63.824
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2014 | 63.856
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2015 | 63.889
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2016 | 63.921
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2017 | 63.953
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2018 | 63.985
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2019 | 64.017
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2020 | 64.049
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2021 | 64.081
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2022 | 64.113
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2023 | 64.146
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2024 | 64.178
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2025 | 64.21
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2026 | 64.242
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2027 | 64.274
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2028 | 64.306
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2029 | 64.338
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2030 | 64.37
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2031 | 64.402
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2032 | 64.435
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2033 | 64.467
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2034 | 64.499
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2035 | 64.531
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2036 | 64.563
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2037 | 64.595
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2038 | 64.627
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2039 | 64.659
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2040 | 64.691
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2041 | 64.723
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2042 | 64.756
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2043 | 64.788
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2044 | 64.82
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2045 | 64.852
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2046 | 64.884
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2047 | 64.916
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2048 | 64.948
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2049 | 64.98
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2050 | 65.012
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2051 | 65.044
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2052 | 65.076
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2053 | 65.109
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2054 | 65.141
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2055 | 65.173
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2056 | 65.205
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2057 | 65.237
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2058 | 65.269
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2059 | 65.301
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2060 | 65.333
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2061 | 65.365
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2062 | 65.397
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2063 | 65.429
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2064 | 65.461
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2065 | 65.493
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2066 | 65.525
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2067 | 65.557
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2068 | 65.59
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2069 | 65.622
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2070 | 65.654
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2071 | 65.686
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2072 | 65.718
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2073 | 65.75
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2074 | 65.782
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2075 | 65.814
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2076 | 65.846
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2077 | 65.878
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2078 | 65.91
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2079 | 65.942
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2080 | 65.974
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2081 | 66.006
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2082 | 66.038
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2083 | 66.07
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2084 | 66.102
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2085 | 66.134
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2086 | 66.166
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2087 | 66.199
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2088 | 66.231
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2089 | 66.263
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2090 | 66.295
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2091 | 66.327
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2092 | 66.359
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2093 | 66.391
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2094 | 66.423
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2095 | 66.455
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2096 | 66.487
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2097 | 66.519
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2098 | 66.551
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2099 | 66.583
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2100 | 66.615
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2101 | 66.647
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2102 | 66.679
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2103 | 66.711
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2104 | 66.743
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2105 | 66.775
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2106 | 66.807
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2107 | 66.839
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2108 | 66.871
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2109 | 66.903
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2110 | 66.935
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2111 | 66.967
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2112 | 66.999
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2113 | 67.031
0 | 32 | 378 | -0.028 | 0 | 31 | 32 | 2114 | 67.063
[7 empty] | 2115 | 67.095
[7 empty] | 2116 | 67.127
[7 empty] | 2117 | 67.159
[7 empty] | 2118 | 67.191
[7 empty] | 2119 | 67.223
[7 empty] | 2120 | 67.255
[7 empty] | 2121 | 67.287
[7 empty] | 2122 | 67.319
[7 empty] | 2123 | 67.351
[7 empty] | 2124 | 67.383
[7 empty] | 2125 | 67.415
[7 empty] | 2126 | 67.447
[7 empty] | 2127 | 67.479
[7 empty] | 2128 | 67.511
[7 empty] | 2129 | 67.543
[7 empty] | 2130 | 67.575
[7 empty] | 2131 | 67.607
[7 empty] | 2132 | 67.639
[7 empty] | 2133 | 67.671
[7 empty] | 2134 | 67.703
[7 empty] | 2135 | 67.735
[7 empty] | 2136 | 67.767
[7 empty] | 2137 | 67.799
[7 empty] | 2138 | 67.831
[7 empty] | 2139 | 67.863
[7 empty] | 2140 | 67.895
[7 empty] | 2141 | 67.927
[7 empty] | 2142 | 67.959
[7 empty] | 2143 | 67.991
[7 empty] | 2144 | 68.023
[7 empty] | 2145 | 68.055
[7 empty] | 2146 | 68.087
[7 empty] | 2147 | 68.119
[7 empty] | 2148 | 68.15
[7 empty] | 2149 | 68.182
[7 empty] | 2150 | 68.214
[7 empty] | 2151 | 68.246
[7 empty] | 2152 | 68.278
[7 empty] | 2153 | 68.31
[7 empty] | 2154 | 68.342
[7 empty] | 2155 | 68.374
[7 empty] | 2156 | 68.406
[7 empty] | 2157 | 68.438
[7 empty] | 2158 | 68.47
[7 empty] | 2159 | 68.502
[7 empty] | 2160 | 68.534
[7 empty] | 2161 | 68.566
[7 empty] | 2162 | 68.597
[7 empty] | 2163 | 68.629
[7 empty] | 2164 | 68.661
[7 empty] | 2165 | 68.693
[7 empty] | 2166 | 68.725
[7 empty] | 2167 | 68.757
[7 empty] | 2168 | 68.789
[7 empty] | 2169 | 68.821
[7 empty] | 2170 | 68.853
[7 empty] | 2171 | 68.884
[7 empty] | 2172 | 68.916
[7 empty] | 2173 | 68.948
[7 empty] | 2174 | 68.98
[7 empty] | 2175 | 69.012
[7 empty] | 2176 | 69.044
[7 empty] | 2177 | 69.076
[7 empty] | 2178 | 69.108
[7 empty] | 2179 | 69.139
[7 empty] | 2180 | 69.171
[7 empty] | 2181 | 69.203
[7 empty] | 2182 | 69.235
[7 empty] | 2183 | 69.267
[7 empty] | 2184 | 69.299
[7 empty] | 2185 | 69.33
[7 empty] | 2186 | 69.362
[7 empty] | 2187 | 69.394
[7 empty] | 2188 | 69.426
[7 empty] | 2189 | 69.458
[7 empty] | 2190 | 69.49
[7 empty] | 2191 | 69.521
[7 empty] | 2192 | 69.553

## Type K
[7 empty] | Temp(F) | mV
mV | Temp(F) | [5 empty] | -454 | -6.458
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -453 | -6.457
17.526 | 800 | 1254 | 17.502 | 17.526 | 799 | 800 | -452 | -6.457
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -451 | -6.456
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -450 | -6.456
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -449 | -6.455
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -448 | -6.454
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -447 | -6.454
5 | 251.5217 | 706 | 4.988 | 5.011 | 251 | 252 | -446 | -6.453
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -445 | -6.452
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -444 | -6.451
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -443 | -6.45
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -442 | -6.449
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -441 | -6.448
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -440 | -6.446
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -439 | -6.445
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -438 | -6.444
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -437 | -6.443
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -436 | -6.441
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -435 | -6.44
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -434 | -6.438
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -433 | -6.436
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -432 | -6.435
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -431 | -6.433
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -430 | -6.431
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -429 | -6.429
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -428 | -6.427
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -427 | -6.425
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -426 | -6.423
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -425 | -6.421
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -424 | -6.419
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -423 | -6.416
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -422 | -6.414
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -421 | -6.411
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -420 | -6.409
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -419 | -6.406
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -418 | -6.404
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -417 | -6.401
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -416 | -6.398
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -415 | -6.395
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -414 | -6.392
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -413 | -6.389
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -412 | -6.386
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -411 | -6.383
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -410 | -6.38
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -409 | -6.377
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -408 | -6.373
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -407 | -6.37
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -406 | -6.366
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -405 | -6.363
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -404 | -6.359
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -403 | -6.355
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -402 | -6.352
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -401 | -6.348
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -400 | -6.344
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -399 | -6.34
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -398 | -6.336
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -397 | -6.332
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -396 | -6.328
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -395 | -6.323
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -394 | -6.319
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -393 | -6.315
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -392 | -6.31
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -391 | -6.306
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -390 | -6.301
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -389 | -6.296
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -388 | -6.292
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -387 | -6.287
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -386 | -6.282
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -385 | -6.277
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -384 | -6.272
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -383 | -6.267
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -382 | -6.262
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -381 | -6.257
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -380 | -6.251
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -379 | -6.246
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -378 | -6.241
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -377 | -6.235
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -376 | -6.23
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -375 | -6.224
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -374 | -6.218
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -373 | -6.213
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -372 | -6.207
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -371 | -6.201
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -370 | -6.195
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -369 | -6.189
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -368 | -6.183
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -367 | -6.177
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -366 | -6.171
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -365 | -6.165
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -364 | -6.158
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -363 | -6.152
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -362 | -6.146
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -361 | -6.139
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -360 | -6.133
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -359 | -6.126
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -358 | -6.119
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -357 | -6.113
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -356 | -6.106
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -355 | -6.099
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -354 | -6.092
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -353 | -6.085
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -352 | -6.078
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -351 | -6.071
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -350 | -6.064
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -349 | -6.057
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -348 | -6.049
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -347 | -6.042
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -346 | -6.035
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -345 | -6.027
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -344 | -6.02
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -343 | -6.012
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -342 | -6.004
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -341 | -5.997
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -340 | -5.989
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -339 | -5.981
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -338 | -5.973
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -337 | -5.965
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -336 | -5.957
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -335 | -5.949
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -334 | -5.941
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -333 | -5.933
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -332 | -5.925
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -331 | -5.917
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -330 | -5.908
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -329 | -5.9
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -328 | -5.891
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -327 | -5.883
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -326 | -5.874
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -325 | -5.866
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -324 | -5.857
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -323 | -5.848
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -322 | -5.84
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -321 | -5.831
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -320 | -5.822
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -319 | -5.813
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -318 | -5.804
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -317 | -5.795
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -316 | -5.786
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -315 | -5.776
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -314 | -5.767
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -313 | -5.758
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -312 | -5.749
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -311 | -5.739
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -310 | -5.73
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -309 | -5.72
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -308 | -5.711
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -307 | -5.701
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -306 | -5.691
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -305 | -5.682
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -304 | -5.672
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -303 | -5.662
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -302 | -5.652
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -301 | -5.642
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -300 | -5.632
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -299 | -5.622
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -298 | -5.612
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -297 | -5.602
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -296 | -5.592
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -295 | -5.581
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -294 | -5.571
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -293 | -5.561
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -292 | -5.55
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -291 | -5.54
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -290 | -5.529
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -289 | -5.519
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -288 | -5.508
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -287 | -5.497
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -286 | -5.487
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -285 | -5.476
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -284 | -5.465
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -283 | -5.454
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -282 | -5.443
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -281 | -5.432
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -280 | -5.421
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -279 | -5.41
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -278 | -5.399
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -277 | -5.388
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -276 | -5.377
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -275 | -5.365
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -274 | -5.354
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -273 | -5.343
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -272 | -5.331
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -271 | -5.32
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -270 | -5.308
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -269 | -5.296
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -268 | -5.285
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -267 | -5.273
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -266 | -5.261
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -265 | -5.25
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -264 | -5.238
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -263 | -5.226
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -262 | -5.214
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -261 | -5.202
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -260 | -5.19
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -259 | -5.178
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -258 | -5.166
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -257 | -5.153
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -256 | -5.141
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -255 | -5.129
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -254 | -5.117
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -253 | -5.104
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -252 | -5.092
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -251 | -5.079
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -250 | -5.067
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -249 | -5.054
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -248 | -5.042
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -247 | -5.029
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -246 | -5.016
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -245 | -5.003
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -244 | -4.991
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -243 | -4.978
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -242 | -4.965
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -241 | -4.952
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -240 | -4.939
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -239 | -4.926
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -238 | -4.913
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -237 | -4.9
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -236 | -4.886
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -235 | -4.873
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -234 | -4.86
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -233 | -4.847
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -232 | -4.833
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -231 | -4.82
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -230 | -4.806
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -229 | -4.793
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -228 | -4.779
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -227 | -4.766
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -226 | -4.752
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -225 | -4.738
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -224 | -4.724
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -223 | -4.711
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -222 | -4.697
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -221 | -4.683
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -220 | -4.669
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -219 | -4.655
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -218 | -4.641
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -217 | -4.627
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -216 | -4.613
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -215 | -4.599
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -214 | -4.584
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -213 | -4.57
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -212 | -4.556
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -211 | -4.542
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -210 | -4.527
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -209 | -4.513
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -208 | -4.498
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -207 | -4.484
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -206 | -4.469
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -205 | -4.455
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -204 | -4.44
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -203 | -4.425
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -202 | -4.411
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -201 | -4.396
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -200 | -4.381
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -199 | -4.366
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -198 | -4.351
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -197 | -4.336
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -196 | -4.321
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -195 | -4.306
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -194 | -4.291
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -193 | -4.276
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -192 | -4.261
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -191 | -4.246
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -190 | -4.231
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -189 | -4.215
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -188 | -4.2
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -187 | -4.185
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -186 | -4.169
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -185 | -4.154
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -184 | -4.138
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -183 | -4.123
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -182 | -4.107
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -181 | -4.091
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -180 | -4.076
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -179 | -4.06
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -178 | -4.044
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -177 | -4.029
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -176 | -4.013
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -175 | -3.997
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -174 | -3.981
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -173 | -3.965
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -172 | -3.949
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -171 | -3.933
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -170 | -3.917
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -169 | -3.901
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -168 | -3.885
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -167 | -3.869
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -166 | -3.852
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -165 | -3.836
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -164 | -3.82
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -163 | -3.803
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -162 | -3.787
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -161 | -3.771
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -160 | -3.754
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -159 | -3.738
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -158 | -3.721
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -157 | -3.705
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -156 | -3.688
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -155 | -3.671
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -154 | -3.655
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -153 | -3.638
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -152 | -3.621
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -151 | -3.604
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -150 | -3.587
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -149 | -3.571
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -148 | -3.554
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -147 | -3.537
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -146 | -3.52
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -145 | -3.503
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -144 | -3.486
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -143 | -3.468
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -142 | -3.451
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -141 | -3.434
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -140 | -3.417
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -139 | -3.4
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -138 | -3.382
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -137 | -3.365
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -136 | -3.348
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -135 | -3.33
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -134 | -3.313
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -133 | -3.295
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -132 | -3.278
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -131 | -3.26
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -130 | -3.243
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -129 | -3.225
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -128 | -3.207
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -127 | -3.19
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -126 | -3.172
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -125 | -3.154
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -124 | -3.136
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -123 | -3.119
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -122 | -3.101
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -121 | -3.083
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -120 | -3.065
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -119 | -3.047
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -118 | -3.029
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -117 | -3.011
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -116 | -2.993
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -115 | -2.975
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -114 | -2.957
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -113 | -2.938
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -112 | -2.92
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -111 | -2.902
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -110 | -2.884
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -109 | -2.865
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -108 | -2.847
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -107 | -2.829
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -106 | -2.81
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -105 | -2.792
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -104 | -2.773
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -103 | -2.755
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -102 | -2.736
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -101 | -2.718
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -100 | -2.699
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -99 | -2.68
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -98 | -2.662
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -97 | -2.643
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -96 | -2.624
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -95 | -2.605
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -94 | -2.587
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -93 | -2.568
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -92 | -2.549
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -91 | -2.53
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -90 | -2.511
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -89 | -2.492
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -88 | -2.473
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -87 | -2.454
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -86 | -2.435
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -85 | -2.416
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -84 | -2.397
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -83 | -2.378
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -82 | -2.359
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -81 | -2.339
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -80 | -2.32
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -79 | -2.301
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -78 | -2.282
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -77 | -2.262
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -76 | -2.243
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -75 | -2.223
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -74 | -2.204
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -73 | -2.185
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -72 | -2.165
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -71 | -2.146
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -70 | -2.126
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -69 | -2.106
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -68 | -2.087
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -67 | -2.067
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -66 | -2.048
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -65 | -2.028
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -64 | -2.008
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -63 | -1.988
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -62 | -1.969
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -61 | -1.949
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -60 | -1.929
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -59 | -1.909
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -58 | -1.889
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -57 | -1.869
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -56 | -1.85
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -55 | -1.83
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -54 | -1.81
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -53 | -1.79
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -52 | -1.77
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -51 | -1.749
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -50 | -1.729
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -49 | -1.709
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -48 | -1.689
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -47 | -1.669
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -46 | -1.649
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -45 | -1.628
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -44 | -1.608
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -43 | -1.588
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -42 | -1.568
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -41 | -1.547
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -40 | -1.527
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -39 | -1.507
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -38 | -1.486
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -37 | -1.466
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -36 | -1.445
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -35 | -1.425
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -34 | -1.404
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -33 | -1.384
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -32 | -1.363
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -31 | -1.343
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -30 | -1.322
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -29 | -1.301
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -28 | -1.281
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -27 | -1.26
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -26 | -1.239
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -25 | -1.218
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -24 | -1.198
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -23 | -1.177
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -22 | -1.156
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -21 | -1.135
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -20 | -1.114
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -19 | -1.094
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -18 | -1.073
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -17 | -1.052
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -16 | -1.031
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -15 | -1.01
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -14 | -0.989
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -13 | -0.968
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -12 | -0.947
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -11 | -0.926
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -10 | -0.905
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -9 | -0.883
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -8 | -0.862
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -7 | -0.841
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -6 | -0.82
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -5 | -0.799
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -4 | -0.778
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -3 | -0.756
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -2 | -0.735
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -1 | -0.714
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 0 | -0.692
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1 | -0.671
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 2 | -0.65
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 3 | -0.628
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 4 | -0.607
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 5 | -0.586
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 6 | -0.564
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 7 | -0.543
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 8 | -0.521
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 9 | -0.5
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 10 | -0.478
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 11 | -0.457
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 12 | -0.435
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 13 | -0.413
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 14 | -0.392
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 15 | -0.37
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 16 | -0.349
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 17 | -0.327
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 18 | -0.305
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 19 | -0.284
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 20 | -0.262
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 21 | -0.24
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 22 | -0.218
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 23 | -0.197
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 24 | -0.175
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 25 | -0.153
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 26 | -0.131
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 27 | -0.109
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 28 | -0.088
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 29 | -0.066
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 30 | -0.044
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 31 | -0.022
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 32 | 0
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 33 | 0.022
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 34 | 0.044
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 35 | 0.066
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 36 | 0.088
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 37 | 0.11
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 38 | 0.132
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 39 | 0.154
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 40 | 0.176
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 41 | 0.198
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 42 | 0.22
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 43 | 0.242
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 44 | 0.264
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 45 | 0.286
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 46 | 0.308
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 47 | 0.33
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 48 | 0.353
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 49 | 0.375
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 50 | 0.397
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 51 | 0.419
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 52 | 0.441
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 53 | 0.463
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 54 | 0.486
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 55 | 0.508
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 56 | 0.53
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 57 | 0.552
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 58 | 0.575
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 59 | 0.597
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 60 | 0.619
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 61 | 0.642
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 62 | 0.664
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 63 | 0.686
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 64 | 0.709
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 65 | 0.731
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 66 | 0.753
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 67 | 0.776
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 68 | 0.798
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 69 | 0.821
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 70 | 0.843
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 71 | 0.865
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 72 | 0.888
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 73 | 0.91
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 74 | 0.933
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 75 | 0.955
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 76 | 0.978
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 77 | 1
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 78 | 1.023
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 79 | 1.045
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 80 | 1.068
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 81 | 1.09
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 82 | 1.113
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 83 | 1.136
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 84 | 1.158
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 85 | 1.181
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 86 | 1.203
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 87 | 1.226
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 88 | 1.249
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 89 | 1.271
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 90 | 1.294
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 91 | 1.316
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 92 | 1.339
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 93 | 1.362
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 94 | 1.384
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 95 | 1.407
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 96 | 1.43
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 97 | 1.453
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 98 | 1.475
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 99 | 1.498
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 100 | 1.521
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 101 | 1.543
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 102 | 1.566
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 103 | 1.589
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 104 | 1.612
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 105 | 1.635
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 106 | 1.657
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 107 | 1.68
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 108 | 1.703
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 109 | 1.726
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 110 | 1.749
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 111 | 1.771
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 112 | 1.794
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 113 | 1.817
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 114 | 1.84
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 115 | 1.863
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 116 | 1.886
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 117 | 1.909
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 118 | 1.931
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 119 | 1.954
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 120 | 1.977
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 121 | 2
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 122 | 2.023
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 123 | 2.046
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 124 | 2.069
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 125 | 2.092
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 126 | 2.115
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 127 | 2.138
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 128 | 2.161
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 129 | 2.184
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 130 | 2.207
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 131 | 2.23
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 132 | 2.253
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 133 | 2.276
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 134 | 2.298
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 135 | 2.321
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 136 | 2.344
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 137 | 2.367
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 138 | 2.39
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 139 | 2.413
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 140 | 2.436
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 141 | 2.459
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 142 | 2.483
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 143 | 2.506
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 144 | 2.529
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 145 | 2.552
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 146 | 2.575
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 147 | 2.598
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 148 | 2.621
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 149 | 2.644
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 150 | 2.667
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 151 | 2.69
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 152 | 2.713
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 153 | 2.736
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 154 | 2.759
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 155 | 2.782
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 156 | 2.805
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 157 | 2.828
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 158 | 2.851
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 159 | 2.874
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 160 | 2.897
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 161 | 2.92
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 162 | 2.944
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 163 | 2.967
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 164 | 2.99
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 165 | 3.013
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 166 | 3.036
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 167 | 3.059
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 168 | 3.082
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 169 | 3.105
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 170 | 3.128
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 171 | 3.151
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 172 | 3.174
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 173 | 3.197
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 174 | 3.22
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 175 | 3.244
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 176 | 3.267
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 177 | 3.29
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 178 | 3.313
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 179 | 3.336
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 180 | 3.359
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 181 | 3.382
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 182 | 3.405
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 183 | 3.428
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 184 | 3.451
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 185 | 3.474
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 186 | 3.497
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 187 | 3.52
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 188 | 3.544
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 189 | 3.567
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 190 | 3.59
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 191 | 3.613
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 192 | 3.636
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 193 | 3.659
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 194 | 3.682
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 195 | 3.705
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 196 | 3.728
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 197 | 3.751
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 198 | 3.774
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 199 | 3.797
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 200 | 3.82
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 201 | 3.843
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 202 | 3.866
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 203 | 3.889
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 204 | 3.912
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 205 | 3.935
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 206 | 3.958
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 207 | 3.981
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 208 | 4.004
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 209 | 4.027
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 210 | 4.05
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 211 | 4.073
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 212 | 4.096
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 213 | 4.119
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 214 | 4.142
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 215 | 4.165
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 216 | 4.188
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 217 | 4.211
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 218 | 4.234
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 219 | 4.257
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 220 | 4.28
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 221 | 4.303
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 222 | 4.326
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 223 | 4.349
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 224 | 4.372
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 225 | 4.395
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 226 | 4.417
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 227 | 4.44
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 228 | 4.463
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 229 | 4.486
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 230 | 4.509
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 231 | 4.532
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 232 | 4.555
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 233 | 4.578
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 234 | 4.601
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 235 | 4.623
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 236 | 4.646
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 237 | 4.669
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 238 | 4.692
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 239 | 4.715
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 240 | 4.738
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 241 | 4.76
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 242 | 4.783
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 243 | 4.806
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 244 | 4.829
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 245 | 4.852
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 246 | 4.874
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 247 | 4.897
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 248 | 4.92
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 249 | 4.943
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 250 | 4.965
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 251 | 4.988
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 252 | 5.011
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 253 | 5.034
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 254 | 5.056
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 255 | 5.079
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 256 | 5.102
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 257 | 5.124
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 258 | 5.147
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 259 | 5.17
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 260 | 5.192
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 261 | 5.215
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 262 | 5.238
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 263 | 5.26
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 264 | 5.283
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 265 | 5.306
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 266 | 5.328
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 267 | 5.351
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 268 | 5.374
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 269 | 5.396
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 270 | 5.419
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 271 | 5.441
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 272 | 5.464
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 273 | 5.487
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 274 | 5.509
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 275 | 5.532
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 276 | 5.554
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 277 | 5.577
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 278 | 5.599
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 279 | 5.622
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 280 | 5.644
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 281 | 5.667
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 282 | 5.69
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 283 | 5.712
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 284 | 5.735
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 285 | 5.757
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 286 | 5.779
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 287 | 5.802
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 288 | 5.824
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 289 | 5.847
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 290 | 5.869
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 291 | 5.892
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 292 | 5.914
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 293 | 5.937
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 294 | 5.959
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 295 | 5.982
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 296 | 6.004
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 297 | 6.026
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 298 | 6.049
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 299 | 6.071
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 300 | 6.094
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 301 | 6.116
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 302 | 6.138
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 303 | 6.161
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 304 | 6.183
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 305 | 6.205
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 306 | 6.228
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 307 | 6.25
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 308 | 6.272
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 309 | 6.295
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 310 | 6.317
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 311 | 6.339
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 312 | 6.362
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 313 | 6.384
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 314 | 6.406
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 315 | 6.429
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 316 | 6.451
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 317 | 6.473
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 318 | 6.496
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 319 | 6.518
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 320 | 6.54
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 321 | 6.562
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 322 | 6.585
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 323 | 6.607
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 324 | 6.629
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 325 | 6.652
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 326 | 6.674
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 327 | 6.696
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 328 | 6.718
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 329 | 6.741
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 330 | 6.763
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 331 | 6.785
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 332 | 6.807
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 333 | 6.829
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 334 | 6.852
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 335 | 6.874
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 336 | 6.896
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 337 | 6.918
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 338 | 6.941
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 339 | 6.963
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 340 | 6.985
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 341 | 7.007
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 342 | 7.029
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 343 | 7.052
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 344 | 7.074
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 345 | 7.096
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 346 | 7.118
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 347 | 7.14
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 348 | 7.163
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 349 | 7.185
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 350 | 7.207
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 351 | 7.229
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 352 | 7.251
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 353 | 7.273
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 354 | 7.296
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 355 | 7.318
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 356 | 7.34
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 357 | 7.362
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 358 | 7.384
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 359 | 7.407
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 360 | 7.429
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 361 | 7.451
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 362 | 7.473
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 363 | 7.495
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 364 | 7.517
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 365 | 7.54
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 366 | 7.562
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 367 | 7.584
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 368 | 7.606
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 369 | 7.628
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 370 | 7.65
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 371 | 7.673
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 372 | 7.695
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 373 | 7.717
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 374 | 7.739
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 375 | 7.761
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 376 | 7.783
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 377 | 7.806
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 378 | 7.828
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 379 | 7.85
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 380 | 7.872
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 381 | 7.894
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 382 | 7.917
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 383 | 7.939
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 384 | 7.961
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 385 | 7.983
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 386 | 8.005
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 387 | 8.027
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 388 | 8.05
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 389 | 8.072
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 390 | 8.094
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 391 | 8.116
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 392 | 8.138
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 393 | 8.161
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 394 | 8.183
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 395 | 8.205
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 396 | 8.227
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 397 | 8.25
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 398 | 8.272
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 399 | 8.294
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 400 | 8.316
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 401 | 8.338
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 402 | 8.361
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 403 | 8.383
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 404 | 8.405
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 405 | 8.427
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 406 | 8.45
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 407 | 8.472
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 408 | 8.494
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 409 | 8.516
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 410 | 8.539
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 411 | 8.561
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 412 | 8.583
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 413 | 8.605
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 414 | 8.628
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 415 | 8.65
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 416 | 8.672
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 417 | 8.694
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 418 | 8.717
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 419 | 8.739
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 420 | 8.761
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 421 | 8.784
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 422 | 8.806
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 423 | 8.828
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 424 | 8.851
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 425 | 8.873
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 426 | 8.895
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 427 | 8.918
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 428 | 8.94
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 429 | 8.962
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 430 | 8.985
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 431 | 9.007
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 432 | 9.029
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 433 | 9.052
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 434 | 9.074
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 435 | 9.096
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 436 | 9.119
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 437 | 9.141
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 438 | 9.163
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 439 | 9.186
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 440 | 9.208
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 441 | 9.231
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 442 | 9.253
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 443 | 9.275
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 444 | 9.298
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 445 | 9.32
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 446 | 9.343
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 447 | 9.365
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 448 | 9.388
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 449 | 9.41
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 450 | 9.432
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 451 | 9.455
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 452 | 9.477
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 453 | 9.5
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 454 | 9.522
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 455 | 9.545
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 456 | 9.567
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 457 | 9.59
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 458 | 9.612
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 459 | 9.635
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 460 | 9.657
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 461 | 9.68
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 462 | 9.702
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 463 | 9.725
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 464 | 9.747
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 465 | 9.77
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 466 | 9.792
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 467 | 9.815
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 468 | 9.837
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 469 | 9.86
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 470 | 9.882
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 471 | 9.905
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 472 | 9.927
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 473 | 9.95
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 474 | 9.973
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 475 | 9.995
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 476 | 10.018
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 477 | 10.04
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 478 | 10.063
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 479 | 10.086
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 480 | 10.108
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 481 | 10.131
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 482 | 10.153
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 483 | 10.176
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 484 | 10.199
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 485 | 10.221
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 486 | 10.244
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 487 | 10.267
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 488 | 10.289
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 489 | 10.312
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 490 | 10.334
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 491 | 10.357
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 492 | 10.38
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 493 | 10.402
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 494 | 10.425
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 495 | 10.448
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 496 | 10.471
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 497 | 10.493
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 498 | 10.516
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 499 | 10.539
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 500 | 10.561
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 501 | 10.584
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 502 | 10.607
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 503 | 10.629
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 504 | 10.652
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 505 | 10.675
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 506 | 10.698
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 507 | 10.72
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 508 | 10.743
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 509 | 10.766
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 510 | 10.789
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 511 | 10.811
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 512 | 10.834
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 513 | 10.857
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 514 | 10.88
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 515 | 10.903
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 516 | 10.925
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 517 | 10.948
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 518 | 10.971
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 519 | 10.994
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 520 | 11.017
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 521 | 11.039
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 522 | 11.062
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 523 | 11.085
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 524 | 11.108
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 525 | 11.131
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 526 | 11.154
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 527 | 11.176
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 528 | 11.199
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 529 | 11.222
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 530 | 11.245
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 531 | 11.268
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 532 | 11.291
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 533 | 11.313
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 534 | 11.336
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 535 | 11.359
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 536 | 11.382
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 537 | 11.405
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 538 | 11.428
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 539 | 11.451
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 540 | 11.474
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 541 | 11.497
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 542 | 11.519
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 543 | 11.542
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 544 | 11.565
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 545 | 11.588
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 546 | 11.611
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 547 | 11.634
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 548 | 11.657
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 549 | 11.68
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 550 | 11.703
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 551 | 11.726
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 552 | 11.749
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 553 | 11.772
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 554 | 11.795
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 555 | 11.818
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 556 | 11.841
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 557 | 11.864
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 558 | 11.887
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 559 | 11.91
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 560 | 11.933
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 561 | 11.956
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 562 | 11.978
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 563 | 12.001
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 564 | 12.024
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 565 | 12.047
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 566 | 12.07
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 567 | 12.093
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 568 | 12.116
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 569 | 12.14
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 570 | 12.163
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 571 | 12.186
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 572 | 12.209
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 573 | 12.232
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 574 | 12.255
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 575 | 12.278
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 576 | 12.301
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 577 | 12.324
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 578 | 12.347
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 579 | 12.37
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 580 | 12.393
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 581 | 12.416
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 582 | 12.439
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 583 | 12.462
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 584 | 12.485
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 585 | 12.508
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 586 | 12.531
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 587 | 12.554
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 588 | 12.577
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 589 | 12.6
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 590 | 12.624
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 591 | 12.647
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 592 | 12.67
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 593 | 12.693
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 594 | 12.716
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 595 | 12.739
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 596 | 12.762
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 597 | 12.785
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 598 | 12.808
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 599 | 12.831
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 600 | 12.855
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 601 | 12.878
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 602 | 12.901
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 603 | 12.924
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 604 | 12.947
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 605 | 12.97
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 606 | 12.993
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 607 | 13.016
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 608 | 13.04
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 609 | 13.063
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 610 | 13.086
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 611 | 13.109
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 612 | 13.132
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 613 | 13.155
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 614 | 13.179
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 615 | 13.202
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 616 | 13.225
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 617 | 13.248
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 618 | 13.271
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 619 | 13.294
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 620 | 13.318
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 621 | 13.341
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 622 | 13.364
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 623 | 13.387
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 624 | 13.41
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 625 | 13.433
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 626 | 13.457
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 627 | 13.48
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 628 | 13.503
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 629 | 13.526
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 630 | 13.549
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 631 | 13.573
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 632 | 13.596
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 633 | 13.619
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 634 | 13.642
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 635 | 13.665
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 636 | 13.689
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 637 | 13.712
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 638 | 13.735
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 639 | 13.758
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 640 | 13.782
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 641 | 13.805
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 642 | 13.828
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 643 | 13.851
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 644 | 13.874
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 645 | 13.898
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 646 | 13.921
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 647 | 13.944
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 648 | 13.967
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 649 | 13.991
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 650 | 14.014
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 651 | 14.037
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 652 | 14.06
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 653 | 14.084
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 654 | 14.107
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 655 | 14.13
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 656 | 14.154
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 657 | 14.177
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 658 | 14.2
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 659 | 14.223
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 660 | 14.247
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 661 | 14.27
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 662 | 14.293
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 663 | 14.316
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 664 | 14.34
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 665 | 14.363
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 666 | 14.386
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 667 | 14.41
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 668 | 14.433
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 669 | 14.456
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 670 | 14.479
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 671 | 14.503
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 672 | 14.526
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 673 | 14.549
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 674 | 14.573
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 675 | 14.596
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 676 | 14.619
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 677 | 14.643
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 678 | 14.666
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 679 | 14.689
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 680 | 14.713
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 681 | 14.736
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 682 | 14.759
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 683 | 14.783
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 684 | 14.806
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 685 | 14.829
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 686 | 14.853
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 687 | 14.876
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 688 | 14.899
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 689 | 14.923
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 690 | 14.946
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 691 | 14.969
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 692 | 14.993
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 693 | 15.016
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 694 | 15.039
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 695 | 15.063
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 696 | 15.086
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 697 | 15.109
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 698 | 15.133
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 699 | 15.156
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 700 | 15.179
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 701 | 15.203
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 702 | 15.226
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 703 | 15.25
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 704 | 15.273
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 705 | 15.296
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 706 | 15.32
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 707 | 15.343
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 708 | 15.366
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 709 | 15.39
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 710 | 15.413
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 711 | 15.437
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 712 | 15.46
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 713 | 15.483
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 714 | 15.507
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 715 | 15.53
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 716 | 15.554
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 717 | 15.577
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 718 | 15.6
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 719 | 15.624
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 720 | 15.647
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 721 | 15.671
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 722 | 15.694
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 723 | 15.717
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 724 | 15.741
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 725 | 15.764
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 726 | 15.788
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 727 | 15.811
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 728 | 15.834
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 729 | 15.858
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 730 | 15.881
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 731 | 15.905
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 732 | 15.928
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 733 | 15.952
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 734 | 15.975
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 735 | 15.998
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 736 | 16.022
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 737 | 16.045
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 738 | 16.069
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 739 | 16.092
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 740 | 16.116
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 741 | 16.139
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 742 | 16.163
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 743 | 16.186
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 744 | 16.209
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 745 | 16.233
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 746 | 16.256
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 747 | 16.28
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 748 | 16.303
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 749 | 16.327
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 750 | 16.35
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 751 | 16.374
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 752 | 16.397
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 753 | 16.421
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 754 | 16.444
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 755 | 16.468
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 756 | 16.491
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 757 | 16.515
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 758 | 16.538
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 759 | 16.561
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 760 | 16.585
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 761 | 16.608
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 762 | 16.632
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 763 | 16.655
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 764 | 16.679
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 765 | 16.702
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 766 | 16.726
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 767 | 16.749
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 768 | 16.773
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 769 | 16.796
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 770 | 16.82
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 771 | 16.843
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 772 | 16.867
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 773 | 16.89
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 774 | 16.914
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 775 | 16.937
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 776 | 16.961
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 777 | 16.984
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 778 | 17.008
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 779 | 17.031
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 780 | 17.055
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 781 | 17.078
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 782 | 17.102
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 783 | 17.125
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 784 | 17.149
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 785 | 17.173
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 786 | 17.196
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 787 | 17.22
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 788 | 17.243
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 789 | 17.267
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 790 | 17.29
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 791 | 17.314
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 792 | 17.337
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 793 | 17.361
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 794 | 17.384
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 795 | 17.408
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 796 | 17.431
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 797 | 17.455
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 798 | 17.478
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 799 | 17.502
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 800 | 17.526
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 801 | 17.549
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 802 | 17.573
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 803 | 17.596
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 804 | 17.62
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 805 | 17.643
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 806 | 17.667
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 807 | 17.69
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 808 | 17.714
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 809 | 17.738
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 810 | 17.761
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 811 | 17.785
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 812 | 17.808
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 813 | 17.832
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 814 | 17.855
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 815 | 17.879
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 816 | 17.902
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 817 | 17.926
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 818 | 17.95
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 819 | 17.973
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 820 | 17.997
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 821 | 18.02
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 822 | 18.044
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 823 | 18.068
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 824 | 18.091
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 825 | 18.115
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 826 | 18.138
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 827 | 18.162
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 828 | 18.185
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 829 | 18.209
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 830 | 18.233
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 831 | 18.256
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 832 | 18.28
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 833 | 18.303
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 834 | 18.327
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 835 | 18.351
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 836 | 18.374
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 837 | 18.398
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 838 | 18.421
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 839 | 18.445
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 840 | 18.469
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 841 | 18.492
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 842 | 18.516
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 843 | 18.539
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 844 | 18.563
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 845 | 18.587
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 846 | 18.61
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 847 | 18.634
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 848 | 18.657
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 849 | 18.681
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 850 | 18.705
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 851 | 18.728
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 852 | 18.752
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 853 | 18.776
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 854 | 18.799
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 855 | 18.823
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 856 | 18.846
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 857 | 18.87
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 858 | 18.894
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 859 | 18.917
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 860 | 18.941
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 861 | 18.965
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 862 | 18.988
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 863 | 19.012
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 864 | 19.035
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 865 | 19.059
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 866 | 19.083
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 867 | 19.106
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 868 | 19.13
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 869 | 19.154
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 870 | 19.177
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 871 | 19.201
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 872 | 19.224
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 873 | 19.248
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 874 | 19.272
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 875 | 19.295
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 876 | 19.319
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 877 | 19.343
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 878 | 19.366
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 879 | 19.39
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 880 | 19.414
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 881 | 19.437
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 882 | 19.461
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 883 | 19.485
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 884 | 19.508
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 885 | 19.532
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 886 | 19.556
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 887 | 19.579
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 888 | 19.603
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 889 | 19.626
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 890 | 19.65
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 891 | 19.674
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 892 | 19.697
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 893 | 19.721
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 894 | 19.745
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 895 | 19.768
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 896 | 19.792
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 897 | 19.816
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 898 | 19.839
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 899 | 19.863
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 900 | 19.887
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 901 | 19.91
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 902 | 19.934
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 903 | 19.958
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 904 | 19.981
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 905 | 20.005
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 906 | 20.029
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 907 | 20.052
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 908 | 20.076
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 909 | 20.1
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 910 | 20.123
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 911 | 20.147
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 912 | 20.171
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 913 | 20.194
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 914 | 20.218
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 915 | 20.242
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 916 | 20.265
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 917 | 20.289
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 918 | 20.313
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 919 | 20.336
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 920 | 20.36
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 921 | 20.384
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 922 | 20.407
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 923 | 20.431
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 924 | 20.455
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 925 | 20.479
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 926 | 20.502
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 927 | 20.526
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 928 | 20.55
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 929 | 20.573
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 930 | 20.597
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 931 | 20.621
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 932 | 20.644
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 933 | 20.668
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 934 | 20.692
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 935 | 20.715
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 936 | 20.739
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 937 | 20.763
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 938 | 20.786
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 939 | 20.81
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 940 | 20.834
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 941 | 20.857
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 942 | 20.881
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 943 | 20.905
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 944 | 20.929
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 945 | 20.952
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 946 | 20.976
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 947 | 21
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 948 | 21.023
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 949 | 21.047
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 950 | 21.071
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 951 | 21.094
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 952 | 21.118
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 953 | 21.142
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 954 | 21.165
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 955 | 21.189
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 956 | 21.213
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 957 | 21.236
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 958 | 21.26
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 959 | 21.284
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 960 | 21.308
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 961 | 21.331
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 962 | 21.355
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 963 | 21.379
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 964 | 21.402
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 965 | 21.426
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 966 | 21.45
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 967 | 21.473
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 968 | 21.497
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 969 | 21.521
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 970 | 21.544
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 971 | 21.568
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 972 | 21.592
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 973 | 21.616
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 974 | 21.639
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 975 | 21.663
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 976 | 21.687
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 977 | 21.71
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 978 | 21.734
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 979 | 21.758
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 980 | 21.781
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 981 | 21.805
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 982 | 21.829
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 983 | 21.852
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 984 | 21.876
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 985 | 21.9
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 986 | 21.924
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 987 | 21.947
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 988 | 21.971
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 989 | 21.995
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 990 | 22.018
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 991 | 22.042
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 992 | 22.066
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 993 | 22.089
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 994 | 22.113
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 995 | 22.137
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 996 | 22.16
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 997 | 22.184
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 998 | 22.208
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 999 | 22.232
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1000 | 22.255
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1001 | 22.279
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1002 | 22.303
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1003 | 22.326
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1004 | 22.35
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1005 | 22.374
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1006 | 22.397
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1007 | 22.421
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1008 | 22.445
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1009 | 22.468
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1010 | 22.492
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1011 | 22.516
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1012 | 22.54
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1013 | 22.563
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1014 | 22.587
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1015 | 22.611
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1016 | 22.634
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1017 | 22.658
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1018 | 22.682
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1019 | 22.705
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1020 | 22.729
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1021 | 22.753
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1022 | 22.776
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1023 | 22.8
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1024 | 22.824
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1025 | 22.847
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1026 | 22.871
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1027 | 22.895
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1028 | 22.919
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1029 | 22.942
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1030 | 22.966
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1031 | 22.99
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1032 | 23.013
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1033 | 23.037
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1034 | 23.061
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1035 | 23.084
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1036 | 23.108
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1037 | 23.132
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1038 | 23.155
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1039 | 23.179
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1040 | 23.203
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1041 | 23.226
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1042 | 23.25
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1043 | 23.274
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1044 | 23.297
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1045 | 23.321
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1046 | 23.345
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1047 | 23.368
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1048 | 23.392
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1049 | 23.416
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1050 | 23.439
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1051 | 23.463
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1052 | 23.487
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1053 | 23.51
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1054 | 23.534
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1055 | 23.558
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1056 | 23.581
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1057 | 23.605
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1058 | 23.629
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1059 | 23.652
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1060 | 23.676
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1061 | 23.7
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1062 | 23.723
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1063 | 23.747
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1064 | 23.771
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1065 | 23.794
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1066 | 23.818
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1067 | 23.842
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1068 | 23.865
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1069 | 23.889
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1070 | 23.913
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1071 | 23.936
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1072 | 23.96
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1073 | 23.984
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1074 | 24.007
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1075 | 24.031
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1076 | 24.055
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1077 | 24.078
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1078 | 24.102
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1079 | 24.126
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1080 | 24.149
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1081 | 24.173
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1082 | 24.197
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1083 | 24.22
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1084 | 24.244
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1085 | 24.267
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1086 | 24.291
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1087 | 24.315
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1088 | 24.338
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1089 | 24.362
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1090 | 24.386
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1091 | 24.409
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1092 | 24.433
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1093 | 24.457
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1094 | 24.48
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1095 | 24.504
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1096 | 24.527
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1097 | 24.551
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1098 | 24.575
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1099 | 24.598
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1100 | 24.622
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1101 | 24.646
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1102 | 24.669
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1103 | 24.693
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1104 | 24.717
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1105 | 24.74
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1106 | 24.764
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1107 | 24.787
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1108 | 24.811
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1109 | 24.835
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1110 | 24.858
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1111 | 24.882
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1112 | 24.905
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1113 | 24.929
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1114 | 24.953
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1115 | 24.976
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1116 | 25
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1117 | 25.024
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1118 | 25.047
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1119 | 25.071
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1120 | 25.094
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1121 | 25.118
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1122 | 25.142
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1123 | 25.165
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1124 | 25.189
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1125 | 25.212
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1126 | 25.236
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1127 | 25.26
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1128 | 25.283
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1129 | 25.307
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1130 | 25.33
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1131 | 25.354
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1132 | 25.377
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1133 | 25.401
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1134 | 25.425
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1135 | 25.448
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1136 | 25.472
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1137 | 25.495
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1138 | 25.519
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1139 | 25.543
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1140 | 25.566
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1141 | 25.59
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1142 | 25.613
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1143 | 25.637
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1144 | 25.66
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1145 | 25.684
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1146 | 25.708
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1147 | 25.731
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1148 | 25.755
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1149 | 25.778
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1150 | 25.802
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1151 | 25.825
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1152 | 25.849
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1153 | 25.873
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1154 | 25.896
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1155 | 25.92
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1156 | 25.943
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1157 | 25.967
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1158 | 25.99
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1159 | 26.014
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1160 | 26.037
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1161 | 26.061
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1162 | 26.084
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1163 | 26.108
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1164 | 26.132
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1165 | 26.155
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1166 | 26.179
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1167 | 26.202
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1168 | 26.226
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1169 | 26.249
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1170 | 26.273
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1171 | 26.296
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1172 | 26.32
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1173 | 26.343
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1174 | 26.367
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1175 | 26.39
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1176 | 26.414
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1177 | 26.437
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1178 | 26.461
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1179 | 26.484
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1180 | 26.508
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1181 | 26.532
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1182 | 26.555
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1183 | 26.579
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1184 | 26.602
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1185 | 26.626
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1186 | 26.649
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1187 | 26.673
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1188 | 26.696
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1189 | 26.72
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1190 | 26.743
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1191 | 26.767
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1192 | 26.79
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1193 | 26.814
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1194 | 26.837
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1195 | 26.861
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1196 | 26.884
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1197 | 26.907
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1198 | 26.931
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1199 | 26.954
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1200 | 26.978
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1201 | 27.001
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1202 | 27.025
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1203 | 27.048
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1204 | 27.072
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1205 | 27.095
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1206 | 27.119
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1207 | 27.142
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1208 | 27.166
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1209 | 27.189
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1210 | 27.213
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1211 | 27.236
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1212 | 27.26
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1213 | 27.283
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1214 | 27.306
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1215 | 27.33
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1216 | 27.353
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1217 | 27.377
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1218 | 27.4
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1219 | 27.424
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1220 | 27.447
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1221 | 27.471
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1222 | 27.494
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1223 | 27.517
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1224 | 27.541
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1225 | 27.564
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1226 | 27.588
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1227 | 27.611
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1228 | 27.635
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1229 | 27.658
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1230 | 27.681
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1231 | 27.705
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1232 | 27.728
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1233 | 27.752
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1234 | 27.775
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1235 | 27.798
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1236 | 27.822
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1237 | 27.845
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1238 | 27.869
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1239 | 27.892
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1240 | 27.915
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1241 | 27.939
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1242 | 27.962
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1243 | 27.986
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1244 | 28.009
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1245 | 28.032
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1246 | 28.056
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1247 | 28.079
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1248 | 28.103
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1249 | 28.126
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1250 | 28.149
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1251 | 28.173
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1252 | 28.196
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1253 | 28.219
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1254 | 28.243
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1255 | 28.266
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1256 | 28.289
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1257 | 28.313
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1258 | 28.336
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1259 | 28.36
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1260 | 28.383
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1261 | 28.406
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1262 | 28.43
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1263 | 28.453
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1264 | 28.476
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1265 | 28.5
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1266 | 28.523
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1267 | 28.546
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1268 | 28.57
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1269 | 28.593
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1270 | 28.616
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1271 | 28.64
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1272 | 28.663
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1273 | 28.686
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1274 | 28.71
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1275 | 28.733
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1276 | 28.756
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1277 | 28.78
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1278 | 28.803
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1279 | 28.826
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1280 | 28.849
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1281 | 28.873
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1282 | 28.896
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1283 | 28.919
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1284 | 28.943
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1285 | 28.966
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1286 | 28.989
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1287 | 29.013
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1288 | 29.036
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1289 | 29.059
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1290 | 29.082
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1291 | 29.106
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1292 | 29.129
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1293 | 29.152
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1294 | 29.176
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1295 | 29.199
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1296 | 29.222
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1297 | 29.245
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1298 | 29.269
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1299 | 29.292
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1300 | 29.315
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1301 | 29.338
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1302 | 29.362
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1303 | 29.385
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1304 | 29.408
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1305 | 29.431
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1306 | 29.455
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1307 | 29.478
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1308 | 29.501
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1309 | 29.524
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1310 | 29.548
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1311 | 29.571
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1312 | 29.594
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1313 | 29.617
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1314 | 29.64
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1315 | 29.664
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1316 | 29.687
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1317 | 29.71
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1318 | 29.733
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1319 | 29.757
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1320 | 29.78
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1321 | 29.803
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1322 | 29.826
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1323 | 29.849
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1324 | 29.873
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1325 | 29.896
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1326 | 29.919
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1327 | 29.942
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1328 | 29.965
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1329 | 29.989
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1330 | 30.012
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1331 | 30.035
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1332 | 30.058
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1333 | 30.081
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1334 | 30.104
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1335 | 30.128
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1336 | 30.151
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1337 | 30.174
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1338 | 30.197
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1339 | 30.22
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1340 | 30.243
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1341 | 30.267
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1342 | 30.29
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1343 | 30.313
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1344 | 30.336
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1345 | 30.359
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1346 | 30.382
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1347 | 30.405
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1348 | 30.429
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1349 | 30.452
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1350 | 30.475
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1351 | 30.498
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1352 | 30.521
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1353 | 30.544
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1354 | 30.567
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1355 | 30.59
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1356 | 30.613
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1357 | 30.637
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1358 | 30.66
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1359 | 30.683
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1360 | 30.706
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1361 | 30.729
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1362 | 30.752
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1363 | 30.775
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1364 | 30.798
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1365 | 30.821
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1366 | 30.844
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1367 | 30.868
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1368 | 30.891
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1369 | 30.914
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1370 | 30.937
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1371 | 30.96
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1372 | 30.983
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1373 | 31.006
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1374 | 31.029
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1375 | 31.052
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1376 | 31.075
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1377 | 31.098
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1378 | 31.121
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1379 | 31.144
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1380 | 31.167
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1381 | 31.19
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1382 | 31.213
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1383 | 31.236
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1384 | 31.26
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1385 | 31.283
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1386 | 31.306
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1387 | 31.329
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1388 | 31.352
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1389 | 31.375
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1390 | 31.398
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1391 | 31.421
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1392 | 31.444
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1393 | 31.467
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1394 | 31.49
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1395 | 31.513
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1396 | 31.536
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1397 | 31.559
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1398 | 31.582
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1399 | 31.605
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1400 | 31.628
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1401 | 31.651
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1402 | 31.674
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1403 | 31.697
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1404 | 31.72
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1405 | 31.743
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1406 | 31.766
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1407 | 31.789
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1408 | 31.812
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1409 | 31.834
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1410 | 31.857
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1411 | 31.88
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1412 | 31.903
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1413 | 31.926
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1414 | 31.949
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1415 | 31.972
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1416 | 31.995
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1417 | 32.018
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1418 | 32.041
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1419 | 32.064
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1420 | 32.087
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1421 | 32.11
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1422 | 32.133
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1423 | 32.156
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1424 | 32.179
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1425 | 32.202
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1426 | 32.224
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1427 | 32.247
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1428 | 32.27
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1429 | 32.293
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1430 | 32.316
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1431 | 32.339
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1432 | 32.362
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1433 | 32.385
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1434 | 32.408
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1435 | 32.431
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1436 | 32.453
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1437 | 32.476
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1438 | 32.499
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1439 | 32.522
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1440 | 32.545
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1441 | 32.568
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1442 | 32.591
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1443 | 32.614
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1444 | 32.636
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1445 | 32.659
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1446 | 32.682
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1447 | 32.705
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1448 | 32.728
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1449 | 32.751
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1450 | 32.774
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1451 | 32.796
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1452 | 32.819
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1453 | 32.842
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1454 | 32.865
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1455 | 32.888
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1456 | 32.911
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1457 | 32.933
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1458 | 32.956
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1459 | 32.979
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1460 | 33.002
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1461 | 33.025
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1462 | 33.047
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1463 | 33.07
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1464 | 33.093
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1465 | 33.116
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1466 | 33.139
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1467 | 33.161
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1468 | 33.184
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1469 | 33.207
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1470 | 33.23
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1471 | 33.253
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1472 | 33.275
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1473 | 33.298
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1474 | 33.321
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1475 | 33.344
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1476 | 33.366
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1477 | 33.389
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1478 | 33.412
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1479 | 33.435
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1480 | 33.458
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1481 | 33.48
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1482 | 33.503
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1483 | 33.526
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1484 | 33.548
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1485 | 33.571
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1486 | 33.594
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1487 | 33.617
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1488 | 33.639
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1489 | 33.662
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1490 | 33.685
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1491 | 33.708
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1492 | 33.73
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1493 | 33.753
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1494 | 33.776
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1495 | 33.798
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1496 | 33.821
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1497 | 33.844
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1498 | 33.867
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1499 | 33.889
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1500 | 33.912
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1501 | 33.935
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1502 | 33.957
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1503 | 33.98
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1504 | 34.003
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1505 | 34.025
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1506 | 34.048
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1507 | 34.071
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1508 | 34.093
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1509 | 34.116
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1510 | 34.139
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1511 | 34.161
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1512 | 34.184
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1513 | 34.207
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1514 | 34.229
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1515 | 34.252
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1516 | 34.275
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1517 | 34.297
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1518 | 34.32
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1519 | 34.343
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1520 | 34.365
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1521 | 34.388
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1522 | 34.41
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1523 | 34.433
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1524 | 34.456
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1525 | 34.478
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1526 | 34.501
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1527 | 34.524
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1528 | 34.546
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1529 | 34.569
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1530 | 34.591
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1531 | 34.614
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1532 | 34.637
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1533 | 34.659
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1534 | 34.682
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1535 | 34.704
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1536 | 34.727
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1537 | 34.75
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1538 | 34.772
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1539 | 34.795
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1540 | 34.817
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1541 | 34.84
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1542 | 34.862
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1543 | 34.885
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1544 | 34.908
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1545 | 34.93
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1546 | 34.953
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1547 | 34.975
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1548 | 34.998
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1549 | 35.02
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1550 | 35.043
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1551 | 35.065
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1552 | 35.088
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1553 | 35.11
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1554 | 35.133
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1555 | 35.156
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1556 | 35.178
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1557 | 35.201
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1558 | 35.223
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1559 | 35.246
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1560 | 35.268
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1561 | 35.291
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1562 | 35.313
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1563 | 35.336
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1564 | 35.358
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1565 | 35.381
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1566 | 35.403
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1567 | 35.426
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1568 | 35.448
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1569 | 35.471
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1570 | 35.493
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1571 | 35.516
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1572 | 35.538
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1573 | 35.56
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1574 | 35.583
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1575 | 35.605
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1576 | 35.628
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1577 | 35.65
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1578 | 35.673
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1579 | 35.695
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1580 | 35.718
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1581 | 35.74
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1582 | 35.763
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1583 | 35.785
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1584 | 35.807
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1585 | 35.83
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1586 | 35.852
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1587 | 35.875
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1588 | 35.897
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1589 | 35.92
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1590 | 35.942
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1591 | 35.964
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1592 | 35.987
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1593 | 36.009
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1594 | 36.032
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1595 | 36.054
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1596 | 36.076
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1597 | 36.099
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1598 | 36.121
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1599 | 36.144
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1600 | 36.166
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1601 | 36.188
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1602 | 36.211
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1603 | 36.233
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1604 | 36.256
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1605 | 36.278
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1606 | 36.3
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1607 | 36.323
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1608 | 36.345
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1609 | 36.367
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1610 | 36.39
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1611 | 36.412
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1612 | 36.434
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1613 | 36.457
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1614 | 36.479
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1615 | 36.501
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1616 | 36.524
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1617 | 36.546
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1618 | 36.568
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1619 | 36.591
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1620 | 36.613
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1621 | 36.635
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1622 | 36.658
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1623 | 36.68
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1624 | 36.702
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1625 | 36.725
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1626 | 36.747
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1627 | 36.769
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1628 | 36.792
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1629 | 36.814
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1630 | 36.836
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1631 | 36.859
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1632 | 36.881
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1633 | 36.903
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1634 | 36.925
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1635 | 36.948
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1636 | 36.97
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1637 | 36.992
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1638 | 37.014
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1639 | 37.037
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1640 | 37.059
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1641 | 37.081
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1642 | 37.104
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1643 | 37.126
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1644 | 37.148
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1645 | 37.17
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1646 | 37.193
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1647 | 37.215
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1648 | 37.237
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1649 | 37.259
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1650 | 37.281
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1651 | 37.304
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1652 | 37.326
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1653 | 37.348
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1654 | 37.37
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1655 | 37.393
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1656 | 37.415
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1657 | 37.437
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1658 | 37.459
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1659 | 37.481
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1660 | 37.504
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1661 | 37.526
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1662 | 37.548
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1663 | 37.57
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1664 | 37.592
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1665 | 37.615
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1666 | 37.637
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1667 | 37.659
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1668 | 37.681
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1669 | 37.703
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1670 | 37.725
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1671 | 37.748
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1672 | 37.77
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1673 | 37.792
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1674 | 37.814
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1675 | 37.836
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1676 | 37.858
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1677 | 37.881
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1678 | 37.903
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1679 | 37.925
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1680 | 37.947
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1681 | 37.969
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1682 | 37.991
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1683 | 38.013
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1684 | 38.036
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1685 | 38.058
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1686 | 38.08
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1687 | 38.102
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1688 | 38.124
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1689 | 38.146
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1690 | 38.168
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1691 | 38.19
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1692 | 38.212
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1693 | 38.235
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1694 | 38.257
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1695 | 38.279
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1696 | 38.301
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1697 | 38.323
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1698 | 38.345
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1699 | 38.367
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1700 | 38.389
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1701 | 38.411
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1702 | 38.433
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1703 | 38.455
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1704 | 38.477
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1705 | 38.499
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1706 | 38.522
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1707 | 38.544
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1708 | 38.566
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1709 | 38.588
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1710 | 38.61
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1711 | 38.632
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1712 | 38.654
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1713 | 38.676
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1714 | 38.698
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1715 | 38.72
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1716 | 38.742
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1717 | 38.764
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1718 | 38.786
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1719 | 38.808
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1720 | 38.83
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1721 | 38.852
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1722 | 38.874
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1723 | 38.896
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1724 | 38.918
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1725 | 38.94
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1726 | 38.962
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1727 | 38.984
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1728 | 39.006
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1729 | 39.028
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1730 | 39.05
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1731 | 39.072
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1732 | 39.094
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1733 | 39.116
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1734 | 39.138
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1735 | 39.16
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1736 | 39.182
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1737 | 39.204
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1738 | 39.226
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1739 | 39.248
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1740 | 39.27
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1741 | 39.292
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1742 | 39.314
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1743 | 39.335
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1744 | 39.357
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1745 | 39.379
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1746 | 39.401
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1747 | 39.423
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1748 | 39.445
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1749 | 39.467
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1750 | 39.489
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1751 | 39.511
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1752 | 39.533
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1753 | 39.555
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1754 | 39.577
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1755 | 39.599
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1756 | 39.62
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1757 | 39.642
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1758 | 39.664
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1759 | 39.686
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1760 | 39.708
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1761 | 39.73
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1762 | 39.752
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1763 | 39.774
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1764 | 39.796
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1765 | 39.817
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1766 | 39.839
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1767 | 39.861
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1768 | 39.883
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1769 | 39.905
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1770 | 39.927
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1771 | 39.949
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1772 | 39.97
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1773 | 39.992
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1774 | 40.014
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1775 | 40.036
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1776 | 40.058
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1777 | 40.08
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1778 | 40.101
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1779 | 40.123
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1780 | 40.145
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1781 | 40.167
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1782 | 40.189
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1783 | 40.211
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1784 | 40.232
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1785 | 40.254
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1786 | 40.276
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1787 | 40.298
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1788 | 40.32
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1789 | 40.341
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1790 | 40.363
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1791 | 40.385
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1792 | 40.407
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1793 | 40.429
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1794 | 40.45
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1795 | 40.472
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1796 | 40.494
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1797 | 40.516
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1798 | 40.537
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1799 | 40.559
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1800 | 40.581
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1801 | 40.603
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1802 | 40.624
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1803 | 40.646
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1804 | 40.668
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1805 | 40.69
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1806 | 40.711
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1807 | 40.733
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1808 | 40.755
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1809 | 40.777
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1810 | 40.798
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1811 | 40.82
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1812 | 40.842
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1813 | 40.864
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1814 | 40.885
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1815 | 40.907
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1816 | 40.929
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1817 | 40.95
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1818 | 40.972
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1819 | 40.994
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1820 | 41.015
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1821 | 41.037
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1822 | 41.059
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1823 | 41.081
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1824 | 41.102
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1825 | 41.124
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1826 | 41.146
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1827 | 41.167
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1828 | 41.189
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1829 | 41.211
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1830 | 41.232
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1831 | 41.254
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1832 | 41.276
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1833 | 41.297
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1834 | 41.319
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1835 | 41.341
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1836 | 41.362
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1837 | 41.384
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1838 | 41.405
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1839 | 41.427
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1840 | 41.449
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1841 | 41.47
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1842 | 41.492
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1843 | 41.514
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1844 | 41.535
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1845 | 41.557
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1846 | 41.578
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1847 | 41.6
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1848 | 41.622
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1849 | 41.643
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1850 | 41.665
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1851 | 41.686
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1852 | 41.708
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1853 | 41.73
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1854 | 41.751
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1855 | 41.773
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1856 | 41.794
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1857 | 41.816
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1858 | 41.838
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1859 | 41.859
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1860 | 41.881
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1861 | 41.902
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1862 | 41.924
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1863 | 41.945
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1864 | 41.967
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1865 | 41.988
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1866 | 42.01
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1867 | 42.032
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1868 | 42.053
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1869 | 42.075
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1870 | 42.096
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1871 | 42.118
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1872 | 42.139
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1873 | 42.161
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1874 | 42.182
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1875 | 42.204
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1876 | 42.225
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1877 | 42.247
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1878 | 42.268
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1879 | 42.29
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1880 | 42.311
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1881 | 42.333
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1882 | 42.354
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1883 | 42.376
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1884 | 42.397
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1885 | 42.419
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1886 | 42.44
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1887 | 42.462
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1888 | 42.483
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1889 | 42.505
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1890 | 42.526
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1891 | 42.548
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1892 | 42.569
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1893 | 42.591
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1894 | 42.612
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1895 | 42.633
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1896 | 42.655
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1897 | 42.676
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1898 | 42.698
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1899 | 42.719
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1900 | 42.741
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1901 | 42.762
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1902 | 42.783
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1903 | 42.805
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1904 | 42.826
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1905 | 42.848
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1906 | 42.869
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1907 | 42.891
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1908 | 42.912
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1909 | 42.933
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1910 | 42.955
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1911 | 42.976
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1912 | 42.998
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1913 | 43.019
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1914 | 43.04
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1915 | 43.062
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1916 | 43.083
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1917 | 43.104
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1918 | 43.126
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1919 | 43.147
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1920 | 43.169
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1921 | 43.19
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1922 | 43.211
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1923 | 43.233
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1924 | 43.254
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1925 | 43.275
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1926 | 43.297
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1927 | 43.318
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1928 | 43.339
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1929 | 43.361
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1930 | 43.382
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1931 | 43.403
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1932 | 43.425
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1933 | 43.446
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1934 | 43.467
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1935 | 43.489
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1936 | 43.51
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1937 | 43.531
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1938 | 43.552
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1939 | 43.574
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1940 | 43.595
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1941 | 43.616
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1942 | 43.638
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1943 | 43.659
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1944 | 43.68
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1945 | 43.701
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1946 | 43.723
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1947 | 43.744
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1948 | 43.765
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1949 | 43.787
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1950 | 43.808
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1951 | 43.829
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1952 | 43.85
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1953 | 43.872
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1954 | 43.893
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1955 | 43.914
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1956 | 43.935
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1957 | 43.957
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1958 | 43.978
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1959 | 43.999
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1960 | 44.02
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1961 | 44.041
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1962 | 44.063
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1963 | 44.084
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1964 | 44.105
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1965 | 44.126
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1966 | 44.147
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1967 | 44.169
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1968 | 44.19
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1969 | 44.211
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1970 | 44.232
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1971 | 44.253
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1972 | 44.275
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1973 | 44.296
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1974 | 44.317
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1975 | 44.338
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1976 | 44.359
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1977 | 44.38
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1978 | 44.402
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1979 | 44.423
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1980 | 44.444
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1981 | 44.465
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1982 | 44.486
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1983 | 44.507
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1984 | 44.528
[7 empty] | 1985 | 44.55
[7 empty] | 1986 | 44.571
[7 empty] | 1987 | 44.592
[7 empty] | 1988 | 44.613
[7 empty] | 1989 | 44.634
[7 empty] | 1990 | 44.655
[7 empty] | 1991 | 44.676
[7 empty] | 1992 | 44.697
[7 empty] | 1993 | 44.719
[7 empty] | 1994 | 44.74
[7 empty] | 1995 | 44.761
[7 empty] | 1996 | 44.782
[7 empty] | 1997 | 44.803
[7 empty] | 1998 | 44.824
[7 empty] | 1999 | 44.845
[7 empty] | 2000 | 44.866
[7 empty] | 2001 | 44.887
[7 empty] | 2002 | 44.908
[7 empty] | 2003 | 44.929
[7 empty] | 2004 | 44.95
[7 empty] | 2005 | 44.971
[7 empty] | 2006 | 44.992
[7 empty] | 2007 | 45.014
[7 empty] | 2008 | 45.035
[7 empty] | 2009 | 45.056
[7 empty] | 2010 | 45.077
[7 empty] | 2011 | 45.098
[7 empty] | 2012 | 45.119
[7 empty] | 2013 | 45.14
[7 empty] | 2014 | 45.161
[7 empty] | 2015 | 45.182
[7 empty] | 2016 | 45.203
[7 empty] | 2017 | 45.224
[7 empty] | 2018 | 45.245
[7 empty] | 2019 | 45.266
[7 empty] | 2020 | 45.287
[7 empty] | 2021 | 45.308
[7 empty] | 2022 | 45.329
[7 empty] | 2023 | 45.35
[7 empty] | 2024 | 45.371
[7 empty] | 2025 | 45.392
[7 empty] | 2026 | 45.413
[7 empty] | 2027 | 45.434
[7 empty] | 2028 | 45.455
[7 empty] | 2029 | 45.476
[7 empty] | 2030 | 45.497
[7 empty] | 2031 | 45.518
[7 empty] | 2032 | 45.539
[7 empty] | 2033 | 45.56
[7 empty] | 2034 | 45.58
[7 empty] | 2035 | 45.601
[7 empty] | 2036 | 45.622
[7 empty] | 2037 | 45.643
[7 empty] | 2038 | 45.664
[7 empty] | 2039 | 45.685
[7 empty] | 2040 | 45.706
[7 empty] | 2041 | 45.727
[7 empty] | 2042 | 45.748
[7 empty] | 2043 | 45.769
[7 empty] | 2044 | 45.79
[7 empty] | 2045 | 45.811
[7 empty] | 2046 | 45.832
[7 empty] | 2047 | 45.852
[7 empty] | 2048 | 45.873
[7 empty] | 2049 | 45.894
[7 empty] | 2050 | 45.915
[7 empty] | 2051 | 45.936
[7 empty] | 2052 | 45.957
[7 empty] | 2053 | 45.978
[7 empty] | 2054 | 45.999
[7 empty] | 2055 | 46.019
[7 empty] | 2056 | 46.04
[7 empty] | 2057 | 46.061
[7 empty] | 2058 | 46.082
[7 empty] | 2059 | 46.103
[7 empty] | 2060 | 46.124
[7 empty] | 2061 | 46.145
[7 empty] | 2062 | 46.165
[7 empty] | 2063 | 46.186
[7 empty] | 2064 | 46.207
[7 empty] | 2065 | 46.228
[7 empty] | 2066 | 46.249
[7 empty] | 2067 | 46.269
[7 empty] | 2068 | 46.29
[7 empty] | 2069 | 46.311
[7 empty] | 2070 | 46.332
[7 empty] | 2071 | 46.353
[7 empty] | 2072 | 46.374
[7 empty] | 2073 | 46.394
[7 empty] | 2074 | 46.415
[7 empty] | 2075 | 46.436
[7 empty] | 2076 | 46.457
[7 empty] | 2077 | 46.477
[7 empty] | 2078 | 46.498
[7 empty] | 2079 | 46.519
[7 empty] | 2080 | 46.54
[7 empty] | 2081 | 46.56
[7 empty] | 2082 | 46.581
[7 empty] | 2083 | 46.602
[7 empty] | 2084 | 46.623
[7 empty] | 2085 | 46.643
[7 empty] | 2086 | 46.664
[7 empty] | 2087 | 46.685
[7 empty] | 2088 | 46.706
[7 empty] | 2089 | 46.726
[7 empty] | 2090 | 46.747
[7 empty] | 2091 | 46.768
[7 empty] | 2092 | 46.789
[7 empty] | 2093 | 46.809
[7 empty] | 2094 | 46.83
[7 empty] | 2095 | 46.851
[7 empty] | 2096 | 46.871
[7 empty] | 2097 | 46.892
[7 empty] | 2098 | 46.913
[7 empty] | 2099 | 46.933
[7 empty] | 2100 | 46.954
[7 empty] | 2101 | 46.975
[7 empty] | 2102 | 46.995
[7 empty] | 2103 | 47.016
[7 empty] | 2104 | 47.037
[7 empty] | 2105 | 47.057
[7 empty] | 2106 | 47.078
[7 empty] | 2107 | 47.099
[7 empty] | 2108 | 47.119
[7 empty] | 2109 | 47.14
[7 empty] | 2110 | 47.161
[7 empty] | 2111 | 47.181
[7 empty] | 2112 | 47.202
[7 empty] | 2113 | 47.223
[7 empty] | 2114 | 47.243
[7 empty] | 2115 | 47.264
[7 empty] | 2116 | 47.284
[7 empty] | 2117 | 47.305
[7 empty] | 2118 | 47.326
[7 empty] | 2119 | 47.346
[7 empty] | 2120 | 47.367
[7 empty] | 2121 | 47.387
[7 empty] | 2122 | 47.408
[7 empty] | 2123 | 47.429
[7 empty] | 2124 | 47.449
[7 empty] | 2125 | 47.47
[7 empty] | 2126 | 47.49
[7 empty] | 2127 | 47.511
[7 empty] | 2128 | 47.531
[7 empty] | 2129 | 47.552
[7 empty] | 2130 | 47.573
[7 empty] | 2131 | 47.593
[7 empty] | 2132 | 47.614
[7 empty] | 2133 | 47.634
[7 empty] | 2134 | 47.655
[7 empty] | 2135 | 47.675
[7 empty] | 2136 | 47.696
[7 empty] | 2137 | 47.716
[7 empty] | 2138 | 47.737
[7 empty] | 2139 | 47.757
[7 empty] | 2140 | 47.778
[7 empty] | 2141 | 47.798
[7 empty] | 2142 | 47.819
[7 empty] | 2143 | 47.839
[7 empty] | 2144 | 47.86
[7 empty] | 2145 | 47.88
[7 empty] | 2146 | 47.901
[7 empty] | 2147 | 47.921
[7 empty] | 2148 | 47.942
[7 empty] | 2149 | 47.962
[7 empty] | 2150 | 47.983
[7 empty] | 2151 | 48.003
[7 empty] | 2152 | 48.024
[7 empty] | 2153 | 48.044
[7 empty] | 2154 | 48.065
[7 empty] | 2155 | 48.085
[7 empty] | 2156 | 48.105
[7 empty] | 2157 | 48.126
[7 empty] | 2158 | 48.146
[7 empty] | 2159 | 48.167
[7 empty] | 2160 | 48.187
[7 empty] | 2161 | 48.208
[7 empty] | 2162 | 48.228
[7 empty] | 2163 | 48.248
[7 empty] | 2164 | 48.269
[7 empty] | 2165 | 48.289
[7 empty] | 2166 | 48.31
[7 empty] | 2167 | 48.33
[7 empty] | 2168 | 48.35
[7 empty] | 2169 | 48.371
[7 empty] | 2170 | 48.391
[7 empty] | 2171 | 48.411
[7 empty] | 2172 | 48.432
[7 empty] | 2173 | 48.452
[7 empty] | 2174 | 48.473
[7 empty] | 2175 | 48.493
[7 empty] | 2176 | 48.513
[7 empty] | 2177 | 48.534
[7 empty] | 2178 | 48.554
[7 empty] | 2179 | 48.574
[7 empty] | 2180 | 48.595
[7 empty] | 2181 | 48.615
[7 empty] | 2182 | 48.635
[7 empty] | 2183 | 48.656
[7 empty] | 2184 | 48.676
[7 empty] | 2185 | 48.696
[7 empty] | 2186 | 48.717
[7 empty] | 2187 | 48.737
[7 empty] | 2188 | 48.757
[7 empty] | 2189 | 48.777
[7 empty] | 2190 | 48.798
[7 empty] | 2191 | 48.818
[7 empty] | 2192 | 48.838
[7 empty] | 2193 | 48.859
[7 empty] | 2194 | 48.879
[7 empty] | 2195 | 48.899
[7 empty] | 2196 | 48.919
[7 empty] | 2197 | 48.94
[7 empty] | 2198 | 48.96
[7 empty] | 2199 | 48.98
[7 empty] | 2200 | 49
[7 empty] | 2201 | 49.021
[7 empty] | 2202 | 49.041
[7 empty] | 2203 | 49.061
[7 empty] | 2204 | 49.081
[7 empty] | 2205 | 49.101
[7 empty] | 2206 | 49.122
[7 empty] | 2207 | 49.142
[7 empty] | 2208 | 49.162
[7 empty] | 2209 | 49.182
[7 empty] | 2210 | 49.202
[7 empty] | 2211 | 49.223
[7 empty] | 2212 | 49.243
[7 empty] | 2213 | 49.263
[7 empty] | 2214 | 49.283
[7 empty] | 2215 | 49.303
[7 empty] | 2216 | 49.323
[7 empty] | 2217 | 49.344
[7 empty] | 2218 | 49.364
[7 empty] | 2219 | 49.384
[7 empty] | 2220 | 49.404
[7 empty] | 2221 | 49.424
[7 empty] | 2222 | 49.444
[7 empty] | 2223 | 49.465
[7 empty] | 2224 | 49.485
[7 empty] | 2225 | 49.505
[7 empty] | 2226 | 49.525
[7 empty] | 2227 | 49.545
[7 empty] | 2228 | 49.565
[7 empty] | 2229 | 49.585
[7 empty] | 2230 | 49.605
[7 empty] | 2231 | 49.625
[7 empty] | 2232 | 49.645
[7 empty] | 2233 | 49.666
[7 empty] | 2234 | 49.686
[7 empty] | 2235 | 49.706
[7 empty] | 2236 | 49.726
[7 empty] | 2237 | 49.746
[7 empty] | 2238 | 49.766
[7 empty] | 2239 | 49.786
[7 empty] | 2240 | 49.806
[7 empty] | 2241 | 49.826
[7 empty] | 2242 | 49.846
[7 empty] | 2243 | 49.866
[7 empty] | 2244 | 49.886
[7 empty] | 2245 | 49.906
[7 empty] | 2246 | 49.926
[7 empty] | 2247 | 49.946
[7 empty] | 2248 | 49.966
[7 empty] | 2249 | 49.986
[7 empty] | 2250 | 50.006
[7 empty] | 2251 | 50.026
[7 empty] | 2252 | 50.046
[7 empty] | 2253 | 50.066
[7 empty] | 2254 | 50.086
[7 empty] | 2255 | 50.106
[7 empty] | 2256 | 50.126
[7 empty] | 2257 | 50.146
[7 empty] | 2258 | 50.166
[7 empty] | 2259 | 50.186
[7 empty] | 2260 | 50.206
[7 empty] | 2261 | 50.226
[7 empty] | 2262 | 50.246
[7 empty] | 2263 | 50.266
[7 empty] | 2264 | 50.286
[7 empty] | 2265 | 50.306
[7 empty] | 2266 | 50.326
[7 empty] | 2267 | 50.346
[7 empty] | 2268 | 50.366
[7 empty] | 2269 | 50.385
[7 empty] | 2270 | 50.405
[7 empty] | 2271 | 50.425
[7 empty] | 2272 | 50.445
[7 empty] | 2273 | 50.465
[7 empty] | 2274 | 50.485
[7 empty] | 2275 | 50.505
[7 empty] | 2276 | 50.525
[7 empty] | 2277 | 50.545
[7 empty] | 2278 | 50.564
[7 empty] | 2279 | 50.584
[7 empty] | 2280 | 50.604
[7 empty] | 2281 | 50.624
[7 empty] | 2282 | 50.644
[7 empty] | 2283 | 50.664
[7 empty] | 2284 | 50.684
[7 empty] | 2285 | 50.703
[7 empty] | 2286 | 50.723
[7 empty] | 2287 | 50.743
[7 empty] | 2288 | 50.763
[7 empty] | 2289 | 50.783
[7 empty] | 2290 | 50.802
[7 empty] | 2291 | 50.822
[7 empty] | 2292 | 50.842
[7 empty] | 2293 | 50.862
[7 empty] | 2294 | 50.882
[7 empty] | 2295 | 50.901
[7 empty] | 2296 | 50.921
[7 empty] | 2297 | 50.941
[7 empty] | 2298 | 50.961
[7 empty] | 2299 | 50.981
[7 empty] | 2300 | 51
[7 empty] | 2301 | 51.02
[7 empty] | 2302 | 51.04
[7 empty] | 2303 | 51.06
[7 empty] | 2304 | 51.079
[7 empty] | 2305 | 51.099
[7 empty] | 2306 | 51.119
[7 empty] | 2307 | 51.139
[7 empty] | 2308 | 51.158
[7 empty] | 2309 | 51.178
[7 empty] | 2310 | 51.198
[7 empty] | 2311 | 51.217
[7 empty] | 2312 | 51.237
[7 empty] | 2313 | 51.257
[7 empty] | 2314 | 51.276
[7 empty] | 2315 | 51.296
[7 empty] | 2316 | 51.316
[7 empty] | 2317 | 51.336
[7 empty] | 2318 | 51.355
[7 empty] | 2319 | 51.375
[7 empty] | 2320 | 51.395
[7 empty] | 2321 | 51.414
[7 empty] | 2322 | 51.434
[7 empty] | 2323 | 51.453
[7 empty] | 2324 | 51.473
[7 empty] | 2325 | 51.493
[7 empty] | 2326 | 51.512
[7 empty] | 2327 | 51.532
[7 empty] | 2328 | 51.552
[7 empty] | 2329 | 51.571
[7 empty] | 2330 | 51.591
[7 empty] | 2331 | 51.611
[7 empty] | 2332 | 51.63
[7 empty] | 2333 | 51.65
[7 empty] | 2334 | 51.669
[7 empty] | 2335 | 51.689
[7 empty] | 2336 | 51.708
[7 empty] | 2337 | 51.728
[7 empty] | 2338 | 51.748
[7 empty] | 2339 | 51.767
[7 empty] | 2340 | 51.787
[7 empty] | 2341 | 51.806
[7 empty] | 2342 | 51.826
[7 empty] | 2343 | 51.845
[7 empty] | 2344 | 51.865
[7 empty] | 2345 | 51.885
[7 empty] | 2346 | 51.904
[7 empty] | 2347 | 51.924
[7 empty] | 2348 | 51.943
[7 empty] | 2349 | 51.963
[7 empty] | 2350 | 51.982
[7 empty] | 2351 | 52.002
[7 empty] | 2352 | 52.021
[7 empty] | 2353 | 52.041
[7 empty] | 2354 | 52.06
[7 empty] | 2355 | 52.08
[7 empty] | 2356 | 52.099
[7 empty] | 2357 | 52.119
[7 empty] | 2358 | 52.138
[7 empty] | 2359 | 52.158
[7 empty] | 2360 | 52.177
[7 empty] | 2361 | 52.197
[7 empty] | 2362 | 52.216
[7 empty] | 2363 | 52.235
[7 empty] | 2364 | 52.255
[7 empty] | 2365 | 52.274
[7 empty] | 2366 | 52.294
[7 empty] | 2367 | 52.313
[7 empty] | 2368 | 52.333
[7 empty] | 2369 | 52.352
[7 empty] | 2370 | 52.371
[7 empty] | 2371 | 52.391
[7 empty] | 2372 | 52.41
[7 empty] | 2373 | 52.43
[7 empty] | 2374 | 52.449
[7 empty] | 2375 | 52.468
[7 empty] | 2376 | 52.488
[7 empty] | 2377 | 52.507
[7 empty] | 2378 | 52.527
[7 empty] | 2379 | 52.546
[7 empty] | 2380 | 52.565
[7 empty] | 2381 | 52.585
[7 empty] | 2382 | 52.604
[7 empty] | 2383 | 52.623
[7 empty] | 2384 | 52.643
[7 empty] | 2385 | 52.662
[7 empty] | 2386 | 52.681
[7 empty] | 2387 | 52.701
[7 empty] | 2388 | 52.72
[7 empty] | 2389 | 52.739
[7 empty] | 2390 | 52.759
[7 empty] | 2391 | 52.778
[7 empty] | 2392 | 52.797
[7 empty] | 2393 | 52.817
[7 empty] | 2394 | 52.836
[7 empty] | 2395 | 52.855
[7 empty] | 2396 | 52.875
[7 empty] | 2397 | 52.894
[7 empty] | 2398 | 52.913
[7 empty] | 2399 | 52.932
[7 empty] | 2400 | 52.952
[7 empty] | 2401 | 52.971
[7 empty] | 2402 | 52.99
[7 empty] | 2403 | 53.01
[7 empty] | 2404 | 53.029
[7 empty] | 2405 | 53.048
[7 empty] | 2406 | 53.067
[7 empty] | 2407 | 53.087
[7 empty] | 2408 | 53.106
[7 empty] | 2409 | 53.125
[7 empty] | 2410 | 53.144
[7 empty] | 2411 | 53.163
[7 empty] | 2412 | 53.183
[7 empty] | 2413 | 53.202
[7 empty] | 2414 | 53.221
[7 empty] | 2415 | 53.24
[7 empty] | 2416 | 53.26
[7 empty] | 2417 | 53.279
[7 empty] | 2418 | 53.298
[7 empty] | 2419 | 53.317
[7 empty] | 2420 | 53.336
[7 empty] | 2421 | 53.355
[7 empty] | 2422 | 53.375
[7 empty] | 2423 | 53.394
[7 empty] | 2424 | 53.413
[7 empty] | 2425 | 53.432
[7 empty] | 2426 | 53.451
[7 empty] | 2427 | 53.47
[7 empty] | 2428 | 53.49
[7 empty] | 2429 | 53.509
[7 empty] | 2430 | 53.528
[7 empty] | 2431 | 53.547
[7 empty] | 2432 | 53.566
[7 empty] | 2433 | 53.585
[7 empty] | 2434 | 53.604
[7 empty] | 2435 | 53.623
[7 empty] | 2436 | 53.643
[7 empty] | 2437 | 53.662
[7 empty] | 2438 | 53.681
[7 empty] | 2439 | 53.7
[7 empty] | 2440 | 53.719
[7 empty] | 2441 | 53.738
[7 empty] | 2442 | 53.757
[7 empty] | 2443 | 53.776
[7 empty] | 2444 | 53.795
[7 empty] | 2445 | 53.814
[7 empty] | 2446 | 53.833
[7 empty] | 2447 | 53.852
[7 empty] | 2448 | 53.871
[7 empty] | 2449 | 53.89
[7 empty] | 2450 | 53.91
[7 empty] | 2451 | 53.929
[7 empty] | 2452 | 53.948
[7 empty] | 2453 | 53.967
[7 empty] | 2454 | 53.986
[7 empty] | 2455 | 54.005
[7 empty] | 2456 | 54.024
[7 empty] | 2457 | 54.043
[7 empty] | 2458 | 54.062
[7 empty] | 2459 | 54.081
[7 empty] | 2460 | 54.1
[7 empty] | 2461 | 54.119
[7 empty] | 2462 | 54.138
[7 empty] | 2463 | 54.157
[7 empty] | 2464 | 54.176
[7 empty] | 2465 | 54.195
[7 empty] | 2466 | 54.214
[7 empty] | 2467 | 54.233
[7 empty] | 2468 | 54.252
[7 empty] | 2469 | 54.271
[7 empty] | 2470 | 54.289
[7 empty] | 2471 | 54.308
[7 empty] | 2472 | 54.327
[7 empty] | 2473 | 54.346
[7 empty] | 2474 | 54.365
[7 empty] | 2475 | 54.384
[7 empty] | 2476 | 54.403
[7 empty] | 2477 | 54.422
[7 empty] | 2478 | 54.441
[7 empty] | 2479 | 54.46
[7 empty] | 2480 | 54.479
[7 empty] | 2481 | 54.498
[7 empty] | 2482 | 54.517
[7 empty] | 2483 | 54.536
[7 empty] | 2484 | 54.554
[7 empty] | 2485 | 54.573
[7 empty] | 2486 | 54.592
[7 empty] | 2487 | 54.611
[7 empty] | 2488 | 54.63
[7 empty] | 2489 | 54.649
[7 empty] | 2490 | 54.668
[7 empty] | 2491 | 54.687
[7 empty] | 2492 | 54.705
[7 empty] | 2493 | 54.724
[7 empty] | 2494 | 54.743
[7 empty] | 2495 | 54.762
[7 empty] | 2496 | 54.781
[7 empty] | 2497 | 54.8
[7 empty] | 2498 | 54.819
[7 empty] | 2499 | 54.837
[7 empty] | 2500 | 54.85

## Type N
[7 empty] | Temp(F) | mV
mV | Temp(F) | [5 empty] | -454 | -4.345
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -453 | -4.345
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -452 | -4.345
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -451 | -4.344
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -450 | -4.344
0.0001 | 32.0071 | 487 | 0 | 0.014 | 32 | 33 | -449 | -4.344
0.5923 | 72.5533 | 527 | 0.584 | 0.599 | 72 | 73 | -448 | -4.343
1.4735 | 130.6562 | 585 | 1.463 | 1.479 | 130 | 131 | -447 | -4.343
2.4638 | 193.0471 | 648 | 2.463 | 2.48 | 193 | 194 | -446 | -4.342
3.5154 | 256.3176 | 711 | 3.51 | 3.527 | 256 | 257 | -445 | -4.342
4.6085 | 319.4722 | 774 | 4.6 | 4.618 | 319 | 320 | -444 | -4.341
5.7339 | 382.1611 | 837 | 5.731 | 5.749 | 382 | 383 | -443 | -4.341
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -442 | -4.34
-50 | #VALUE! | 0 | mV | -4.345 | Temp(F) | -454 | -441 | -4.34
-50 | #VALUE! | 0 | mV | -4.345 | Temp(F) | -454 | -440 | -4.339
-50 | #VALUE! | 0 | mV | -4.345 | Temp(F) | -454 | -439 | -4.338
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -438 | -4.337
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -437 | -4.337
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -436 | -4.336
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -435 | -4.335
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -434 | -4.334
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -433 | -4.333
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -432 | -4.332
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -431 | -4.331
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -430 | -4.33
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -429 | -4.329
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -428 | -4.327
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -427 | -4.326
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -426 | -4.325
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -425 | -4.324
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -424 | -4.322
200 | 9984.6358 | 2827 | 47.513 | 0 | 2372 | 0 | -423 | -4.321
200 | 9984.6358 | 2827 | 47.513 | 0 | 2372 | 0 | -422 | -4.319
200 | 9984.6358 | 2827 | 47.513 | 0 | 2372 | 0 | -421 | -4.318
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -420 | -4.316
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -419 | -4.315
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -418 | -4.313
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -417 | -4.312
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -416 | -4.31
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -415 | -4.308
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -414 | -4.306
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -413 | -4.305
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -412 | -4.303
450 | 22465.4305 | 2827 | 47.513 | 0 | 2372 | 0 | -411 | -4.301
450 | 22465.4305 | 2827 | 47.513 | 0 | 2372 | 0 | -410 | -4.299
450 | 22465.4305 | 2827 | 47.513 | 0 | 2372 | 0 | -409 | -4.297
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -408 | -4.295
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -407 | -4.293
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -406 | -4.291
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -405 | -4.288
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -404 | -4.286
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -403 | -4.284
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -402 | -4.282
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -401 | -4.279
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -400 | -4.277
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -399 | -4.275
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -398 | -4.272
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -397 | -4.27
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -396 | -4.267
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -395 | -4.264
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -394 | -4.262
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -393 | -4.259
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -392 | -4.256
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -391 | -4.254
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -390 | -4.251
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -389 | -4.248
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -388 | -4.245
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -387 | -4.242
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -386 | -4.239
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -385 | -4.236
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -384 | -4.233
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -383 | -4.23
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -382 | -4.226
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -381 | -4.223
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -380 | -4.22
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -379 | -4.217
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -378 | -4.213
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -377 | -4.21
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -376 | -4.206
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -375 | -4.203
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -374 | -4.199
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -373 | -4.196
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -372 | -4.192
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -371 | -4.189
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -370 | -4.185
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -369 | -4.181
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -368 | -4.177
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -367 | -4.174
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -366 | -4.17
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -365 | -4.166
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -364 | -4.162
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -363 | -4.158
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -362 | -4.154
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -361 | -4.15
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -360 | -4.145
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -359 | -4.141
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -358 | -4.137
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -357 | -4.133
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -356 | -4.128
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -355 | -4.124
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -354 | -4.12
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -353 | -4.115
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -352 | -4.111
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -351 | -4.106
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -350 | -4.102
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -349 | -4.097
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -348 | -4.092
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -347 | -4.088
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -346 | -4.083
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -345 | -4.078
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -344 | -4.073
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -343 | -4.068
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -342 | -4.064
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -341 | -4.059
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -340 | -4.054
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -339 | -4.049
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -338 | -4.043
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -337 | -4.038
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -336 | -4.033
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -335 | -4.028
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -334 | -4.023
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -333 | -4.017
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -332 | -4.012
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -331 | -4.007
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -330 | -4.001
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -329 | -3.996
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -328 | -3.99
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -327 | -3.985
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -326 | -3.979
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -325 | -3.974
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -324 | -3.968
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -323 | -3.962
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -322 | -3.957
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -321 | -3.951
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -320 | -3.945
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -319 | -3.939
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -318 | -3.933
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -317 | -3.927
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -316 | -3.921
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -315 | -3.915
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -314 | -3.909
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -313 | -3.903
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -312 | -3.897
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -311 | -3.891
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -310 | -3.884
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -309 | -3.878
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -308 | -3.872
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -307 | -3.866
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -306 | -3.859
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -305 | -3.853
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -304 | -3.846
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -303 | -3.84
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -302 | -3.833
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -301 | -3.827
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -300 | -3.82
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -299 | -3.813
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -298 | -3.807
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -297 | -3.8
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -296 | -3.793
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -295 | -3.786
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -294 | -3.779
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -293 | -3.773
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -292 | -3.766
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -291 | -3.759
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -290 | -3.752
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -289 | -3.745
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -288 | -3.738
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -287 | -3.73
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -286 | -3.723
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -285 | -3.716
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -284 | -3.709
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -283 | -3.702
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -282 | -3.694
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -281 | -3.687
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -280 | -3.679
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -279 | -3.672
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -278 | -3.665
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -277 | -3.657
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -276 | -3.65
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -275 | -3.642
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -274 | -3.634
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -273 | -3.627
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -272 | -3.619
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -271 | -3.611
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -270 | -3.604
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -269 | -3.596
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -268 | -3.588
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -267 | -3.58
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -266 | -3.572
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -265 | -3.564
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -264 | -3.556
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -263 | -3.548
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -262 | -3.54
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -261 | -3.532
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -260 | -3.524
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -259 | -3.516
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -258 | -3.508
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -257 | -3.499
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -256 | -3.491
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -255 | -3.483
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -254 | -3.474
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -253 | -3.466
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -252 | -3.458
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -251 | -3.449
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -250 | -3.441
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -249 | -3.432
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -248 | -3.424
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -247 | -3.415
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -246 | -3.407
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -245 | -3.398
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -244 | -3.389
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -243 | -3.38
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -242 | -3.372
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -241 | -3.363
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -240 | -3.354
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -239 | -3.345
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -238 | -3.336
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -237 | -3.327
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -236 | -3.318
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -235 | -3.309
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -234 | -3.3
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -233 | -3.291
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -232 | -3.282
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -231 | -3.273
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -230 | -3.264
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -229 | -3.255
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -228 | -3.246
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -227 | -3.236
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -226 | -3.227
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -225 | -3.218
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -224 | -3.208
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -223 | -3.199
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -222 | -3.189
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -221 | -3.18
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -220 | -3.171
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -219 | -3.161
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -218 | -3.151
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -217 | -3.142
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -216 | -3.132
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -215 | -3.123
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -214 | -3.113
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -213 | -3.103
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -212 | -3.093
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -211 | -3.084
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -210 | -3.074
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -209 | -3.064
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -208 | -3.054
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -207 | -3.044
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -206 | -3.034
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -205 | -3.024
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -204 | -3.014
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -203 | -3.004
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -202 | -2.994
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -201 | -2.984
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -200 | -2.974
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -199 | -2.964
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -198 | -2.954
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -197 | -2.943
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -196 | -2.933
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -195 | -2.923
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -194 | -2.912
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -193 | -2.902
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -192 | -2.892
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -191 | -2.881
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -190 | -2.871
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -189 | -2.86
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -188 | -2.85
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -187 | -2.839
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -186 | -2.829
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -185 | -2.818
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -184 | -2.808
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -183 | -2.797
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -182 | -2.786
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -181 | -2.776
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -180 | -2.765
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -179 | -2.754
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -178 | -2.743
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -177 | -2.733
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -176 | -2.722
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -175 | -2.711
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -174 | -2.7
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -173 | -2.689
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -172 | -2.678
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -171 | -2.667
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -170 | -2.656
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -169 | -2.645
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -168 | -2.634
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -167 | -2.623
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -166 | -2.612
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -165 | -2.601
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -164 | -2.589
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -163 | -2.578
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -162 | -2.567
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -161 | -2.556
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -160 | -2.544
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -159 | -2.533
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -158 | -2.522
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -157 | -2.51
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -156 | -2.499
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -155 | -2.488
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -154 | -2.476
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -153 | -2.465
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -152 | -2.453
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -151 | -2.442
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -150 | -2.43
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -149 | -2.418
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -148 | -2.407
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -147 | -2.395
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -146 | -2.384
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -145 | -2.372
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -144 | -2.36
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -143 | -2.348
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -142 | -2.337
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -141 | -2.325
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -140 | -2.313
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -139 | -2.301
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -138 | -2.289
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -137 | -2.277
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -136 | -2.265
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -135 | -2.254
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -134 | -2.242
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -133 | -2.23
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -132 | -2.218
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -131 | -2.206
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -130 | -2.193
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -129 | -2.181
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -128 | -2.169
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -127 | -2.157
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -126 | -2.145
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -125 | -2.133
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -124 | -2.121
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -123 | -2.108
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -122 | -2.096
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -121 | -2.084
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -120 | -2.072
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -119 | -2.059
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -118 | -2.047
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -117 | -2.035
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -116 | -2.022
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -115 | -2.01
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -114 | -1.997
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -113 | -1.985
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -112 | -1.972
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -111 | -1.96
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -110 | -1.947
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -109 | -1.935
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -108 | -1.922
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -107 | -1.91
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -106 | -1.897
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -105 | -1.884
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -104 | -1.872
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -103 | -1.859
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -102 | -1.846
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -101 | -1.834
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -100 | -1.821
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -99 | -1.808
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -98 | -1.795
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -97 | -1.783
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -96 | -1.77
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -95 | -1.757
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -94 | -1.744
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -93 | -1.731
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -92 | -1.718
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -91 | -1.705
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -90 | -1.692
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -89 | -1.679
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -88 | -1.666
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -87 | -1.653
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -86 | -1.64
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -85 | -1.627
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -84 | -1.614
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -83 | -1.601
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -82 | -1.588
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -81 | -1.575
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -80 | -1.562
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -79 | -1.549
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -78 | -1.536
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -77 | -1.522
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -76 | -1.509
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -75 | -1.496
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -74 | -1.483
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -73 | -1.47
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -72 | -1.456
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -71 | -1.443
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -70 | -1.43
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -69 | -1.416
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -68 | -1.403
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -67 | -1.39
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -66 | -1.376
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -65 | -1.363
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -64 | -1.349
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -63 | -1.336
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -62 | -1.323
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -61 | -1.309
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -60 | -1.296
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -59 | -1.282
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -58 | -1.269
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -57 | -1.255
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -56 | -1.242
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -55 | -1.228
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -54 | -1.214
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -53 | -1.201
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -52 | -1.187
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -51 | -1.174
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -50 | -1.16
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -49 | -1.146
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -48 | -1.133
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -47 | -1.119
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -46 | -1.105
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -45 | -1.092
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -44 | -1.078
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -43 | -1.064
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -42 | -1.05
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -41 | -1.037
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -40 | -1.023
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -39 | -1.009
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -38 | -0.995
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -37 | -0.981
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -36 | -0.967
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -35 | -0.954
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -34 | -0.94
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -33 | -0.926
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -32 | -0.912
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -31 | -0.898
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -30 | -0.884
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -29 | -0.87
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -28 | -0.856
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -27 | -0.842
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -26 | -0.828
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -25 | -0.814
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -24 | -0.8
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -23 | -0.786
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -22 | -0.772
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -21 | -0.758
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -20 | -0.744
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -19 | -0.73
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -18 | -0.716
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -17 | -0.702
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -16 | -0.688
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -15 | -0.674
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -14 | -0.66
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -13 | -0.646
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -12 | -0.632
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -11 | -0.617
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -10 | -0.603
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -9 | -0.589
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -8 | -0.575
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -7 | -0.561
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -6 | -0.546
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -5 | -0.532
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -4 | -0.518
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -3 | -0.504
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -2 | -0.49
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | -1 | -0.475
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 0 | -0.461
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1 | -0.447
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2 | -0.433
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 3 | -0.418
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 4 | -0.404
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 5 | -0.39
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 6 | -0.375
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 7 | -0.361
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 8 | -0.347
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 9 | -0.332
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 10 | -0.318
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 11 | -0.304
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 12 | -0.289
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 13 | -0.275
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 14 | -0.26
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 15 | -0.246
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 16 | -0.232
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 17 | -0.217
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 18 | -0.203
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 19 | -0.188
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 20 | -0.174
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 21 | -0.159
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 22 | -0.145
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 23 | -0.131
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 24 | -0.116
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 25 | -0.102
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 26 | -0.087
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 27 | -0.073
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 28 | -0.058
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 29 | -0.044
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 30 | -0.029
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 31 | -0.015
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 32 | 0
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 33 | 0.014
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 34 | 0.029
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 35 | 0.043
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 36 | 0.058
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 37 | 0.072
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 38 | 0.087
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 39 | 0.101
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 40 | 0.116
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 41 | 0.13
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 42 | 0.145
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 43 | 0.159
[7 empty] | 44 | 0.174
[7 empty] | 45 | 0.188
[7 empty] | 46 | 0.203
[7 empty] | 47 | 0.217
[7 empty] | 48 | 0.232
[7 empty] | 49 | 0.246
[7 empty] | 50 | 0.261
[7 empty] | 51 | 0.275
[7 empty] | 52 | 0.29
[7 empty] | 53 | 0.305
[7 empty] | 54 | 0.319
[7 empty] | 55 | 0.334
[7 empty] | 56 | 0.349
[7 empty] | 57 | 0.363
[7 empty] | 58 | 0.378
[7 empty] | 59 | 0.393
[7 empty] | 60 | 0.407
[7 empty] | 61 | 0.422
[7 empty] | 62 | 0.437
[7 empty] | 63 | 0.451
[7 empty] | 64 | 0.466
[7 empty] | 65 | 0.481
[7 empty] | 66 | 0.496
[7 empty] | 67 | 0.51
[7 empty] | 68 | 0.525
[7 empty] | 69 | 0.54
[7 empty] | 70 | 0.555
[7 empty] | 71 | 0.57
[7 empty] | 72 | 0.584
[7 empty] | 73 | 0.599
[7 empty] | 74 | 0.614
[7 empty] | 75 | 0.629
[7 empty] | 76 | 0.644
[7 empty] | 77 | 0.659
[7 empty] | 78 | 0.674
[7 empty] | 79 | 0.688
[7 empty] | 80 | 0.703
[7 empty] | 81 | 0.718
[7 empty] | 82 | 0.733
[7 empty] | 83 | 0.748
[7 empty] | 84 | 0.763
[7 empty] | 85 | 0.778
[7 empty] | 86 | 0.793
[7 empty] | 87 | 0.808
[7 empty] | 88 | 0.823
[7 empty] | 89 | 0.838
[7 empty] | 90 | 0.853
[7 empty] | 91 | 0.868
[7 empty] | 92 | 0.883
[7 empty] | 93 | 0.898
[7 empty] | 94 | 0.913
[7 empty] | 95 | 0.928
[7 empty] | 96 | 0.943
[7 empty] | 97 | 0.958
[7 empty] | 98 | 0.974
[7 empty] | 99 | 0.989
[7 empty] | 100 | 1.004
[7 empty] | 101 | 1.019
[7 empty] | 102 | 1.034
[7 empty] | 103 | 1.049
[7 empty] | 104 | 1.065
[7 empty] | 105 | 1.08
[7 empty] | 106 | 1.095
[7 empty] | 107 | 1.11
[7 empty] | 108 | 1.125
[7 empty] | 109 | 1.141
[7 empty] | 110 | 1.156
[7 empty] | 111 | 1.171
[7 empty] | 112 | 1.186
[7 empty] | 113 | 1.202
[7 empty] | 114 | 1.217
[7 empty] | 115 | 1.232
[7 empty] | 116 | 1.248
[7 empty] | 117 | 1.263
[7 empty] | 118 | 1.278
[7 empty] | 119 | 1.294
[7 empty] | 120 | 1.309
[7 empty] | 121 | 1.324
[7 empty] | 122 | 1.34
[7 empty] | 123 | 1.355
[7 empty] | 124 | 1.371
[7 empty] | 125 | 1.386
[7 empty] | 126 | 1.402
[7 empty] | 127 | 1.417
[7 empty] | 128 | 1.432
[7 empty] | 129 | 1.448
[7 empty] | 130 | 1.463
[7 empty] | 131 | 1.479
[7 empty] | 132 | 1.494
[7 empty] | 133 | 1.51
[7 empty] | 134 | 1.525
[7 empty] | 135 | 1.541
[7 empty] | 136 | 1.557
[7 empty] | 137 | 1.572
[7 empty] | 138 | 1.588
[7 empty] | 139 | 1.603
[7 empty] | 140 | 1.619
[7 empty] | 141 | 1.635
[7 empty] | 142 | 1.65
[7 empty] | 143 | 1.666
[7 empty] | 144 | 1.682
[7 empty] | 145 | 1.697
[7 empty] | 146 | 1.713
[7 empty] | 147 | 1.729
[7 empty] | 148 | 1.744
[7 empty] | 149 | 1.76
[7 empty] | 150 | 1.776
[7 empty] | 151 | 1.791
[7 empty] | 152 | 1.807
[7 empty] | 153 | 1.823
[7 empty] | 154 | 1.839
[7 empty] | 155 | 1.855
[7 empty] | 156 | 1.87
[7 empty] | 157 | 1.886
[7 empty] | 158 | 1.902
[7 empty] | 159 | 1.918
[7 empty] | 160 | 1.934
[7 empty] | 161 | 1.95
[7 empty] | 162 | 1.965
[7 empty] | 163 | 1.981
[7 empty] | 164 | 1.997
[7 empty] | 165 | 2.013
[7 empty] | 166 | 2.029
[7 empty] | 167 | 2.045
[7 empty] | 168 | 2.061
[7 empty] | 169 | 2.077
[7 empty] | 170 | 2.093
[7 empty] | 171 | 2.109
[7 empty] | 172 | 2.125
[7 empty] | 173 | 2.141
[7 empty] | 174 | 2.157
[7 empty] | 175 | 2.173
[7 empty] | 176 | 2.189
[7 empty] | 177 | 2.205
[7 empty] | 178 | 2.221
[7 empty] | 179 | 2.237
[7 empty] | 180 | 2.253
[7 empty] | 181 | 2.269
[7 empty] | 182 | 2.285
[7 empty] | 183 | 2.301
[7 empty] | 184 | 2.318
[7 empty] | 185 | 2.334
[7 empty] | 186 | 2.35
[7 empty] | 187 | 2.366
[7 empty] | 188 | 2.382
[7 empty] | 189 | 2.398
[7 empty] | 190 | 2.415
[7 empty] | 191 | 2.431
[7 empty] | 192 | 2.447
[7 empty] | 193 | 2.463
[7 empty] | 194 | 2.48
[7 empty] | 195 | 2.496
[7 empty] | 196 | 2.512
[7 empty] | 197 | 2.528
[7 empty] | 198 | 2.545
[7 empty] | 199 | 2.561
[7 empty] | 200 | 2.577
[7 empty] | 201 | 2.594
[7 empty] | 202 | 2.61
[7 empty] | 203 | 2.626
[7 empty] | 204 | 2.643
[7 empty] | 205 | 2.659
[7 empty] | 206 | 2.676
[7 empty] | 207 | 2.692
[7 empty] | 208 | 2.708
[7 empty] | 209 | 2.725
[7 empty] | 210 | 2.741
[7 empty] | 211 | 2.758
[7 empty] | 212 | 2.774
[7 empty] | 213 | 2.791
[7 empty] | 214 | 2.807
[7 empty] | 215 | 2.824
[7 empty] | 216 | 2.84
[7 empty] | 217 | 2.857
[7 empty] | 218 | 2.873
[7 empty] | 219 | 2.89
[7 empty] | 220 | 2.906
[7 empty] | 221 | 2.923
[7 empty] | 222 | 2.939
[7 empty] | 223 | 2.956
[7 empty] | 224 | 2.973
[7 empty] | 225 | 2.989
[7 empty] | 226 | 3.006
[7 empty] | 227 | 3.022
[7 empty] | 228 | 3.039
[7 empty] | 229 | 3.056
[7 empty] | 230 | 3.072
[7 empty] | 231 | 3.089
[7 empty] | 232 | 3.106
[7 empty] | 233 | 3.123
[7 empty] | 234 | 3.139
[7 empty] | 235 | 3.156
[7 empty] | 236 | 3.173
[7 empty] | 237 | 3.189
[7 empty] | 238 | 3.206
[7 empty] | 239 | 3.223
[7 empty] | 240 | 3.24
[7 empty] | 241 | 3.257
[7 empty] | 242 | 3.273
[7 empty] | 243 | 3.29
[7 empty] | 244 | 3.307
[7 empty] | 245 | 3.324
[7 empty] | 246 | 3.341
[7 empty] | 247 | 3.358
[7 empty] | 248 | 3.374
[7 empty] | 249 | 3.391
[7 empty] | 250 | 3.408
[7 empty] | 251 | 3.425
[7 empty] | 252 | 3.442
[7 empty] | 253 | 3.459
[7 empty] | 254 | 3.476
[7 empty] | 255 | 3.493
[7 empty] | 256 | 3.51
[7 empty] | 257 | 3.527
[7 empty] | 258 | 3.544
[7 empty] | 259 | 3.561
[7 empty] | 260 | 3.578
[7 empty] | 261 | 3.595
[7 empty] | 262 | 3.612
[7 empty] | 263 | 3.629
[7 empty] | 264 | 3.646
[7 empty] | 265 | 3.663
[7 empty] | 266 | 3.68
[7 empty] | 267 | 3.697
[7 empty] | 268 | 3.714
[7 empty] | 269 | 3.731
[7 empty] | 270 | 3.748
[7 empty] | 271 | 3.766
[7 empty] | 272 | 3.783
[7 empty] | 273 | 3.8
[7 empty] | 274 | 3.817
[7 empty] | 275 | 3.834
[7 empty] | 276 | 3.851
[7 empty] | 277 | 3.869
[7 empty] | 278 | 3.886
[7 empty] | 279 | 3.903
[7 empty] | 280 | 3.92
[7 empty] | 281 | 3.937
[7 empty] | 282 | 3.955
[7 empty] | 283 | 3.972
[7 empty] | 284 | 3.989
[7 empty] | 285 | 4.007
[7 empty] | 286 | 4.024
[7 empty] | 287 | 4.041
[7 empty] | 288 | 4.058
[7 empty] | 289 | 4.076
[7 empty] | 290 | 4.093
[7 empty] | 291 | 4.11
[7 empty] | 292 | 4.128
[7 empty] | 293 | 4.145
[7 empty] | 294 | 4.162
[7 empty] | 295 | 4.18
[7 empty] | 296 | 4.197
[7 empty] | 297 | 4.215
[7 empty] | 298 | 4.232
[7 empty] | 299 | 4.25
[7 empty] | 300 | 4.267
[7 empty] | 301 | 4.284
[7 empty] | 302 | 4.302
[7 empty] | 303 | 4.319
[7 empty] | 304 | 4.337
[7 empty] | 305 | 4.354
[7 empty] | 306 | 4.372
[7 empty] | 307 | 4.389
[7 empty] | 308 | 4.407
[7 empty] | 309 | 4.424
[7 empty] | 310 | 4.442
[7 empty] | 311 | 4.459
[7 empty] | 312 | 4.477
[7 empty] | 313 | 4.495
[7 empty] | 314 | 4.512
[7 empty] | 315 | 4.53
[7 empty] | 316 | 4.547
[7 empty] | 317 | 4.565
[7 empty] | 318 | 4.583
[7 empty] | 319 | 4.6
[7 empty] | 320 | 4.618
[7 empty] | 321 | 4.635
[7 empty] | 322 | 4.653
[7 empty] | 323 | 4.671
[7 empty] | 324 | 4.688
[7 empty] | 325 | 4.706
[7 empty] | 326 | 4.724
[7 empty] | 327 | 4.742
[7 empty] | 328 | 4.759
[7 empty] | 329 | 4.777
[7 empty] | 330 | 4.795
[7 empty] | 331 | 4.813
[7 empty] | 332 | 4.83
[7 empty] | 333 | 4.848
[7 empty] | 334 | 4.866
[7 empty] | 335 | 4.884
[7 empty] | 336 | 4.901
[7 empty] | 337 | 4.919
[7 empty] | 338 | 4.937
[7 empty] | 339 | 4.955
[7 empty] | 340 | 4.973
[7 empty] | 341 | 4.991
[7 empty] | 342 | 5.008
[7 empty] | 343 | 5.026
[7 empty] | 344 | 5.044
[7 empty] | 345 | 5.062
[7 empty] | 346 | 5.08
[7 empty] | 347 | 5.098
[7 empty] | 348 | 5.116
[7 empty] | 349 | 5.134
[7 empty] | 350 | 5.152
[7 empty] | 351 | 5.17
[7 empty] | 352 | 5.188
[7 empty] | 353 | 5.206
[7 empty] | 354 | 5.224
[7 empty] | 355 | 5.241
[7 empty] | 356 | 5.259
[7 empty] | 357 | 5.277
[7 empty] | 358 | 5.295
[7 empty] | 359 | 5.314
[7 empty] | 360 | 5.332
[7 empty] | 361 | 5.35
[7 empty] | 362 | 5.368
[7 empty] | 363 | 5.386
[7 empty] | 364 | 5.404
[7 empty] | 365 | 5.422
[7 empty] | 366 | 5.44
[7 empty] | 367 | 5.458
[7 empty] | 368 | 5.476
[7 empty] | 369 | 5.494
[7 empty] | 370 | 5.512
[7 empty] | 371 | 5.531
[7 empty] | 372 | 5.549
[7 empty] | 373 | 5.567
[7 empty] | 374 | 5.585
[7 empty] | 375 | 5.603
[7 empty] | 376 | 5.621
[7 empty] | 377 | 5.639
[7 empty] | 378 | 5.658
[7 empty] | 379 | 5.676
[7 empty] | 380 | 5.694
[7 empty] | 381 | 5.712
[7 empty] | 382 | 5.731
[7 empty] | 383 | 5.749
[7 empty] | 384 | 5.767
[7 empty] | 385 | 5.785
[7 empty] | 386 | 5.804
[7 empty] | 387 | 5.822
[7 empty] | 388 | 5.84
[7 empty] | 389 | 5.858
[7 empty] | 390 | 5.877
[7 empty] | 391 | 5.895
[7 empty] | 392 | 5.913
[7 empty] | 393 | 5.932
[7 empty] | 394 | 5.95
[7 empty] | 395 | 5.968
[7 empty] | 396 | 5.987
[7 empty] | 397 | 6.005
[7 empty] | 398 | 6.024
[7 empty] | 399 | 6.042
[7 empty] | 400 | 6.06
[7 empty] | 401 | 6.079
[7 empty] | 402 | 6.097
[7 empty] | 403 | 6.116
[7 empty] | 404 | 6.134
[7 empty] | 405 | 6.152
[7 empty] | 406 | 6.171
[7 empty] | 407 | 6.189
[7 empty] | 408 | 6.208
[7 empty] | 409 | 6.226
[7 empty] | 410 | 6.245
[7 empty] | 411 | 6.263
[7 empty] | 412 | 6.282
[7 empty] | 413 | 6.3
[7 empty] | 414 | 6.319
[7 empty] | 415 | 6.337
[7 empty] | 416 | 6.356
[7 empty] | 417 | 6.374
[7 empty] | 418 | 6.393
[7 empty] | 419 | 6.411
[7 empty] | 420 | 6.43
[7 empty] | 421 | 6.449
[7 empty] | 422 | 6.467
[7 empty] | 423 | 6.486
[7 empty] | 424 | 6.504
[7 empty] | 425 | 6.523
[7 empty] | 426 | 6.542
[7 empty] | 427 | 6.56
[7 empty] | 428 | 6.579
[7 empty] | 429 | 6.597
[7 empty] | 430 | 6.616
[7 empty] | 431 | 6.635
[7 empty] | 432 | 6.653
[7 empty] | 433 | 6.672
[7 empty] | 434 | 6.691
[7 empty] | 435 | 6.71
[7 empty] | 436 | 6.728
[7 empty] | 437 | 6.747
[7 empty] | 438 | 6.766
[7 empty] | 439 | 6.784
[7 empty] | 440 | 6.803
[7 empty] | 441 | 6.822
[7 empty] | 442 | 6.841
[7 empty] | 443 | 6.859
[7 empty] | 444 | 6.878
[7 empty] | 445 | 6.897
[7 empty] | 446 | 6.916
[7 empty] | 447 | 6.934
[7 empty] | 448 | 6.953
[7 empty] | 449 | 6.972
[7 empty] | 450 | 6.991
[7 empty] | 451 | 7.01
[7 empty] | 452 | 7.029
[7 empty] | 453 | 7.047
[7 empty] | 454 | 7.066
[7 empty] | 455 | 7.085
[7 empty] | 456 | 7.104
[7 empty] | 457 | 7.123
[7 empty] | 458 | 7.142
[7 empty] | 459 | 7.161
[7 empty] | 460 | 7.179
[7 empty] | 461 | 7.198
[7 empty] | 462 | 7.217
[7 empty] | 463 | 7.236
[7 empty] | 464 | 7.255
[7 empty] | 465 | 7.274
[7 empty] | 466 | 7.293
[7 empty] | 467 | 7.312
[7 empty] | 468 | 7.331
[7 empty] | 469 | 7.35
[7 empty] | 470 | 7.369
[7 empty] | 471 | 7.388
[7 empty] | 472 | 7.407
[7 empty] | 473 | 7.426
[7 empty] | 474 | 7.445
[7 empty] | 475 | 7.464
[7 empty] | 476 | 7.483
[7 empty] | 477 | 7.502
[7 empty] | 478 | 7.521
[7 empty] | 479 | 7.54
[7 empty] | 480 | 7.559
[7 empty] | 481 | 7.578
[7 empty] | 482 | 7.597
[7 empty] | 483 | 7.616
[7 empty] | 484 | 7.635
[7 empty] | 485 | 7.654
[7 empty] | 486 | 7.673
[7 empty] | 487 | 7.692
[7 empty] | 488 | 7.711
[7 empty] | 489 | 7.731
[7 empty] | 490 | 7.75
[7 empty] | 491 | 7.769
[7 empty] | 492 | 7.788
[7 empty] | 493 | 7.807
[7 empty] | 494 | 7.826
[7 empty] | 495 | 7.845
[7 empty] | 496 | 7.865
[7 empty] | 497 | 7.884
[7 empty] | 498 | 7.903
[7 empty] | 499 | 7.922
[7 empty] | 500 | 7.941
[7 empty] | 501 | 7.96
[7 empty] | 502 | 7.98
[7 empty] | 503 | 7.999
[7 empty] | 504 | 8.018
[7 empty] | 505 | 8.037
[7 empty] | 506 | 8.057
[7 empty] | 507 | 8.076
[7 empty] | 508 | 8.095
[7 empty] | 509 | 8.114
[7 empty] | 510 | 8.134
[7 empty] | 511 | 8.153
[7 empty] | 512 | 8.172
[7 empty] | 513 | 8.191
[7 empty] | 514 | 8.211
[7 empty] | 515 | 8.23
[7 empty] | 516 | 8.249
[7 empty] | 517 | 8.269
[7 empty] | 518 | 8.288
[7 empty] | 519 | 8.307
[7 empty] | 520 | 8.327
[7 empty] | 521 | 8.346
[7 empty] | 522 | 8.365
[7 empty] | 523 | 8.385
[7 empty] | 524 | 8.404
[7 empty] | 525 | 8.423
[7 empty] | 526 | 8.443
[7 empty] | 527 | 8.462
[7 empty] | 528 | 8.482
[7 empty] | 529 | 8.501
[7 empty] | 530 | 8.52
[7 empty] | 531 | 8.54
[7 empty] | 532 | 8.559
[7 empty] | 533 | 8.579
[7 empty] | 534 | 8.598
[7 empty] | 535 | 8.617
[7 empty] | 536 | 8.637
[7 empty] | 537 | 8.656
[7 empty] | 538 | 8.676
[7 empty] | 539 | 8.695
[7 empty] | 540 | 8.715
[7 empty] | 541 | 8.734
[7 empty] | 542 | 8.754
[7 empty] | 543 | 8.773
[7 empty] | 544 | 8.793
[7 empty] | 545 | 8.812
[7 empty] | 546 | 8.832
[7 empty] | 547 | 8.851
[7 empty] | 548 | 8.871
[7 empty] | 549 | 8.89
[7 empty] | 550 | 8.91
[7 empty] | 551 | 8.929
[7 empty] | 552 | 8.949
[7 empty] | 553 | 8.968
[7 empty] | 554 | 8.988
[7 empty] | 555 | 9.008
[7 empty] | 556 | 9.027
[7 empty] | 557 | 9.047
[7 empty] | 558 | 9.066
[7 empty] | 559 | 9.086
[7 empty] | 560 | 9.105
[7 empty] | 561 | 9.125
[7 empty] | 562 | 9.145
[7 empty] | 563 | 9.164
[7 empty] | 564 | 9.184
[7 empty] | 565 | 9.204
[7 empty] | 566 | 9.223
[7 empty] | 567 | 9.243
[7 empty] | 568 | 9.262
[7 empty] | 569 | 9.282
[7 empty] | 570 | 9.302
[7 empty] | 571 | 9.321
[7 empty] | 572 | 9.341
[7 empty] | 573 | 9.361
[7 empty] | 574 | 9.381
[7 empty] | 575 | 9.4
[7 empty] | 576 | 9.42
[7 empty] | 577 | 9.44
[7 empty] | 578 | 9.459
[7 empty] | 579 | 9.479
[7 empty] | 580 | 9.499
[7 empty] | 581 | 9.519
[7 empty] | 582 | 9.538
[7 empty] | 583 | 9.558
[7 empty] | 584 | 9.578
[7 empty] | 585 | 9.598
[7 empty] | 586 | 9.617
[7 empty] | 587 | 9.637
[7 empty] | 588 | 9.657
[7 empty] | 589 | 9.677
[7 empty] | 590 | 9.696
[7 empty] | 591 | 9.716
[7 empty] | 592 | 9.736
[7 empty] | 593 | 9.756
[7 empty] | 594 | 9.776
[7 empty] | 595 | 9.795
[7 empty] | 596 | 9.815
[7 empty] | 597 | 9.835
[7 empty] | 598 | 9.855
[7 empty] | 599 | 9.875
[7 empty] | 600 | 9.895
[7 empty] | 601 | 9.914
[7 empty] | 602 | 9.934
[7 empty] | 603 | 9.954
[7 empty] | 604 | 9.974
[7 empty] | 605 | 9.994
[7 empty] | 606 | 10.014
[7 empty] | 607 | 10.034
[7 empty] | 608 | 10.054
[7 empty] | 609 | 10.073
[7 empty] | 610 | 10.093
[7 empty] | 611 | 10.113
[7 empty] | 612 | 10.133
[7 empty] | 613 | 10.153
[7 empty] | 614 | 10.173
[7 empty] | 615 | 10.193
[7 empty] | 616 | 10.213
[7 empty] | 617 | 10.233
[7 empty] | 618 | 10.253
[7 empty] | 619 | 10.273
[7 empty] | 620 | 10.293
[7 empty] | 621 | 10.313
[7 empty] | 622 | 10.333
[7 empty] | 623 | 10.353
[7 empty] | 624 | 10.373
[7 empty] | 625 | 10.393
[7 empty] | 626 | 10.413
[7 empty] | 627 | 10.433
[7 empty] | 628 | 10.453
[7 empty] | 629 | 10.473
[7 empty] | 630 | 10.493
[7 empty] | 631 | 10.513
[7 empty] | 632 | 10.533
[7 empty] | 633 | 10.553
[7 empty] | 634 | 10.573
[7 empty] | 635 | 10.593
[7 empty] | 636 | 10.613
[7 empty] | 637 | 10.633
[7 empty] | 638 | 10.653
[7 empty] | 639 | 10.673
[7 empty] | 640 | 10.693
[7 empty] | 641 | 10.713
[7 empty] | 642 | 10.733
[7 empty] | 643 | 10.753
[7 empty] | 644 | 10.774
[7 empty] | 645 | 10.794
[7 empty] | 646 | 10.814
[7 empty] | 647 | 10.834
[7 empty] | 648 | 10.854
[7 empty] | 649 | 10.874
[7 empty] | 650 | 10.894
[7 empty] | 651 | 10.914
[7 empty] | 652 | 10.934
[7 empty] | 653 | 10.955
[7 empty] | 654 | 10.975
[7 empty] | 655 | 10.995
[7 empty] | 656 | 11.015
[7 empty] | 657 | 11.035
[7 empty] | 658 | 11.055
[7 empty] | 659 | 11.076
[7 empty] | 660 | 11.096
[7 empty] | 661 | 11.116
[7 empty] | 662 | 11.136
[7 empty] | 663 | 11.156
[7 empty] | 664 | 11.177
[7 empty] | 665 | 11.197
[7 empty] | 666 | 11.217
[7 empty] | 667 | 11.237
[7 empty] | 668 | 11.257
[7 empty] | 669 | 11.278
[7 empty] | 670 | 11.298
[7 empty] | 671 | 11.318
[7 empty] | 672 | 11.338
[7 empty] | 673 | 11.359
[7 empty] | 674 | 11.379
[7 empty] | 675 | 11.399
[7 empty] | 676 | 11.419
[7 empty] | 677 | 11.44
[7 empty] | 678 | 11.46
[7 empty] | 679 | 11.48
[7 empty] | 680 | 11.501
[7 empty] | 681 | 11.521
[7 empty] | 682 | 11.541
[7 empty] | 683 | 11.561
[7 empty] | 684 | 11.582
[7 empty] | 685 | 11.602
[7 empty] | 686 | 11.622
[7 empty] | 687 | 11.643
[7 empty] | 688 | 11.663
[7 empty] | 689 | 11.683
[7 empty] | 690 | 11.704
[7 empty] | 691 | 11.724
[7 empty] | 692 | 11.744
[7 empty] | 693 | 11.765
[7 empty] | 694 | 11.785
[7 empty] | 695 | 11.805
[7 empty] | 696 | 11.826
[7 empty] | 697 | 11.846
[7 empty] | 698 | 11.867
[7 empty] | 699 | 11.887
[7 empty] | 700 | 11.907
[7 empty] | 701 | 11.928
[7 empty] | 702 | 11.948
[7 empty] | 703 | 11.968
[7 empty] | 704 | 11.989
[7 empty] | 705 | 12.009
[7 empty] | 706 | 12.03
[7 empty] | 707 | 12.05
[7 empty] | 708 | 12.071
[7 empty] | 709 | 12.091
[7 empty] | 710 | 12.111
[7 empty] | 711 | 12.132
[7 empty] | 712 | 12.152
[7 empty] | 713 | 12.173
[7 empty] | 714 | 12.193
[7 empty] | 715 | 12.214
[7 empty] | 716 | 12.234
[7 empty] | 717 | 12.255
[7 empty] | 718 | 12.275
[7 empty] | 719 | 12.295
[7 empty] | 720 | 12.316
[7 empty] | 721 | 12.336
[7 empty] | 722 | 12.357
[7 empty] | 723 | 12.377
[7 empty] | 724 | 12.398
[7 empty] | 725 | 12.418
[7 empty] | 726 | 12.439
[7 empty] | 727 | 12.459
[7 empty] | 728 | 12.48
[7 empty] | 729 | 12.5
[7 empty] | 730 | 12.521
[7 empty] | 731 | 12.542
[7 empty] | 732 | 12.562
[7 empty] | 733 | 12.583
[7 empty] | 734 | 12.603
[7 empty] | 735 | 12.624
[7 empty] | 736 | 12.644
[7 empty] | 737 | 12.665
[7 empty] | 738 | 12.685
[7 empty] | 739 | 12.706
[7 empty] | 740 | 12.726
[7 empty] | 741 | 12.747
[7 empty] | 742 | 12.768
[7 empty] | 743 | 12.788
[7 empty] | 744 | 12.809
[7 empty] | 745 | 12.829
[7 empty] | 746 | 12.85
[7 empty] | 747 | 12.871
[7 empty] | 748 | 12.891
[7 empty] | 749 | 12.912
[7 empty] | 750 | 12.932
[7 empty] | 751 | 12.953
[7 empty] | 752 | 12.974
[7 empty] | 753 | 12.994
[7 empty] | 754 | 13.015
[7 empty] | 755 | 13.036
[7 empty] | 756 | 13.056
[7 empty] | 757 | 13.077
[7 empty] | 758 | 13.098
[7 empty] | 759 | 13.118
[7 empty] | 760 | 13.139
[7 empty] | 761 | 13.159
[7 empty] | 762 | 13.18
[7 empty] | 763 | 13.201
[7 empty] | 764 | 13.221
[7 empty] | 765 | 13.242
[7 empty] | 766 | 13.263
[7 empty] | 767 | 13.284
[7 empty] | 768 | 13.304
[7 empty] | 769 | 13.325
[7 empty] | 770 | 13.346
[7 empty] | 771 | 13.366
[7 empty] | 772 | 13.387
[7 empty] | 773 | 13.408
[7 empty] | 774 | 13.428
[7 empty] | 775 | 13.449
[7 empty] | 776 | 13.47
[7 empty] | 777 | 13.491
[7 empty] | 778 | 13.511
[7 empty] | 779 | 13.532
[7 empty] | 780 | 13.553
[7 empty] | 781 | 13.574
[7 empty] | 782 | 13.594
[7 empty] | 783 | 13.615
[7 empty] | 784 | 13.636
[7 empty] | 785 | 13.657
[7 empty] | 786 | 13.677
[7 empty] | 787 | 13.698
[7 empty] | 788 | 13.719
[7 empty] | 789 | 13.74
[7 empty] | 790 | 13.76
[7 empty] | 791 | 13.781
[7 empty] | 792 | 13.802
[7 empty] | 793 | 13.823
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 794 | 13.844
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 795 | 13.864
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 796 | 13.885
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 797 | 13.906
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 798 | 13.927
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 799 | 13.948
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 800 | 13.969
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 801 | 13.989
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 802 | 14.01
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 803 | 14.031
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 804 | 14.052
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 805 | 14.073
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 806 | 14.094
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 807 | 14.114
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 808 | 14.135
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 809 | 14.156
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 810 | 14.177
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 811 | 14.198
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 812 | 14.219
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 813 | 14.24
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 814 | 14.26
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 815 | 14.281
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 816 | 14.302
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 817 | 14.323
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 818 | 14.344
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 819 | 14.365
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 820 | 14.386
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 821 | 14.407
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 822 | 14.428
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 823 | 14.448
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 824 | 14.469
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 825 | 14.49
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 826 | 14.511
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 827 | 14.532
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 828 | 14.553
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 829 | 14.574
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 830 | 14.595
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 831 | 14.616
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 832 | 14.637
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 833 | 14.658
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 834 | 14.679
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 835 | 14.7
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 836 | 14.721
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 837 | 14.742
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 838 | 14.763
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 839 | 14.784
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 840 | 14.804
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 841 | 14.825
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 842 | 14.846
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 843 | 14.867
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 844 | 14.888
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 845 | 14.909
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 846 | 14.93
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 847 | 14.951
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 848 | 14.972
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 849 | 14.993
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 850 | 15.014
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 851 | 15.035
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 852 | 15.056
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 853 | 15.077
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 854 | 15.098
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 855 | 15.119
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 856 | 15.14
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 857 | 15.162
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 858 | 15.183
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 859 | 15.204
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 860 | 15.225
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 861 | 15.246
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 862 | 15.267
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 863 | 15.288
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 864 | 15.309
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 865 | 15.33
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 866 | 15.351
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 867 | 15.372
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 868 | 15.393
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 869 | 15.414
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 870 | 15.435
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 871 | 15.456
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 872 | 15.477
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 873 | 15.498
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 874 | 15.52
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 875 | 15.541
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 876 | 15.562
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 877 | 15.583
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 878 | 15.604
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 879 | 15.625
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 880 | 15.646
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 881 | 15.667
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 882 | 15.688
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 883 | 15.709
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 884 | 15.731
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 885 | 15.752
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 886 | 15.773
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 887 | 15.794
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 888 | 15.815
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 889 | 15.836
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 890 | 15.857
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 891 | 15.878
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 892 | 15.9
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 893 | 15.921
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 894 | 15.942
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 895 | 15.963
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 896 | 15.984
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 897 | 16.005
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 898 | 16.027
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 899 | 16.048
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 900 | 16.069
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 901 | 16.09
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 902 | 16.111
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 903 | 16.132
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 904 | 16.154
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 905 | 16.175
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 906 | 16.196
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 907 | 16.217
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 908 | 16.238
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 909 | 16.26
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 910 | 16.281
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 911 | 16.302
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 912 | 16.323
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 913 | 16.344
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 914 | 16.366
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 915 | 16.387
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 916 | 16.408
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 917 | 16.429
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 918 | 16.45
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 919 | 16.472
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 920 | 16.493
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 921 | 16.514
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 922 | 16.535
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 923 | 16.557
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 924 | 16.578
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 925 | 16.599
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 926 | 16.62
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 927 | 16.642
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 928 | 16.663
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 929 | 16.684
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 930 | 16.705
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 931 | 16.727
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 932 | 16.748
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 933 | 16.769
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 934 | 16.79
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 935 | 16.812
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 936 | 16.833
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 937 | 16.854
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 938 | 16.875
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 939 | 16.897
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 940 | 16.918
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 941 | 16.939
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 942 | 16.961
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 943 | 16.982
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 944 | 17.003
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 945 | 17.025
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 946 | 17.046
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 947 | 17.067
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 948 | 17.088
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 949 | 17.11
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 950 | 17.131
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 951 | 17.152
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 952 | 17.174
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 953 | 17.195
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 954 | 17.216
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 955 | 17.238
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 956 | 17.259
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 957 | 17.28
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 958 | 17.302
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 959 | 17.323
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 960 | 17.344
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 961 | 17.366
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 962 | 17.387
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 963 | 17.408
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 964 | 17.43
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 965 | 17.451
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 966 | 17.472
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 967 | 17.494
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 968 | 17.515
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 969 | 17.536
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 970 | 17.558
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 971 | 17.579
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 972 | 17.601
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 973 | 17.622
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 974 | 17.643
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 975 | 17.665
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 976 | 17.686
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 977 | 17.707
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 978 | 17.729
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 979 | 17.75
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 980 | 17.772
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 981 | 17.793
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 982 | 17.814
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 983 | 17.836
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 984 | 17.857
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 985 | 17.879
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 986 | 17.9
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 987 | 17.921
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 988 | 17.943
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 989 | 17.964
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 990 | 17.986
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 991 | 18.007
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 992 | 18.028
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 993 | 18.05
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 994 | 18.071
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 995 | 18.093
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 996 | 18.114
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 997 | 18.136
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 998 | 18.157
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 999 | 18.178
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1000 | 18.2
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1001 | 18.221
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1002 | 18.243
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1003 | 18.264
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1004 | 18.286
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1005 | 18.307
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1006 | 18.328
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1007 | 18.35
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1008 | 18.371
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1009 | 18.393
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1010 | 18.414
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1011 | 18.436
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1012 | 18.457
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1013 | 18.479
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1014 | 18.5
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1015 | 18.522
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1016 | 18.543
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1017 | 18.565
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1018 | 18.586
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1019 | 18.608
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1020 | 18.629
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1021 | 18.65
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1022 | 18.672
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1023 | 18.693
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1024 | 18.715
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1025 | 18.736
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1026 | 18.758
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1027 | 18.779
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1028 | 18.801
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1029 | 18.822
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1030 | 18.844
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1031 | 18.865
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1032 | 18.887
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1033 | 18.908
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1034 | 18.93
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1035 | 18.951
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1036 | 18.973
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1037 | 18.994
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1038 | 19.016
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1039 | 19.037
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1040 | 19.059
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1041 | 19.081
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1042 | 19.102
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1043 | 19.124
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1044 | 19.145
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1045 | 19.167
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1046 | 19.188
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1047 | 19.21
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1048 | 19.231
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1049 | 19.253
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1050 | 19.274
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1051 | 19.296
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1052 | 19.317
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1053 | 19.339
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1054 | 19.36
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1055 | 19.382
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1056 | 19.404
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1057 | 19.425
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1058 | 19.447
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1059 | 19.468
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1060 | 19.49
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1061 | 19.511
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1062 | 19.533
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1063 | 19.554
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1064 | 19.576
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1065 | 19.598
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1066 | 19.619
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1067 | 19.641
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1068 | 19.662
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1069 | 19.684
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1070 | 19.705
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1071 | 19.727
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1072 | 19.749
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1073 | 19.77
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1074 | 19.792
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1075 | 19.813
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1076 | 19.835
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1077 | 19.857
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1078 | 19.878
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1079 | 19.9
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1080 | 19.921
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1081 | 19.943
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1082 | 19.964
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1083 | 19.986
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1084 | 20.008
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1085 | 20.029
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1086 | 20.051
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1087 | 20.072
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1088 | 20.094
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1089 | 20.116
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1090 | 20.137
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1091 | 20.159
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1092 | 20.181
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1093 | 20.202
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1094 | 20.224
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1095 | 20.245
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1096 | 20.267
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1097 | 20.289
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1098 | 20.31
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1099 | 20.332
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1100 | 20.353
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1101 | 20.375
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1102 | 20.397
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1103 | 20.418
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1104 | 20.44
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1105 | 20.462
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1106 | 20.483
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1107 | 20.505
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1108 | 20.527
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1109 | 20.548
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1110 | 20.57
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1111 | 20.591
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1112 | 20.613
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1113 | 20.635
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1114 | 20.656
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1115 | 20.678
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1116 | 20.7
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1117 | 20.721
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1118 | 20.743
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1119 | 20.765
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1120 | 20.786
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1121 | 20.808
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1122 | 20.83
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1123 | 20.851
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1124 | 20.873
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1125 | 20.895
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1126 | 20.916
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1127 | 20.938
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1128 | 20.96
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1129 | 20.981
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1130 | 21.003
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1131 | 21.025
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1132 | 21.046
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1133 | 21.068
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1134 | 21.09
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1135 | 21.111
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1136 | 21.133
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1137 | 21.155
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1138 | 21.176
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1139 | 21.198
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1140 | 21.22
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1141 | 21.241
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1142 | 21.263
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1143 | 21.285
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1144 | 21.306
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1145 | 21.328
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1146 | 21.35
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1147 | 21.371
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1148 | 21.393
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1149 | 21.415
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1150 | 21.437
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1151 | 21.458
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1152 | 21.48
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1153 | 21.502
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1154 | 21.523
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1155 | 21.545
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1156 | 21.567
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1157 | 21.588
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1158 | 21.61
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1159 | 21.632
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1160 | 21.654
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1161 | 21.675
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1162 | 21.697
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1163 | 21.719
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1164 | 21.74
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1165 | 21.762
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1166 | 21.784
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1167 | 21.806
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1168 | 21.827
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1169 | 21.849
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1170 | 21.871
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1171 | 21.892
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1172 | 21.914
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1173 | 21.936
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1174 | 21.958
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1175 | 21.979
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1176 | 22.001
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1177 | 22.023
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1178 | 22.044
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1179 | 22.066
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1180 | 22.088
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1181 | 22.11
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1182 | 22.131
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1183 | 22.153
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1184 | 22.175
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1185 | 22.197
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1186 | 22.218
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1187 | 22.24
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1188 | 22.262
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1189 | 22.284
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1190 | 22.305
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1191 | 22.327
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1192 | 22.349
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1193 | 22.37
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1194 | 22.392
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1195 | 22.414
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1196 | 22.436
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1197 | 22.457
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1198 | 22.479
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1199 | 22.501
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1200 | 22.523
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1201 | 22.544
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1202 | 22.566
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1203 | 22.588
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1204 | 22.61
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1205 | 22.631
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1206 | 22.653
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1207 | 22.675
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1208 | 22.697
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1209 | 22.718
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1210 | 22.74
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1211 | 22.762
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1212 | 22.784
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1213 | 22.805
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1214 | 22.827
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1215 | 22.849
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1216 | 22.871
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1217 | 22.893
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1218 | 22.914
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1219 | 22.936
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1220 | 22.958
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1221 | 22.98
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1222 | 23.001
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1223 | 23.023
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1224 | 23.045
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1225 | 23.067
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1226 | 23.088
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1227 | 23.11
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1228 | 23.132
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1229 | 23.154
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1230 | 23.176
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1231 | 23.197
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1232 | 23.219
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1233 | 23.241
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1234 | 23.263
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1235 | 23.284
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1236 | 23.306
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1237 | 23.328
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1238 | 23.35
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1239 | 23.372
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1240 | 23.393
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1241 | 23.415
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1242 | 23.437
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1243 | 23.459
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1244 | 23.48
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1245 | 23.502
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1246 | 23.524
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1247 | 23.546
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1248 | 23.568
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1249 | 23.589
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1250 | 23.611
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1251 | 23.633
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1252 | 23.655
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1253 | 23.676
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1254 | 23.698
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1255 | 23.72
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1256 | 23.742
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1257 | 23.764
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1258 | 23.785
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1259 | 23.807
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1260 | 23.829
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1261 | 23.851
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1262 | 23.873
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1263 | 23.894
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1264 | 23.916
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1265 | 23.938
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1266 | 23.96
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1267 | 23.982
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1268 | 24.003
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1269 | 24.025
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1270 | 24.047
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1271 | 24.069
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1272 | 24.091
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1273 | 24.112
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1274 | 24.134
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1275 | 24.156
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1276 | 24.178
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1277 | 24.2
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1278 | 24.221
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1279 | 24.243
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1280 | 24.265
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1281 | 24.287
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1282 | 24.309
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1283 | 24.33
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1284 | 24.352
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1285 | 24.374
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1286 | 24.396
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1287 | 24.418
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1288 | 24.439
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1289 | 24.461
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1290 | 24.483
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1291 | 24.505
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1292 | 24.527
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1293 | 24.548
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1294 | 24.57
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1295 | 24.592
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1296 | 24.614
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1297 | 24.636
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1298 | 24.658
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1299 | 24.679
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1300 | 24.701
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1301 | 24.723
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1302 | 24.745
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1303 | 24.767
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1304 | 24.788
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1305 | 24.81
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1306 | 24.832
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1307 | 24.854
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1308 | 24.876
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1309 | 24.897
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1310 | 24.919
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1311 | 24.941
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1312 | 24.963
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1313 | 24.985
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1314 | 25.007
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1315 | 25.028
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1316 | 25.05
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1317 | 25.072
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1318 | 25.094
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1319 | 25.116
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1320 | 25.137
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1321 | 25.159
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1322 | 25.181
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1323 | 25.203
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1324 | 25.225
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1325 | 25.247
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1326 | 25.268
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1327 | 25.29
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1328 | 25.312
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1329 | 25.334
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1330 | 25.356
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1331 | 25.377
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1332 | 25.399
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1333 | 25.421
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1334 | 25.443
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1335 | 25.465
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1336 | 25.487
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1337 | 25.508
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1338 | 25.53
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1339 | 25.552
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1340 | 25.574
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1341 | 25.596
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1342 | 25.618
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1343 | 25.639
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1344 | 25.661
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1345 | 25.683
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1346 | 25.705
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1347 | 25.727
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1348 | 25.748
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1349 | 25.77
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1350 | 25.792
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1351 | 25.814
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1352 | 25.836
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1353 | 25.858
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1354 | 25.879
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1355 | 25.901
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1356 | 25.923
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1357 | 25.945
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1358 | 25.967
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1359 | 25.989
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1360 | 26.01
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1361 | 26.032
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1362 | 26.054
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1363 | 26.076
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1364 | 26.098
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1365 | 26.119
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1366 | 26.141
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1367 | 26.163
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1368 | 26.185
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1369 | 26.207
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1370 | 26.229
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1371 | 26.25
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1372 | 26.272
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1373 | 26.294
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1374 | 26.316
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1375 | 26.338
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1376 | 26.36
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1377 | 26.381
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1378 | 26.403
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1379 | 26.425
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1380 | 26.447
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1381 | 26.469
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1382 | 26.491
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1383 | 26.512
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1384 | 26.534
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1385 | 26.556
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1386 | 26.578
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1387 | 26.6
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1388 | 26.622
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1389 | 26.643
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1390 | 26.665
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1391 | 26.687
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1392 | 26.709
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1393 | 26.731
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1394 | 26.752
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1395 | 26.774
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1396 | 26.796
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1397 | 26.818
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1398 | 26.84
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1399 | 26.862
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1400 | 26.883
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1401 | 26.905
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1402 | 26.927
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1403 | 26.949
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1404 | 26.971
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1405 | 26.993
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1406 | 27.014
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1407 | 27.036
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1408 | 27.058
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1409 | 27.08
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1410 | 27.102
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1411 | 27.124
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1412 | 27.145
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1413 | 27.167
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1414 | 27.189
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1415 | 27.211
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1416 | 27.233
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1417 | 27.254
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1418 | 27.276
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1419 | 27.298
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1420 | 27.32
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1421 | 27.342
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1422 | 27.364
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1423 | 27.385
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1424 | 27.407
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1425 | 27.429
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1426 | 27.451
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1427 | 27.473
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1428 | 27.495
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1429 | 27.516
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1430 | 27.538
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1431 | 27.56
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1432 | 27.582
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1433 | 27.604
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1434 | 27.625
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1435 | 27.647
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1436 | 27.669
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1437 | 27.691
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1438 | 27.713
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1439 | 27.735
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1440 | 27.756
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1441 | 27.778
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1442 | 27.8
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1443 | 27.822
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1444 | 27.844
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1445 | 27.866
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1446 | 27.887
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1447 | 27.909
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1448 | 27.931
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1449 | 27.953
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1450 | 27.975
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1451 | 27.996
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1452 | 28.018
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1453 | 28.04
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1454 | 28.062
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1455 | 28.084
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1456 | 28.105
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1457 | 28.127
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1458 | 28.149
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1459 | 28.171
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1460 | 28.193
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1461 | 28.215
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1462 | 28.236
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1463 | 28.258
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1464 | 28.28
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1465 | 28.302
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1466 | 28.324
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1467 | 28.345
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1468 | 28.367
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1469 | 28.389
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1470 | 28.411
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1471 | 28.433
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1472 | 28.455
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1473 | 28.476
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1474 | 28.498
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1475 | 28.52
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1476 | 28.542
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1477 | 28.564
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1478 | 28.585
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1479 | 28.607
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1480 | 28.629
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1481 | 28.651
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1482 | 28.673
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1483 | 28.694
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1484 | 28.716
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1485 | 28.738
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1486 | 28.76
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1487 | 28.782
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1488 | 28.803
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1489 | 28.825
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1490 | 28.847
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1491 | 28.869
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1492 | 28.891
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1493 | 28.912
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1494 | 28.934
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1495 | 28.956
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1496 | 28.978
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1497 | 29
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1498 | 29.021
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1499 | 29.043
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1500 | 29.065
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1501 | 29.087
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1502 | 29.109
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1503 | 29.13
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1504 | 29.152
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1505 | 29.174
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1506 | 29.196
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1507 | 29.218
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1508 | 29.239
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1509 | 29.261
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1510 | 29.283
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1511 | 29.305
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1512 | 29.327
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1513 | 29.348
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1514 | 29.37
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1515 | 29.392
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1516 | 29.414
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1517 | 29.436
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1518 | 29.457
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1519 | 29.479
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1520 | 29.501
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1521 | 29.523
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1522 | 29.545
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1523 | 29.566
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1524 | 29.588
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1525 | 29.61
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1526 | 29.632
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1527 | 29.653
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1528 | 29.675
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1529 | 29.697
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1530 | 29.719
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1531 | 29.741
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1532 | 29.762
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1533 | 29.784
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1534 | 29.806
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1535 | 29.828
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1536 | 29.85
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1537 | 29.871
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1538 | 29.893
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1539 | 29.915
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1540 | 29.937
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1541 | 29.958
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1542 | 29.98
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1543 | 30.002
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1544 | 30.024
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1545 | 30.046
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1546 | 30.067
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1547 | 30.089
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1548 | 30.111
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1549 | 30.133
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1550 | 30.154
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1551 | 30.176
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1552 | 30.198
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1553 | 30.22
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1554 | 30.242
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1555 | 30.263
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1556 | 30.285
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1557 | 30.307
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1558 | 30.329
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1559 | 30.35
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1560 | 30.372
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1561 | 30.394
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1562 | 30.416
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1563 | 30.437
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1564 | 30.459
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1565 | 30.481
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1566 | 30.503
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1567 | 30.524
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1568 | 30.546
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1569 | 30.568
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1570 | 30.59
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1571 | 30.611
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1572 | 30.633
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1573 | 30.655
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1574 | 30.677
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1575 | 30.699
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1576 | 30.72
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1577 | 30.742
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1578 | 30.764
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1579 | 30.786
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1580 | 30.807
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1581 | 30.829
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1582 | 30.851
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1583 | 30.873
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1584 | 30.894
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1585 | 30.916
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1586 | 30.938
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1587 | 30.96
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1588 | 30.981
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1589 | 31.003
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1590 | 31.025
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1591 | 31.047
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1592 | 31.068
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1593 | 31.09
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1594 | 31.112
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1595 | 31.133
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1596 | 31.155
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1597 | 31.177
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1598 | 31.199
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1599 | 31.22
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1600 | 31.242
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1601 | 31.264
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1602 | 31.286
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1603 | 31.307
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1604 | 31.329
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1605 | 31.351
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1606 | 31.373
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1607 | 31.394
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1608 | 31.416
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1609 | 31.438
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1610 | 31.459
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1611 | 31.481
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1612 | 31.503
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1613 | 31.525
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1614 | 31.546
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1615 | 31.568
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1616 | 31.59
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1617 | 31.612
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1618 | 31.633
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1619 | 31.655
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1620 | 31.677
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1621 | 31.698
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1622 | 31.72
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1623 | 31.742
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1624 | 31.764
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1625 | 31.785
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1626 | 31.807
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1627 | 31.829
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1628 | 31.85
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1629 | 31.872
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1630 | 31.894
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1631 | 31.916
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1632 | 31.937
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1633 | 31.959
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1634 | 31.981
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1635 | 32.002
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1636 | 32.024
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1637 | 32.046
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1638 | 32.068
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1639 | 32.089
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1640 | 32.111
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1641 | 32.133
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1642 | 32.154
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1643 | 32.176
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1644 | 32.198
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1645 | 32.219
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1646 | 32.241
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1647 | 32.263
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1648 | 32.284
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1649 | 32.306
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1650 | 32.328
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1651 | 32.35
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1652 | 32.371
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1653 | 32.393
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1654 | 32.415
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1655 | 32.436
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1656 | 32.458
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1657 | 32.48
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1658 | 32.501
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1659 | 32.523
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1660 | 32.545
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1661 | 32.566
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1662 | 32.588
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1663 | 32.61
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1664 | 32.631
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1665 | 32.653
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1666 | 32.675
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1667 | 32.696
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1668 | 32.718
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1669 | 32.74
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1670 | 32.761
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1671 | 32.783
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1672 | 32.805
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1673 | 32.826
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1674 | 32.848
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1675 | 32.87
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1676 | 32.891
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1677 | 32.913
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1678 | 32.935
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1679 | 32.956
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1680 | 32.978
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1681 | 33
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1682 | 33.021
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1683 | 33.043
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1684 | 33.065
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1685 | 33.086
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1686 | 33.108
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1687 | 33.13
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1688 | 33.151
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1689 | 33.173
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1690 | 33.195
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1691 | 33.216
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1692 | 33.238
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1693 | 33.26
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1694 | 33.281
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1695 | 33.303
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1696 | 33.325
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1697 | 33.346
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1698 | 33.368
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1699 | 33.389
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1700 | 33.411
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1701 | 33.433
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1702 | 33.454
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1703 | 33.476
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1704 | 33.498
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1705 | 33.519
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1706 | 33.541
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1707 | 33.563
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1708 | 33.584
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1709 | 33.606
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1710 | 33.627
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1711 | 33.649
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1712 | 33.671
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1713 | 33.692
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1714 | 33.714
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1715 | 33.736
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1716 | 33.757
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1717 | 33.779
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1718 | 33.8
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1719 | 33.822
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1720 | 33.844
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1721 | 33.865
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1722 | 33.887
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1723 | 33.908
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1724 | 33.93
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1725 | 33.952
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1726 | 33.973
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1727 | 33.995
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1728 | 34.016
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1729 | 34.038
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1730 | 34.06
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1731 | 34.081
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1732 | 34.103
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1733 | 34.124
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1734 | 34.146
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1735 | 34.168
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1736 | 34.189
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1737 | 34.211
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1738 | 34.232
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1739 | 34.254
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1740 | 34.276
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1741 | 34.297
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1742 | 34.319
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1743 | 34.34
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1744 | 34.362
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1745 | 34.384
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1746 | 34.405
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1747 | 34.427
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1748 | 34.448
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1749 | 34.47
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1750 | 34.491
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1751 | 34.513
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1752 | 34.535
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1753 | 34.556
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1754 | 34.578
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1755 | 34.599
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1756 | 34.621
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1757 | 34.642
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1758 | 34.664
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1759 | 34.686
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1760 | 34.707
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1761 | 34.729
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1762 | 34.75
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1763 | 34.772
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1764 | 34.793
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1765 | 34.815
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1766 | 34.836
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1767 | 34.858
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1768 | 34.879
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1769 | 34.901
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1770 | 34.923
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1771 | 34.944
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1772 | 34.966
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1773 | 34.987
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1774 | 35.009
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1775 | 35.03
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1776 | 35.052
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1777 | 35.073
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1778 | 35.095
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1779 | 35.116
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1780 | 35.138
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1781 | 35.16
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1782 | 35.181
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1783 | 35.203
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1784 | 35.224
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1785 | 35.246
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1786 | 35.267
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1787 | 35.289
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1788 | 35.31
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1789 | 35.332
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1790 | 35.353
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1791 | 35.375
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1792 | 35.396
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1793 | 35.418
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1794 | 35.439
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1795 | 35.461
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1796 | 35.482
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1797 | 35.504
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1798 | 35.525
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1799 | 35.547
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1800 | 35.568
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1801 | 35.59
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1802 | 35.611
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1803 | 35.633
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1804 | 35.654
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1805 | 35.676
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1806 | 35.697
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1807 | 35.719
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1808 | 35.74
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1809 | 35.762
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1810 | 35.783
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1811 | 35.805
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1812 | 35.826
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1813 | 35.848
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1814 | 35.869
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1815 | 35.891
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1816 | 35.912
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1817 | 35.934
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1818 | 35.955
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1819 | 35.977
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1820 | 35.998
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1821 | 36.019
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1822 | 36.041
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1823 | 36.062
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1824 | 36.084
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1825 | 36.105
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1826 | 36.127
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1827 | 36.148
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1828 | 36.17
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1829 | 36.191
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1830 | 36.213
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1831 | 36.234
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1832 | 36.256
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1833 | 36.277
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1834 | 36.298
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1835 | 36.32
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1836 | 36.341
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1837 | 36.363
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1838 | 36.384
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1839 | 36.406
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1840 | 36.427
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1841 | 36.449
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1842 | 36.47
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1843 | 36.491
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1844 | 36.513
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1845 | 36.534
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1846 | 36.556
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1847 | 36.577
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1848 | 36.599
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1849 | 36.62
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1850 | 36.641
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1851 | 36.663
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1852 | 36.684
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1853 | 36.706
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1854 | 36.727
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1855 | 36.748
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1856 | 36.77
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1857 | 36.791
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1858 | 36.813
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1859 | 36.834
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1860 | 36.855
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1861 | 36.877
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1862 | 36.898
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1863 | 36.92
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1864 | 36.941
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1865 | 36.962
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1866 | 36.984
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1867 | 37.005
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1868 | 37.027
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1869 | 37.048
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1870 | 37.069
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1871 | 37.091
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1872 | 37.112
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1873 | 37.134
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1874 | 37.155
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1875 | 37.176
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1876 | 37.198
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1877 | 37.219
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1878 | 37.24
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1879 | 37.262
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1880 | 37.283
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1881 | 37.305
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1882 | 37.326
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1883 | 37.347
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1884 | 37.369
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1885 | 37.39
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1886 | 37.411
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1887 | 37.433
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1888 | 37.454
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1889 | 37.475
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1890 | 37.497
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1891 | 37.518
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1892 | 37.539
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1893 | 37.561
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1894 | 37.582
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1895 | 37.603
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1896 | 37.625
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1897 | 37.646
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1898 | 37.668
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1899 | 37.689
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1900 | 37.71
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1901 | 37.731
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1902 | 37.753
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1903 | 37.774
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1904 | 37.795
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1905 | 37.817
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1906 | 37.838
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1907 | 37.859
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1908 | 37.881
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1909 | 37.902
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1910 | 37.923
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1911 | 37.945
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1912 | 37.966
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1913 | 37.987
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1914 | 38.009
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1915 | 38.03
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1916 | 38.051
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1917 | 38.073
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1918 | 38.094
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1919 | 38.115
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1920 | 38.136
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1921 | 38.158
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1922 | 38.179
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1923 | 38.2
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1924 | 38.222
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1925 | 38.243
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1926 | 38.264
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1927 | 38.285
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1928 | 38.307
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1929 | 38.328
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1930 | 38.349
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1931 | 38.37
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1932 | 38.392
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1933 | 38.413
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1934 | 38.434
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1935 | 38.456
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1936 | 38.477
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1937 | 38.498
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1938 | 38.519
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1939 | 38.541
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1940 | 38.562
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1941 | 38.583
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1942 | 38.604
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1943 | 38.626
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1944 | 38.647
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1945 | 38.668
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1946 | 38.689
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1947 | 38.711
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1948 | 38.732
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1949 | 38.753
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1950 | 38.774
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1951 | 38.795
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1952 | 38.817
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1953 | 38.838
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1954 | 38.859
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1955 | 38.88
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1956 | 38.902
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1957 | 38.923
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1958 | 38.944
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1959 | 38.965
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1960 | 38.986
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1961 | 39.008
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1962 | 39.029
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1963 | 39.05
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1964 | 39.071
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1965 | 39.093
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1966 | 39.114
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1967 | 39.135
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1968 | 39.156
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1969 | 39.177
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1970 | 39.198
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1971 | 39.22
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1972 | 39.241
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1973 | 39.262
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1974 | 39.283
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1975 | 39.304
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1976 | 39.326
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1977 | 39.347
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1978 | 39.368
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1979 | 39.389
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1980 | 39.41
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1981 | 39.431
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1982 | 39.453
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1983 | 39.474
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1984 | 39.495
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1985 | 39.516
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1986 | 39.537
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1987 | 39.558
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1988 | 39.58
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1989 | 39.601
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1990 | 39.622
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1991 | 39.643
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1992 | 39.664
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1993 | 39.685
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1994 | 39.706
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1995 | 39.728
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1996 | 39.749
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1997 | 39.77
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1998 | 39.791
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 1999 | 39.812
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2000 | 39.833
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2001 | 39.854
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2002 | 39.875
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2003 | 39.897
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2004 | 39.918
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2005 | 39.939
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2006 | 39.96
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2007 | 39.981
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2008 | 40.002
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2009 | 40.023
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2010 | 40.044
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2011 | 40.066
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2012 | 40.087
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2013 | 40.108
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2014 | 40.129
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2015 | 40.15
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2016 | 40.171
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2017 | 40.192
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2018 | 40.213
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2019 | 40.234
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2020 | 40.255
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2021 | 40.276
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2022 | 40.297
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2023 | 40.319
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2024 | 40.34
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2025 | 40.361
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2026 | 40.382
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2027 | 40.403
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2028 | 40.424
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2029 | 40.445
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2030 | 40.466
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2031 | 40.487
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2032 | 40.508
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2033 | 40.529
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2034 | 40.55
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2035 | 40.571
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2036 | 40.592
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2037 | 40.613
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2038 | 40.634
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2039 | 40.655
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2040 | 40.677
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2041 | 40.698
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2042 | 40.719
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2043 | 40.74
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2044 | 40.761
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2045 | 40.782
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2046 | 40.803
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2047 | 40.824
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2048 | 40.845
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2049 | 40.866
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2050 | 40.887
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2051 | 40.908
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2052 | 40.929
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2053 | 40.95
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2054 | 40.971
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2055 | 40.992
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2056 | 41.013
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2057 | 41.034
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2058 | 41.055
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2059 | 41.076
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2060 | 41.097
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2061 | 41.118
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2062 | 41.139
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2063 | 41.16
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2064 | 41.181
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2065 | 41.202
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2066 | 41.223
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2067 | 41.244
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2068 | 41.265
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2069 | 41.286
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2070 | 41.307
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2071 | 41.328
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2072 | 41.349
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2073 | 41.37
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2074 | 41.39
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2075 | 41.411
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2076 | 41.432
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2077 | 41.453
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2078 | 41.474
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2079 | 41.495
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2080 | 41.516
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2081 | 41.537
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2082 | 41.558
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2083 | 41.579
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2084 | 41.6
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2085 | 41.621
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2086 | 41.642
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2087 | 41.663
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2088 | 41.684
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2089 | 41.705
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2090 | 41.725
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2091 | 41.746
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2092 | 41.767
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2093 | 41.788
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2094 | 41.809
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2095 | 41.83
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2096 | 41.851
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2097 | 41.872
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2098 | 41.893
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2099 | 41.914
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2100 | 41.935
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2101 | 41.955
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2102 | 41.976
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2103 | 41.997
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2104 | 42.018
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2105 | 42.039
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2106 | 42.06
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2107 | 42.081
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2108 | 42.102
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2109 | 42.123
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2110 | 42.143
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2111 | 42.164
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2112 | 42.185
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2113 | 42.206
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2114 | 42.227
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2115 | 42.248
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2116 | 42.269
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2117 | 42.289
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2118 | 42.31
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2119 | 42.331
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2120 | 42.352
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2121 | 42.373
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2122 | 42.394
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2123 | 42.415
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2124 | 42.435
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2125 | 42.456
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2126 | 42.477
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2127 | 42.498
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2128 | 42.519
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2129 | 42.54
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2130 | 42.56
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2131 | 42.581
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2132 | 42.602
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2133 | 42.623
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2134 | 42.644
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2135 | 42.664
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2136 | 42.685
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2137 | 42.706
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2138 | 42.727
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2139 | 42.748
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2140 | 42.768
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2141 | 42.789
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2142 | 42.81
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2143 | 42.831
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2144 | 42.852
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2145 | 42.872
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2146 | 42.893
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2147 | 42.914
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2148 | 42.935
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2149 | 42.956
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2150 | 42.976
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2151 | 42.997
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2152 | 43.018
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2153 | 43.039
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2154 | 43.059
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2155 | 43.08
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2156 | 43.101
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2157 | 43.122
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2158 | 43.142
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2159 | 43.163
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2160 | 43.184
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2161 | 43.205
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2162 | 43.225
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2163 | 43.246
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2164 | 43.267
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2165 | 43.288
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2166 | 43.308
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2167 | 43.329
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2168 | 43.35
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2169 | 43.37
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2170 | 43.391
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2171 | 43.412
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2172 | 43.433
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2173 | 43.453
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2174 | 43.474
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2175 | 43.495
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2176 | 43.515
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2177 | 43.536
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2178 | 43.557
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2179 | 43.578
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2180 | 43.598
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2181 | 43.619
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2182 | 43.64
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2183 | 43.66
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2184 | 43.681
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2185 | 43.702
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2186 | 43.722
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2187 | 43.743
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2188 | 43.764
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2189 | 43.784
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2190 | 43.805
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2191 | 43.826
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2192 | 43.846
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2193 | 43.867
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2194 | 43.888
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2195 | 43.908
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2196 | 43.929
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2197 | 43.95
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2198 | 43.97
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2199 | 43.991
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2200 | 44.012
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2201 | 44.032
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2202 | 44.053
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2203 | 44.073
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2204 | 44.094
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2205 | 44.115
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2206 | 44.135
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2207 | 44.156
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2208 | 44.177
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2209 | 44.197
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2210 | 44.218
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2211 | 44.238
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2212 | 44.259
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2213 | 44.28
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2214 | 44.3
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2215 | 44.321
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2216 | 44.341
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2217 | 44.362
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2218 | 44.383
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2219 | 44.403
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2220 | 44.424
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2221 | 44.444
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2222 | 44.465
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2223 | 44.485
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2224 | 44.506
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2225 | 44.527
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2226 | 44.547
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2227 | 44.568
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2228 | 44.588
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2229 | 44.609
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2230 | 44.629
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2231 | 44.65
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2232 | 44.671
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2233 | 44.691
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2234 | 44.712
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2235 | 44.732
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2236 | 44.753
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2237 | 44.773
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2238 | 44.794
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2239 | 44.814
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2240 | 44.835
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2241 | 44.855
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2242 | 44.876
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2243 | 44.896
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2244 | 44.917
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2245 | 44.937
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2246 | 44.958
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2247 | 44.978
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2248 | 44.999
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2249 | 45.019
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2250 | 45.04
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2251 | 45.06
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2252 | 45.081
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2253 | 45.101
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2254 | 45.122
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2255 | 45.142
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2256 | 45.163
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2257 | 45.183
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2258 | 45.204
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2259 | 45.224
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2260 | 45.245
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2261 | 45.265
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2262 | 45.286
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2263 | 45.306
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2264 | 45.326
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2265 | 45.347
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2266 | 45.367
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2267 | 45.388
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2268 | 45.408
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2269 | 45.429
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2270 | 45.449
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2271 | 45.469
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2272 | 45.49
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2273 | 45.51
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2274 | 45.531
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2275 | 45.551
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2276 | 45.572
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2277 | 45.592
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2278 | 45.612
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2279 | 45.633
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2280 | 45.653
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2281 | 45.674
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2282 | 45.694
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2283 | 45.714
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2284 | 45.735
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2285 | 45.755
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2286 | 45.775
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2287 | 45.796
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2288 | 45.816
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2289 | 45.837
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2290 | 45.857
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2291 | 45.877
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2292 | 45.898
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2293 | 45.918
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2294 | 45.938
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2295 | 45.959
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2296 | 45.979
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2297 | 45.999
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2298 | 46.02
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2299 | 46.04
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2300 | 46.06
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2301 | 46.081
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2302 | 46.101
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2303 | 46.121
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2304 | 46.142
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2305 | 46.162
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2306 | 46.182
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2307 | 46.202
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2308 | 46.223
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2309 | 46.243
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2310 | 46.263
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2311 | 46.284
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2312 | 46.304
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2313 | 46.324
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2314 | 46.344
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2315 | 46.365
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2316 | 46.385
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2317 | 46.405
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2318 | 46.425
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2319 | 46.446
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2320 | 46.466
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2321 | 46.486
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2322 | 46.506
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2323 | 46.527
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2324 | 46.547
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2325 | 46.567
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2326 | 46.587
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2327 | 46.608
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2328 | 46.628
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2329 | 46.648
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2330 | 46.668
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2331 | 46.688
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2332 | 46.709
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2333 | 46.729
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2334 | 46.749
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2335 | 46.769
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2336 | 46.789
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2337 | 46.81
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2338 | 46.83
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2339 | 46.85
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2340 | 46.87
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2341 | 46.89
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2342 | 46.91
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2343 | 46.931
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2344 | 46.951
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2345 | 46.971
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2346 | 46.991
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2347 | 47.011
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2348 | 47.031
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2349 | 47.051
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2350 | 47.071
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2351 | 47.092
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2352 | 47.112
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2353 | 47.132
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2354 | 47.152
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2355 | 47.172
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2356 | 47.192
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2357 | 47.212
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2358 | 47.232
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2359 | 47.252
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2360 | 47.272
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2361 | 47.292
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2362 | 47.312
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2363 | 47.333
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2364 | 47.353
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2365 | 47.373
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2366 | 47.393
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2367 | 47.413
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2368 | 47.433
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2369 | 47.453
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2370 | 47.473
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2371 | 47.493
0 | 32 | 486 | -0.015 | 0 | 31 | 32 | 2372 | 47.513
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32
0 | 32 | 486 | -0.015 | 0 | 31 | 32

## Type T
[7 empty] | Temp(F) | mV
mV | Temp(F) | [5 empty] | -454 | -6.258
-50 | #VALUE! | 0 | mV | -6.258 | Temp(F) | -454 | -453 | -6.257
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -452 | -6.256
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -451 | -6.255
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -450 | -6.254
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -449 | -6.253
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -448 | -6.252
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -447 | -6.251
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -446 | -6.25
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -445 | -6.248
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -444 | -6.247
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -443 | -6.245
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -442 | -6.243
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -441 | -6.242
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -440 | -6.24
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -439 | -6.238
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -438 | -6.236
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -437 | -6.234
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -436 | -6.232
200 | 7205.826 | 1207 | 20.872 | 0 | 752 | 0 | -435 | -6.23
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -434 | -6.227
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -433 | -6.225
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -432 | -6.222
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -431 | -6.22
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -430 | -6.217
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -429 | -6.215
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -428 | -6.212
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -427 | -6.209
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -426 | -6.206
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -425 | -6.203
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -424 | -6.2
450 | 16213.1085 | 1207 | 20.872 | 0 | 752 | 0 | -423 | -6.197
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -422 | -6.194
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -421 | -6.191
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -420 | -6.187
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -419 | -6.184
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -418 | -6.18
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -417 | -6.177
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -416 | -6.173
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -415 | -6.17
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -414 | -6.166
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -413 | -6.162
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -412 | -6.158
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -411 | -6.154
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -410 | -6.15
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -409 | -6.146
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -408 | -6.141
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -407 | -6.137
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -406 | -6.133
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -405 | -6.128
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -404 | -6.124
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -403 | -6.119
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -402 | -6.115
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -401 | -6.11
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -400 | -6.105
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -399 | -6.1
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -398 | -6.095
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -397 | -6.09
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -396 | -6.085
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -395 | -6.08
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -394 | -6.075
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -393 | -6.069
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -392 | -6.064
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -391 | -6.059
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -390 | -6.053
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -389 | -6.047
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -388 | -6.042
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -387 | -6.036
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -386 | -6.03
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -385 | -6.025
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -384 | -6.019
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -383 | -6.013
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -382 | -6.007
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -381 | -6.001
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -380 | -5.994
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -379 | -5.988
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -378 | -5.982
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -377 | -5.976
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -376 | -5.969
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -375 | -5.963
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -374 | -5.956
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -373 | -5.95
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -372 | -5.943
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -371 | -5.937
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -370 | -5.93
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -369 | -5.923
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -368 | -5.916
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -367 | -5.909
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -366 | -5.902
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -365 | -5.896
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -364 | -5.888
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -363 | -5.881
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -362 | -5.874
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -361 | -5.867
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -360 | -5.86
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -359 | -5.853
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -358 | -5.845
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -357 | -5.838
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -356 | -5.83
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -355 | -5.823
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -354 | -5.815
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -353 | -5.808
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -352 | -5.8
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -351 | -5.792
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -350 | -5.785
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -349 | -5.777
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -348 | -5.769
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -347 | -5.761
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -346 | -5.753
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -345 | -5.745
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -344 | -5.737
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -343 | -5.729
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -342 | -5.721
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -341 | -5.713
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -340 | -5.705
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -339 | -5.697
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -338 | -5.688
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -337 | -5.68
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -336 | -5.672
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -335 | -5.663
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -334 | -5.655
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -333 | -5.646
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -332 | -5.638
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -331 | -5.629
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -330 | -5.62
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -329 | -5.612
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -328 | -5.603
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -327 | -5.594
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -326 | -5.585
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -325 | -5.577
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -324 | -5.568
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -323 | -5.559
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -322 | -5.55
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -321 | -5.541
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -320 | -5.532
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -319 | -5.523
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -318 | -5.513
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -317 | -5.504
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -316 | -5.495
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -315 | -5.486
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -314 | -5.476
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -313 | -5.467
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -312 | -5.458
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -311 | -5.448
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -310 | -5.439
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -309 | -5.429
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -308 | -5.42
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -307 | -5.41
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -306 | -5.4
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -305 | -5.391
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -304 | -5.381
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -303 | -5.371
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -302 | -5.361
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -301 | -5.351
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -300 | -5.341
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -299 | -5.332
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -298 | -5.322
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -297 | -5.312
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -296 | -5.301
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -295 | -5.291
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -294 | -5.281
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -293 | -5.271
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -292 | -5.261
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -291 | -5.25
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -290 | -5.24
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -289 | -5.23
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -288 | -5.219
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -287 | -5.209
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -286 | -5.198
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -285 | -5.188
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -284 | -5.177
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -283 | -5.167
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -282 | -5.156
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -281 | -5.145
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -280 | -5.135
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -279 | -5.124
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -278 | -5.113
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -277 | -5.102
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -276 | -5.091
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -275 | -5.081
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -274 | -5.07
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -273 | -5.059
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -272 | -5.048
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -271 | -5.036
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -270 | -5.025
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -269 | -5.014
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -268 | -5.003
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -267 | -4.992
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -266 | -4.98
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -265 | -4.969
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -264 | -4.958
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -263 | -4.946
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -262 | -4.935
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -261 | -4.923
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -260 | -4.912
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -259 | -4.9
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -258 | -4.889
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -257 | -4.877
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -256 | -4.865
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -255 | -4.854
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -254 | -4.842
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -253 | -4.83
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -252 | -4.818
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -251 | -4.806
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -250 | -4.794
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -249 | -4.783
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -248 | -4.771
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -247 | -4.759
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -246 | -4.746
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -245 | -4.734
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -244 | -4.722
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -243 | -4.71
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -242 | -4.698
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -241 | -4.685
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -240 | -4.673
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -239 | -4.661
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -238 | -4.648
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -237 | -4.636
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -236 | -4.624
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -235 | -4.611
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -234 | -4.599
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -233 | -4.586
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -232 | -4.573
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -231 | -4.561
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -230 | -4.548
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -229 | -4.535
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -228 | -4.523
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -227 | -4.51
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -226 | -4.497
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -225 | -4.484
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -224 | -4.471
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -223 | -4.458
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -222 | -4.445
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -221 | -4.432
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -220 | -4.419
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -219 | -4.406
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -218 | -4.393
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -217 | -4.38
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -216 | -4.366
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -215 | -4.353
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -214 | -4.34
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -213 | -4.326
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -212 | -4.313
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -211 | -4.3
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -210 | -4.286
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -209 | -4.273
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -208 | -4.259
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -207 | -4.246
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -206 | -4.232
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -205 | -4.218
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -204 | -4.205
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -203 | -4.191
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -202 | -4.177
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -201 | -4.163
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -200 | -4.149
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -199 | -4.136
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -198 | -4.122
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -197 | -4.108
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -196 | -4.094
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -195 | -4.08
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -194 | -4.066
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -193 | -4.052
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -192 | -4.037
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -191 | -4.023
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -190 | -4.009
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -189 | -3.995
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -188 | -3.98
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -187 | -3.966
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -186 | -3.952
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -185 | -3.937
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -184 | -3.923
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -183 | -3.908
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -182 | -3.894
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -181 | -3.879
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -180 | -3.865
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -179 | -3.85
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -178 | -3.836
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -177 | -3.821
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -176 | -3.806
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -175 | -3.791
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -174 | -3.777
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -173 | -3.762
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -172 | -3.747
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -171 | -3.732
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -170 | -3.717
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -169 | -3.702
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -168 | -3.687
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -167 | -3.672
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -166 | -3.657
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -165 | -3.642
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -164 | -3.626
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -163 | -3.611
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -162 | -3.596
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -161 | -3.581
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -160 | -3.565
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -159 | -3.55
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -158 | -3.535
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -157 | -3.519
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -156 | -3.504
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -155 | -3.488
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -154 | -3.473
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -153 | -3.457
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -152 | -3.441
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -151 | -3.426
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -150 | -3.41
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -149 | -3.394
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -148 | -3.379
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -147 | -3.363
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -146 | -3.347
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -145 | -3.331
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -144 | -3.315
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -143 | -3.299
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -142 | -3.283
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -141 | -3.267
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -140 | -3.251
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -139 | -3.235
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -138 | -3.219
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -137 | -3.203
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -136 | -3.187
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -135 | -3.171
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -134 | -3.154
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -133 | -3.138
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -132 | -3.122
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -131 | -3.105
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -130 | -3.089
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -129 | -3.072
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -128 | -3.056
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -127 | -3.04
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -126 | -3.023
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -125 | -3.006
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -124 | -2.99
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -123 | -2.973
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -122 | -2.956
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -121 | -2.94
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -120 | -2.923
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -119 | -2.906
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -118 | -2.889
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -117 | -2.873
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -116 | -2.856
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -115 | -2.839
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -114 | -2.822
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -113 | -2.805
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -112 | -2.788
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -111 | -2.771
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -110 | -2.754
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -109 | -2.737
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -108 | -2.719
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -107 | -2.702
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -106 | -2.685
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -105 | -2.668
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -104 | -2.651
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -103 | -2.633
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -102 | -2.616
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -101 | -2.598
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -100 | -2.581
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -99 | -2.564
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -98 | -2.546
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -97 | -2.529
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -96 | -2.511
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -95 | -2.493
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -94 | -2.476
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -93 | -2.458
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -92 | -2.44
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -91 | -2.423
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -90 | -2.405
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -89 | -2.387
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -88 | -2.369
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -87 | -2.351
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -86 | -2.334
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -85 | -2.316
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -84 | -2.298
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -83 | -2.28
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -82 | -2.262
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -81 | -2.244
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -80 | -2.225
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -79 | -2.207
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -78 | -2.189
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -77 | -2.171
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -76 | -2.153
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -75 | -2.134
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -74 | -2.116
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -73 | -2.098
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -72 | -2.079
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -71 | -2.061
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -70 | -2.043
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -69 | -2.024
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -68 | -2.006
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -67 | -1.987
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -66 | -1.969
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -65 | -1.95
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -64 | -1.931
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -63 | -1.913
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -62 | -1.894
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -61 | -1.875
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -60 | -1.857
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -59 | -1.838
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -58 | -1.819
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -57 | -1.8
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -56 | -1.781
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -55 | -1.762
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -54 | -1.743
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -53 | -1.724
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -52 | -1.705
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -51 | -1.686
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -50 | -1.667
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -49 | -1.648
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -48 | -1.629
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -47 | -1.61
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -46 | -1.591
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -45 | -1.572
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -44 | -1.552
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -43 | -1.533
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -42 | -1.514
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -41 | -1.494
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -40 | -1.475
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -39 | -1.456
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -38 | -1.436
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -37 | -1.417
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -36 | -1.397
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -35 | -1.378
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -34 | -1.358
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -33 | -1.338
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -32 | -1.319
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -31 | -1.299
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -30 | -1.279
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -29 | -1.26
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -28 | -1.24
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -27 | -1.22
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -26 | -1.2
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -25 | -1.181
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -24 | -1.161
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -23 | -1.141
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -22 | -1.121
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -21 | -1.101
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -20 | -1.081
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -19 | -1.061
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -18 | -1.041
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -17 | -1.021
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -16 | -1.001
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -15 | -0.98
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -14 | -0.96
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -13 | -0.94
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -12 | -0.92
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -11 | -0.9
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -10 | -0.879
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -9 | -0.859
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -8 | -0.839
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -7 | -0.818
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -6 | -0.798
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -5 | -0.777
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -4 | -0.757
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -3 | -0.736
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -2 | -0.716
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | -1 | -0.695
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 0 | -0.675
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 1 | -0.654
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 2 | -0.633
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 3 | -0.613
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 4 | -0.592
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 5 | -0.571
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 6 | -0.55
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 7 | -0.53
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 8 | -0.509
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 9 | -0.488
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 10 | -0.467
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 11 | -0.446
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 12 | -0.425
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 13 | -0.404
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 14 | -0.383
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 15 | -0.362
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 16 | -0.341
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 17 | -0.32
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 18 | -0.299
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 19 | -0.278
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 20 | -0.256
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 21 | -0.235
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 22 | -0.214
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 23 | -0.193
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 24 | -0.171
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 25 | -0.15
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 26 | -0.129
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 27 | -0.107
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 28 | -0.086
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 29 | -0.064
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 30 | -0.043
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 31 | -0.022
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 32 | 0
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 33 | 0.022
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 34 | 0.043
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 35 | 0.065
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 36 | 0.086
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 37 | 0.108
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 38 | 0.13
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 39 | 0.151
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 40 | 0.173
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 41 | 0.195
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 42 | 0.216
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 43 | 0.238
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 44 | 0.26
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 45 | 0.282
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 46 | 0.303
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 47 | 0.325
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 48 | 0.347
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 49 | 0.369
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 50 | 0.391
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 51 | 0.413
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 52 | 0.435
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 53 | 0.457
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 54 | 0.479
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 55 | 0.501
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 56 | 0.523
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 57 | 0.545
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 58 | 0.567
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 59 | 0.589
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 60 | 0.611
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 61 | 0.634
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 62 | 0.656
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 63 | 0.678
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 64 | 0.7
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 65 | 0.723
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 66 | 0.745
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 67 | 0.767
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 68 | 0.79
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 69 | 0.812
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 70 | 0.834
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 71 | 0.857
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 72 | 0.879
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 73 | 0.902
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 74 | 0.924
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 75 | 0.947
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 76 | 0.969
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 77 | 0.992
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 78 | 1.015
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 79 | 1.037
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 80 | 1.06
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 81 | 1.083
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 82 | 1.105
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 83 | 1.128
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 84 | 1.151
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 85 | 1.174
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 86 | 1.196
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 87 | 1.219
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 88 | 1.242
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 89 | 1.265
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 90 | 1.288
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 91 | 1.311
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 92 | 1.334
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 93 | 1.357
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 94 | 1.38
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 95 | 1.403
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 96 | 1.426
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 97 | 1.449
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 98 | 1.472
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 99 | 1.496
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 100 | 1.519
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 101 | 1.542
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 102 | 1.565
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 103 | 1.588
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 104 | 1.612
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 105 | 1.635
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 106 | 1.658
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 107 | 1.682
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 108 | 1.705
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 109 | 1.729
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 110 | 1.752
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 111 | 1.776
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 112 | 1.799
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 113 | 1.823
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 114 | 1.846
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 115 | 1.87
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 116 | 1.893
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 117 | 1.917
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 118 | 1.941
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 119 | 1.964
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 120 | 1.988
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 121 | 2.012
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 122 | 2.036
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 123 | 2.06
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 124 | 2.083
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 125 | 2.107
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 126 | 2.131
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 127 | 2.155
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 128 | 2.179
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 129 | 2.203
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 130 | 2.227
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 131 | 2.251
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 132 | 2.275
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 133 | 2.299
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 134 | 2.323
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 135 | 2.347
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 136 | 2.371
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 137 | 2.395
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 138 | 2.42
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 139 | 2.444
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 140 | 2.468
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 141 | 2.492
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 142 | 2.517
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 143 | 2.541
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 144 | 2.565
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 145 | 2.59
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 146 | 2.614
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 147 | 2.639
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 148 | 2.663
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 149 | 2.687
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 150 | 2.712
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 151 | 2.737
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 152 | 2.761
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 153 | 2.786
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 154 | 2.81
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 155 | 2.835
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 156 | 2.86
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 157 | 2.884
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 158 | 2.909
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 159 | 2.934
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 160 | 2.958
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 161 | 2.983
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 162 | 3.008
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 163 | 3.033
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 164 | 3.058
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 165 | 3.082
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 166 | 3.107
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 167 | 3.132
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 168 | 3.157
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 169 | 3.182
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 170 | 3.207
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 171 | 3.232
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 172 | 3.257
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 173 | 3.282
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 174 | 3.307
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 175 | 3.333
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 176 | 3.358
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 177 | 3.383
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 178 | 3.408
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 179 | 3.433
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 180 | 3.459
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 181 | 3.484
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 182 | 3.509
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 183 | 3.534
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 184 | 3.56
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 185 | 3.585
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 186 | 3.61
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 187 | 3.636
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 188 | 3.661
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 189 | 3.687
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 190 | 3.712
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 191 | 3.738
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 192 | 3.763
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 193 | 3.789
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 194 | 3.814
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 195 | 3.84
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 196 | 3.866
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 197 | 3.891
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 198 | 3.917
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 199 | 3.943
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 200 | 3.968
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 201 | 3.994
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 202 | 4.02
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 203 | 4.046
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 204 | 4.071
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 205 | 4.097
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 206 | 4.123
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 207 | 4.149
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 208 | 4.175
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 209 | 4.201
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 210 | 4.227
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 211 | 4.253
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 212 | 4.279
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 213 | 4.305
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 214 | 4.331
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 215 | 4.357
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 216 | 4.383
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 217 | 4.409
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 218 | 4.435
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 219 | 4.461
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 220 | 4.487
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 221 | 4.513
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 222 | 4.54
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 223 | 4.566
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 224 | 4.592
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 225 | 4.618
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 226 | 4.645
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 227 | 4.671
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 228 | 4.697
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 229 | 4.724
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 230 | 4.75
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 231 | 4.776
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 232 | 4.803
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 233 | 4.829
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 234 | 4.856
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 235 | 4.882
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 236 | 4.909
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 237 | 4.935
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 238 | 4.962
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 239 | 4.988
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 240 | 5.015
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 241 | 5.042
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 242 | 5.068
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 243 | 5.095
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 244 | 5.122
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 245 | 5.148
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 246 | 5.175
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 247 | 5.202
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 248 | 5.228
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 249 | 5.255
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 250 | 5.282
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 251 | 5.309
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 252 | 5.336
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 253 | 5.363
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 254 | 5.389
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 255 | 5.416
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 256 | 5.443
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 257 | 5.47
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 258 | 5.497
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 259 | 5.524
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 260 | 5.551
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 261 | 5.578
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 262 | 5.605
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 263 | 5.632
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 264 | 5.66
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 265 | 5.687
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 266 | 5.714
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 267 | 5.741
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 268 | 5.768
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 269 | 5.795
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 270 | 5.823
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 271 | 5.85
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 272 | 5.877
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 273 | 5.904
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 274 | 5.932
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 275 | 5.959
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 276 | 5.986
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 277 | 6.014
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 278 | 6.041
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 279 | 6.068
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 280 | 6.096
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 281 | 6.123
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 282 | 6.151
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 283 | 6.178
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 284 | 6.206
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 285 | 6.233
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 286 | 6.261
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 287 | 6.288
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 288 | 6.316
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 289 | 6.343
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 290 | 6.371
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 291 | 6.399
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 292 | 6.426
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 293 | 6.454
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 294 | 6.482
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 295 | 6.51
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 296 | 6.537
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 297 | 6.565
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 298 | 6.593
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 299 | 6.621
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 300 | 6.648
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 301 | 6.676
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 302 | 6.704
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 303 | 6.732
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 304 | 6.76
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 305 | 6.788
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 306 | 6.816
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 307 | 6.844
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 308 | 6.872
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 309 | 6.9
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 310 | 6.928
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 311 | 6.956
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 312 | 6.984
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 313 | 7.012
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 314 | 7.04
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 315 | 7.068
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 316 | 7.096
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 317 | 7.124
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 318 | 7.152
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 319 | 7.181
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 320 | 7.209
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 321 | 7.237
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 322 | 7.265
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 323 | 7.294
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 324 | 7.322
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 325 | 7.35
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 326 | 7.378
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 327 | 7.407
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 328 | 7.435
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 329 | 7.463
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 330 | 7.492
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 331 | 7.52
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 332 | 7.549
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 333 | 7.577
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 334 | 7.606
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 335 | 7.634
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 336 | 7.663
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 337 | 7.691
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 338 | 7.72
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 339 | 7.748
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 340 | 7.777
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 341 | 7.805
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 342 | 7.834
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 343 | 7.863
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 344 | 7.891
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 345 | 7.92
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 346 | 7.949
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 347 | 7.977
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 348 | 8.006
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 349 | 8.035
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 350 | 8.064
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 351 | 8.092
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 352 | 8.121
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 353 | 8.15
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 354 | 8.179
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 355 | 8.208
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 356 | 8.237
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 357 | 8.266
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 358 | 8.294
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 359 | 8.323
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 360 | 8.352
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 361 | 8.381
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 362 | 8.41
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 363 | 8.439
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 364 | 8.468
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 365 | 8.497
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 366 | 8.526
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 367 | 8.555
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 368 | 8.585
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 369 | 8.614
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 370 | 8.643
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 371 | 8.672
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 372 | 8.701
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 373 | 8.73
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 374 | 8.759
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 375 | 8.789
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 376 | 8.818
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 377 | 8.847
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 378 | 8.876
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 379 | 8.906
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 380 | 8.935
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 381 | 8.964
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 382 | 8.994
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 383 | 9.023
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 384 | 9.052
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 385 | 9.082
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 386 | 9.111
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 387 | 9.141
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 388 | 9.17
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 389 | 9.2
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 390 | 9.229
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 391 | 9.259
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 392 | 9.288
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 393 | 9.318
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 394 | 9.347
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 395 | 9.377
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 396 | 9.406
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 397 | 9.436
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 398 | 9.466
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 399 | 9.495
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 400 | 9.525
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 401 | 9.555
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 402 | 9.584
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 403 | 9.614
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 404 | 9.644
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 405 | 9.673
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 406 | 9.703
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 407 | 9.733
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 408 | 9.763
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 409 | 9.793
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 410 | 9.822
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 411 | 9.852
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 412 | 9.882
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 413 | 9.912
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 414 | 9.942
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 415 | 9.972
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 416 | 10.002
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 417 | 10.032
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 418 | 10.062
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 419 | 10.092
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 420 | 10.122
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 421 | 10.152
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 422 | 10.182
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 423 | 10.212
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 424 | 10.242
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 425 | 10.272
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 426 | 10.302
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 427 | 10.332
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 428 | 10.362
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 429 | 10.392
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 430 | 10.423
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 431 | 10.453
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 432 | 10.483
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 433 | 10.513
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 434 | 10.543
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 435 | 10.574
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 436 | 10.604
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 437 | 10.634
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 438 | 10.664
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 439 | 10.695
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 440 | 10.725
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 441 | 10.755
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 442 | 10.786
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 443 | 10.816
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 444 | 10.847
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 445 | 10.877
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 446 | 10.907
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 447 | 10.938
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 448 | 10.968
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 449 | 10.999
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 450 | 11.029
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 451 | 11.06
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 452 | 11.09
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 453 | 11.121
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 454 | 11.151
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 455 | 11.182
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 456 | 11.213
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 457 | 11.243
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 458 | 11.274
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 459 | 11.304
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 460 | 11.335
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 461 | 11.366
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 462 | 11.396
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 463 | 11.427
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 464 | 11.458
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 465 | 11.489
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 466 | 11.519
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 467 | 11.55
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 468 | 11.581
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 469 | 11.612
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 470 | 11.643
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 471 | 11.673
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 472 | 11.704
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 473 | 11.735
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 474 | 11.766
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 475 | 11.797
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 476 | 11.828
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 477 | 11.859
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 478 | 11.89
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 479 | 11.92
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 480 | 11.951
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 481 | 11.982
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 482 | 12.013
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 483 | 12.044
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 484 | 12.075
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 485 | 12.106
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 486 | 12.138
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 487 | 12.169
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 488 | 12.2
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 489 | 12.231
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 490 | 12.262
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 491 | 12.293
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 492 | 12.324
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 493 | 12.355
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 494 | 12.386
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 495 | 12.418
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 496 | 12.449
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 497 | 12.48
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 498 | 12.511
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 499 | 12.543
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 500 | 12.574
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 501 | 12.605
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 502 | 12.636
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 503 | 12.668
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 504 | 12.699
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 505 | 12.73
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 506 | 12.762
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 507 | 12.793
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 508 | 12.824
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 509 | 12.856
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 510 | 12.887
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 511 | 12.919
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 512 | 12.95
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 513 | 12.982
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 514 | 13.013
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 515 | 13.045
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 516 | 13.076
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 517 | 13.108
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 518 | 13.139
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 519 | 13.171
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 520 | 13.202
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 521 | 13.234
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 522 | 13.265
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 523 | 13.297
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 524 | 13.328
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 525 | 13.36
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 526 | 13.392
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 527 | 13.423
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 528 | 13.455
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 529 | 13.487
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 530 | 13.518
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 531 | 13.55
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 532 | 13.582
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 533 | 13.614
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 534 | 13.645
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 535 | 13.677
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 536 | 13.709
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 537 | 13.741
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 538 | 13.772
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 539 | 13.804
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 540 | 13.836
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 541 | 13.868
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 542 | 13.9
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 543 | 13.932
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 544 | 13.964
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 545 | 13.995
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 546 | 14.027
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 547 | 14.059
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 548 | 14.091
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 549 | 14.123
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 550 | 14.155
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 551 | 14.187
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 552 | 14.219
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 553 | 14.251
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 554 | 14.283
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 555 | 14.315
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 556 | 14.347
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 557 | 14.379
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 558 | 14.411
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 559 | 14.444
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 560 | 14.476
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 561 | 14.508
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 562 | 14.54
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 563 | 14.572
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 564 | 14.604
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 565 | 14.636
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 566 | 14.669
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 567 | 14.701
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 568 | 14.733
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 569 | 14.765
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 570 | 14.797
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 571 | 14.83
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 572 | 14.862
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 573 | 14.894
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 574 | 14.926
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 575 | 14.959
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 576 | 14.991
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 577 | 15.023
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 578 | 15.056
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 579 | 15.088
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 580 | 15.121
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 581 | 15.153
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 582 | 15.185
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 583 | 15.218
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 584 | 15.25
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 585 | 15.283
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 586 | 15.315
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 587 | 15.347
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 588 | 15.38
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 589 | 15.412
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 590 | 15.445
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 591 | 15.477
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 592 | 15.51
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 593 | 15.543
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 594 | 15.575
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 595 | 15.608
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 596 | 15.64
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 597 | 15.673
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 598 | 15.705
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 599 | 15.738
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 600 | 15.771
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 601 | 15.803
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 602 | 15.836
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 603 | 15.869
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 604 | 15.901
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 605 | 15.934
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 606 | 15.967
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 607 | 15.999
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 608 | 16.032
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 609 | 16.065
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 610 | 16.098
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 611 | 16.13
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 612 | 16.163
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 613 | 16.196
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 614 | 16.229
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 615 | 16.262
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 616 | 16.295
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 617 | 16.327
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 618 | 16.36
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 619 | 16.393
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 620 | 16.426
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 621 | 16.459
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 622 | 16.492
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 623 | 16.525
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 624 | 16.558
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 625 | 16.591
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 626 | 16.624
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 627 | 16.657
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 628 | 16.69
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 629 | 16.723
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 630 | 16.756
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 631 | 16.789
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 632 | 16.822
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 633 | 16.855
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 634 | 16.888
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 635 | 16.921
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 636 | 16.954
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 637 | 16.987
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 638 | 17.02
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 639 | 17.053
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 640 | 17.086
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 641 | 17.12
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 642 | 17.153
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 643 | 17.186
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 644 | 17.219
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 645 | 17.252
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 646 | 17.286
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 647 | 17.319
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 648 | 17.352
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 649 | 17.385
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 650 | 17.418
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 651 | 17.452
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 652 | 17.485
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 653 | 17.518
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 654 | 17.552
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 655 | 17.585
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 656 | 17.618
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 657 | 17.652
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 658 | 17.685
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 659 | 17.718
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 660 | 17.752
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 661 | 17.785
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 662 | 17.819
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 663 | 17.852
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 664 | 17.886
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 665 | 17.919
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 666 | 17.952
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 667 | 17.986
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 668 | 18.019
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 669 | 18.053
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 670 | 18.086
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 671 | 18.12
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 672 | 18.154
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 673 | 18.187
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 674 | 18.221
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 675 | 18.254
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 676 | 18.288
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 677 | 18.321
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 678 | 18.355
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 679 | 18.389
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 680 | 18.422
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 681 | 18.456
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 682 | 18.49
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 683 | 18.523
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 684 | 18.557
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 685 | 18.591
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 686 | 18.624
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 687 | 18.658
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 688 | 18.692
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 689 | 18.725
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 690 | 18.759
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 691 | 18.793
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 692 | 18.827
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 693 | 18.861
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 694 | 18.894
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 695 | 18.928
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 696 | 18.962
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 697 | 18.996
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 698 | 19.03
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 699 | 19.064
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 700 | 19.097
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 701 | 19.131
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 702 | 19.165
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 703 | 19.199
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 704 | 19.233
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 705 | 19.267
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 706 | 19.301
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 707 | 19.335
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 708 | 19.369
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 709 | 19.403
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 710 | 19.437
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 711 | 19.471
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 712 | 19.505
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 713 | 19.539
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 714 | 19.573
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 715 | 19.607
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 716 | 19.641
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 717 | 19.675
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 718 | 19.709
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 719 | 19.743
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 720 | 19.777
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 721 | 19.811
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 722 | 19.845
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 723 | 19.879
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 724 | 19.913
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 725 | 19.947
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 726 | 19.982
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 727 | 20.016
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 728 | 20.05
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 729 | 20.084
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 730 | 20.118
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 731 | 20.152
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 732 | 20.187
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 733 | 20.221
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 734 | 20.255
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 735 | 20.289
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 736 | 20.323
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 737 | 20.358
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 738 | 20.392
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 739 | 20.426
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 740 | 20.46
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 741 | 20.495
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 742 | 20.529
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 743 | 20.563
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 744 | 20.597
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 745 | 20.632
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 746 | 20.666
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 747 | 20.7
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 748 | 20.735
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 749 | 20.769
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 750 | 20.803
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 751 | 20.838
0 | 32 | 486 | -0.022 | 0 | 31 | 32 | 752 | 20.872
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32
0 | 32 | 486 | -0.022 | 0 | 31 | 32

## Type S
[9 empty] | Temp(F) | mV
mV | Temp(F) | [7 empty] | -58 | -0.236
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -57 | -0.233
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -56 | -0.231
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -55 | -0.229
0.0001 | 32.0333 | 91 | 0 | 0.003 | 32 | 33 |  |  | -54 | -0.227
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -53 | -0.224
0.5979 | 200.225 | 259 | 0.597 | 0.601 | 200 | 201 |  |  | -52 | -0.222
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -51 | -0.22
1.4762 | 399.55 | 458 | 1.474 | 1.478 | 399 | 400 |  |  | -50 | -0.218
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -49 | -0.215
2.4648 | 599.76 | 658 | 2.461 | 2.466 | 599 | 600 |  |  | -48 | -0.213
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -47 | -0.211
3.5165 | 800.1 | 859 | 3.516 | 3.521 | 800 | 801 |  |  | -46 | -0.208
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -45 | -0.206
4.6113 | 1000.26 | 1059 | 4.61 | 4.615 | 1000 | 1001 |  |  | -44 | -0.204
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -43 | -0.201
5.74 | 1199.8 | 1258 | 5.736 | 5.741 | 1199 | 1200 |  |  | -42 | -0.199
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -41 | -0.197
6.8989 | 1397.65 | 1456 | 6.895 | 6.901 | 1397 | 1398 |  |  | -40 | -0.194
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -39 | -0.192
8.1146 | 1598.0857 | 1657 | 8.114 | 8.121 | 1598 | 1599 |  |  | -38 | -0.19
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -37 | -0.187
9.37 | 1798 | 1856 | 9.363 | 9.37 | 1797 | 1798 |  |  | -36 | -0.185
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -35 | -0.182
10.6897 | 2001.8143 | 2060 | 10.684 | 10.691 | 2001 | 2002 |  |  | -34 | -0.18
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -33 | -0.178
12.003 | 2199.8571 | 2258 | 11.997 | 12.004 | 2199 | 2200 |  |  | -32 | -0.175
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -31 | -0.173
NO | #VALUE! | 0 | mV | -0.236 | Temp(F) | -58 |  |  | -30 | -0.17
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -29 | -0.168
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -28 | -0.165
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -27 | -0.163
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -26 | -0.16
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -25 | -0.158
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -24 | -0.155
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -23 | -0.153
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -22 | -0.15
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -21 | -0.148
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -20 | -0.145
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -19 | -0.142
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -18 | -0.14
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -17 | -0.137
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -16 | -0.135
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -15 | -0.132
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -14 | -0.129
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -13 | -0.127
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -12 | -0.124
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -11 | -0.122
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -10 | -0.119
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -9 | -0.116
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -8 | -0.114
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -7 | -0.111
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -6 | -0.108
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -5 | -0.106
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -4 | -0.103
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -3 | -0.1
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -2 | -0.097
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | -1 | -0.095
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 0 | -0.092
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1 | -0.089
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 2 | -0.086
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 3 | -0.084
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 4 | -0.081
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 5 | -0.078
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 6 | -0.075
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 7 | -0.073
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 8 | -0.07
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 9 | -0.067
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 10 | -0.064
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 11 | -0.061
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 12 | -0.058
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 13 | -0.056
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 14 | -0.053
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 15 | -0.05
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 16 | -0.047
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 17 | -0.044
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 18 | -0.041
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 19 | -0.038
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 20 | -0.035
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 21 | -0.033
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 22 | -0.03
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 23 | -0.027
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 24 | -0.024
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 25 | -0.021
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 26 | -0.018
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 27 | -0.015
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 28 | -0.012
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 29 | -0.009
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 30 | -0.006
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 31 | -0.003
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 32 | 0
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 33 | 0.003
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 34 | 0.006
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 35 | 0.009
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 36 | 0.012
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 37 | 0.015
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 38 | 0.018
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 39 | 0.021
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 40 | 0.024
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 41 | 0.027
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 42 | 0.03
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 43 | 0.033
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 44 | 0.037
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 45 | 0.04
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 46 | 0.043
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 47 | 0.046
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 48 | 0.049
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 49 | 0.052
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 50 | 0.055
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 51 | 0.058
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 52 | 0.062
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 53 | 0.065
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 54 | 0.068
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 55 | 0.071
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 56 | 0.074
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 57 | 0.077
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 58 | 0.081
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 59 | 0.084
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 60 | 0.087
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 61 | 0.09
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 62 | 0.093
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 63 | 0.097
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 64 | 0.1
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 65 | 0.103
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 66 | 0.106
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 67 | 0.11
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 68 | 0.113
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 69 | 0.116
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 70 | 0.119
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 71 | 0.123
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 72 | 0.126
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 73 | 0.129
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 74 | 0.133
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 75 | 0.136
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 76 | 0.139
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 77 | 0.143
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 78 | 0.146
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 79 | 0.149
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 80 | 0.153
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 81 | 0.156
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 82 | 0.159
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 83 | 0.163
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 84 | 0.166
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 85 | 0.169
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 86 | 0.173
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 87 | 0.176
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 88 | 0.18
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 89 | 0.183
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 90 | 0.186
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 91 | 0.19
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 92 | 0.193
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 93 | 0.197
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 94 | 0.2
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 95 | 0.204
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 96 | 0.207
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 97 | 0.21
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 98 | 0.214
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 99 | 0.217
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 100 | 0.221
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 101 | 0.224
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 102 | 0.228
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 103 | 0.231
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 104 | 0.235
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 105 | 0.238
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 106 | 0.242
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 107 | 0.245
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 108 | 0.249
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 109 | 0.252
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 110 | 0.256
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 111 | 0.26
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 112 | 0.263
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 113 | 0.267
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 114 | 0.27
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 115 | 0.274
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 116 | 0.277
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 117 | 0.281
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 118 | 0.285
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 119 | 0.288
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 120 | 0.292
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 121 | 0.295
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 122 | 0.299
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 123 | 0.303
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 124 | 0.306
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 125 | 0.31
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 126 | 0.313
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 127 | 0.317
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 128 | 0.321
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 129 | 0.324
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 130 | 0.328
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 131 | 0.332
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 132 | 0.335
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 133 | 0.339
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 134 | 0.343
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 135 | 0.346
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 136 | 0.35
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 137 | 0.354
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 138 | 0.357
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 139 | 0.361
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 140 | 0.365
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 141 | 0.369
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 142 | 0.372
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 143 | 0.376
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 144 | 0.38
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 145 | 0.384
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 146 | 0.387
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 147 | 0.391
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 148 | 0.395
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 149 | 0.399
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 150 | 0.402
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 151 | 0.406
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 152 | 0.41
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 153 | 0.414
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 154 | 0.417
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 155 | 0.421
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 156 | 0.425
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 157 | 0.429
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 158 | 0.433
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 159 | 0.436
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 160 | 0.44
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 161 | 0.444
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 162 | 0.448
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 163 | 0.452
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 164 | 0.456
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 165 | 0.459
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 166 | 0.463
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 167 | 0.467
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 168 | 0.471
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 169 | 0.475
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 170 | 0.479
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 171 | 0.483
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 172 | 0.487
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 173 | 0.49
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 174 | 0.494
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 175 | 0.498
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 176 | 0.502
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 177 | 0.506
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 178 | 0.51
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 179 | 0.514
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 180 | 0.518
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 181 | 0.522
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 182 | 0.526
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 183 | 0.53
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 184 | 0.534
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 185 | 0.538
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 186 | 0.541
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 187 | 0.545
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 188 | 0.549
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 189 | 0.553
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 190 | 0.557
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 191 | 0.561
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 192 | 0.565
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 193 | 0.569
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 194 | 0.573
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 195 | 0.577
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 196 | 0.581
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 197 | 0.585
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 198 | 0.589
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 199 | 0.593
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 200 | 0.597
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 201 | 0.601
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 202 | 0.605
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 203 | 0.609
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 204 | 0.613
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 205 | 0.617
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 206 | 0.622
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 207 | 0.626
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 208 | 0.63
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 209 | 0.634
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 210 | 0.638
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 211 | 0.642
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 212 | 0.646
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 213 | 0.65
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 214 | 0.654
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 215 | 0.658
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 216 | 0.662
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 217 | 0.666
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 218 | 0.67
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 219 | 0.675
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 220 | 0.679
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 221 | 0.683
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 222 | 0.687
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 223 | 0.691
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 224 | 0.695
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 225 | 0.699
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 226 | 0.703
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 227 | 0.708
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 228 | 0.712
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 229 | 0.716
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 230 | 0.72
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 231 | 0.724
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 232 | 0.728
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 233 | 0.732
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 234 | 0.737
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 235 | 0.741
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 236 | 0.745
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 237 | 0.749
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 238 | 0.753
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 239 | 0.758
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 240 | 0.762
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 241 | 0.766
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 242 | 0.77
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 243 | 0.774
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 244 | 0.779
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 245 | 0.783
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 246 | 0.787
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 247 | 0.791
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 248 | 0.795
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 249 | 0.8
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 250 | 0.804
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 251 | 0.808
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 252 | 0.812
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 253 | 0.817
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 254 | 0.821
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 255 | 0.825
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 256 | 0.829
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 257 | 0.834
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 258 | 0.838
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 259 | 0.842
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 260 | 0.847
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 261 | 0.851
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 262 | 0.855
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 263 | 0.859
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 264 | 0.864
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 265 | 0.868
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 266 | 0.872
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 267 | 0.877
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 268 | 0.881
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 269 | 0.885
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 270 | 0.889
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 271 | 0.894
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 272 | 0.898
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 273 | 0.902
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 274 | 0.907
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 275 | 0.911
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 276 | 0.915
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 277 | 0.92
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 278 | 0.924
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 279 | 0.928
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 280 | 0.933
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 281 | 0.937
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 282 | 0.942
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 283 | 0.946
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 284 | 0.95
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 285 | 0.955
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 286 | 0.959
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 287 | 0.963
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 288 | 0.968
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 289 | 0.972
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 290 | 0.977
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 291 | 0.981
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 292 | 0.985
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 293 | 0.99
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 294 | 0.994
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 295 | 0.998
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 296 | 1.003
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 297 | 1.007
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 298 | 1.012
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 299 | 1.016
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 300 | 1.021
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 301 | 1.025
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 302 | 1.029
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 303 | 1.034
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 304 | 1.038
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 305 | 1.043
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 306 | 1.047
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 307 | 1.052
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 308 | 1.056
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 309 | 1.061
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 310 | 1.065
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 311 | 1.069
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 312 | 1.074
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 313 | 1.078
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 314 | 1.083
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 315 | 1.087
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 316 | 1.092
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 317 | 1.096
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 318 | 1.101
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 319 | 1.105
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 320 | 1.11
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 321 | 1.114
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 322 | 1.119
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 323 | 1.123
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 324 | 1.128
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 325 | 1.132
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 326 | 1.137
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 327 | 1.141
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 328 | 1.146
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 329 | 1.15
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 330 | 1.155
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 331 | 1.159
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 332 | 1.164
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 333 | 1.168
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 334 | 1.173
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 335 | 1.177
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 336 | 1.182
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 337 | 1.186
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 338 | 1.191
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 339 | 1.196
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 340 | 1.2
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 341 | 1.205
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 342 | 1.209
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 343 | 1.214
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 344 | 1.218
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 345 | 1.223
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 346 | 1.227
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 347 | 1.232
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 348 | 1.237
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 349 | 1.241
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 350 | 1.246
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 351 | 1.25
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 352 | 1.255
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 353 | 1.26
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 354 | 1.264
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 355 | 1.269
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 356 | 1.273
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 357 | 1.278
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 358 | 1.283
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 359 | 1.287
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 360 | 1.292
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 361 | 1.296
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 362 | 1.301
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 363 | 1.306
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 364 | 1.31
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 365 | 1.315
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 366 | 1.319
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 367 | 1.324
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 368 | 1.329
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 369 | 1.333
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 370 | 1.338
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 371 | 1.343
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 372 | 1.347
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 373 | 1.352
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 374 | 1.357
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 375 | 1.361
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 376 | 1.366
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 377 | 1.371
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 378 | 1.375
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 379 | 1.38
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 380 | 1.385
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 381 | 1.389
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 382 | 1.394
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 383 | 1.399
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 384 | 1.403
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 385 | 1.408
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 386 | 1.413
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 387 | 1.417
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 388 | 1.422
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 389 | 1.427
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 390 | 1.431
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 391 | 1.436
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 392 | 1.441
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 393 | 1.445
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 394 | 1.45
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 395 | 1.455
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 396 | 1.46
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 397 | 1.464
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 398 | 1.469
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 399 | 1.474
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 400 | 1.478
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 401 | 1.483
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 402 | 1.488
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 403 | 1.493
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 404 | 1.497
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 405 | 1.502
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 406 | 1.507
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 407 | 1.512
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 408 | 1.516
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 409 | 1.521
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 410 | 1.526
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 411 | 1.531
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 412 | 1.535
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 413 | 1.54
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 414 | 1.545
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 415 | 1.55
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 416 | 1.554
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 417 | 1.559
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 418 | 1.564
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 419 | 1.569
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 420 | 1.573
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 421 | 1.578
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 422 | 1.583
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 423 | 1.588
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 424 | 1.592
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 425 | 1.597
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 426 | 1.602
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 427 | 1.607
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 428 | 1.612
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 429 | 1.616
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 430 | 1.621
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 431 | 1.626
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 432 | 1.631
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 433 | 1.636
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 434 | 1.64
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 435 | 1.645
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 436 | 1.65
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 437 | 1.655
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 438 | 1.66
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 439 | 1.664
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 440 | 1.669
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 441 | 1.674
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 442 | 1.679
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 443 | 1.684
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 444 | 1.689
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 445 | 1.693
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 446 | 1.698
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 447 | 1.703
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 448 | 1.708
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 449 | 1.713
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 450 | 1.718
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 451 | 1.722
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 452 | 1.727
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 453 | 1.732
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 454 | 1.737
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 455 | 1.742
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 456 | 1.747
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 457 | 1.752
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 458 | 1.756
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 459 | 1.761
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 460 | 1.766
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 461 | 1.771
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 462 | 1.776
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 463 | 1.781
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 464 | 1.786
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 465 | 1.79
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 466 | 1.795
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 467 | 1.8
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 468 | 1.805
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 469 | 1.81
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 470 | 1.815
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 471 | 1.82
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 472 | 1.825
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 473 | 1.829
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 474 | 1.834
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 475 | 1.839
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 476 | 1.844
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 477 | 1.849
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 478 | 1.854
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 479 | 1.859
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 480 | 1.864
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 481 | 1.869
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 482 | 1.874
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 483 | 1.878
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 484 | 1.883
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 485 | 1.888
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 486 | 1.893
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 487 | 1.898
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 488 | 1.903
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 489 | 1.908
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 490 | 1.913
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 491 | 1.918
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 492 | 1.923
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 493 | 1.928
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 494 | 1.933
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 495 | 1.938
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 496 | 1.942
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 497 | 1.947
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 498 | 1.952
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 499 | 1.957
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 500 | 1.962
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 501 | 1.967
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 502 | 1.972
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 503 | 1.977
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 504 | 1.982
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 505 | 1.987
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 506 | 1.992
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 507 | 1.997
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 508 | 2.002
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 509 | 2.007
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 510 | 2.012
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 511 | 2.017
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 512 | 2.022
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 513 | 2.027
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 514 | 2.032
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 515 | 2.037
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 516 | 2.042
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 517 | 2.047
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 518 | 2.052
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 519 | 2.057
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 520 | 2.062
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 521 | 2.067
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 522 | 2.072
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 523 | 2.076
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 524 | 2.081
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 525 | 2.086
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 526 | 2.091
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 527 | 2.096
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 528 | 2.101
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 529 | 2.106
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 530 | 2.111
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 531 | 2.116
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 532 | 2.121
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 533 | 2.126
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 534 | 2.131
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 535 | 2.136
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 536 | 2.141
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 537 | 2.147
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 538 | 2.152
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 539 | 2.157
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 540 | 2.162
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 541 | 2.167
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 542 | 2.172
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 543 | 2.177
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 544 | 2.182
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 545 | 2.187
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 546 | 2.192
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 547 | 2.197
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 548 | 2.202
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 549 | 2.207
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 550 | 2.212
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 551 | 2.217
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 552 | 2.222
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 553 | 2.227
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 554 | 2.232
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 555 | 2.237
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 556 | 2.242
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 557 | 2.247
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 558 | 2.252
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 559 | 2.257
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 560 | 2.262
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 561 | 2.267
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 562 | 2.272
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 563 | 2.277
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 564 | 2.283
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 565 | 2.288
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 566 | 2.293
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 567 | 2.298
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 568 | 2.303
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 569 | 2.308
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 570 | 2.313
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 571 | 2.318
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 572 | 2.323
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 573 | 2.328
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 574 | 2.333
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 575 | 2.338
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 576 | 2.343
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 577 | 2.348
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 578 | 2.354
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 579 | 2.359
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 580 | 2.364
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 581 | 2.369
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 582 | 2.374
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 583 | 2.379
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 584 | 2.384
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 585 | 2.389
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 586 | 2.394
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 587 | 2.399
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 588 | 2.404
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 589 | 2.41
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 590 | 2.415
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 591 | 2.42
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 592 | 2.425
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 593 | 2.43
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 594 | 2.435
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 595 | 2.44
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 596 | 2.445
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 597 | 2.45
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 598 | 2.455
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 599 | 2.461
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 600 | 2.466
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 601 | 2.471
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 602 | 2.476
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 603 | 2.481
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 604 | 2.486
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 605 | 2.491
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 606 | 2.496
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 607 | 2.502
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 608 | 2.507
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 609 | 2.512
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 610 | 2.517
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 611 | 2.522
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 612 | 2.527
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 613 | 2.532
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 614 | 2.538
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 615 | 2.543
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 616 | 2.548
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 617 | 2.553
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 618 | 2.558
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 619 | 2.563
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 620 | 2.568
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 621 | 2.574
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 622 | 2.579
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 623 | 2.584
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 624 | 2.589
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 625 | 2.594
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 626 | 2.599
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 627 | 2.604
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 628 | 2.61
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 629 | 2.615
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 630 | 2.62
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 631 | 2.625
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 632 | 2.63
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 633 | 2.635
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 634 | 2.641
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 635 | 2.646
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 636 | 2.651
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 637 | 2.656
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 638 | 2.661
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 639 | 2.666
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 640 | 2.672
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 641 | 2.677
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 642 | 2.682
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 643 | 2.687
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 644 | 2.692
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 645 | 2.697
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 646 | 2.703
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 647 | 2.708
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 648 | 2.713
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 649 | 2.718
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 650 | 2.723
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 651 | 2.729
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 652 | 2.734
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 653 | 2.739
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 654 | 2.744
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 655 | 2.749
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 656 | 2.755
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 657 | 2.76
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 658 | 2.765
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 659 | 2.77
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 660 | 2.775
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 661 | 2.781
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 662 | 2.786
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 663 | 2.791
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 664 | 2.796
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 665 | 2.801
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 666 | 2.807
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 667 | 2.812
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 668 | 2.817
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 669 | 2.822
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 670 | 2.827
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 671 | 2.833
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 672 | 2.838
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 673 | 2.843
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 674 | 2.848
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 675 | 2.854
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 676 | 2.859
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 677 | 2.864
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 678 | 2.869
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 679 | 2.874
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 680 | 2.88
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 681 | 2.885
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 682 | 2.89
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 683 | 2.895
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 684 | 2.901
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 685 | 2.906
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 686 | 2.911
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 687 | 2.916
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 688 | 2.922
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 689 | 2.927
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 690 | 2.932
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 691 | 2.937
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 692 | 2.943
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 693 | 2.948
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 694 | 2.953
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 695 | 2.958
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 696 | 2.964
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 697 | 2.969
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 698 | 2.974
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 699 | 2.979
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 700 | 2.985
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 701 | 2.99
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 702 | 2.995
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 703 | 3
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 704 | 3.006
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 705 | 3.011
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 706 | 3.016
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 707 | 3.021
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 708 | 3.027
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 709 | 3.032
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 710 | 3.037
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 711 | 3.042
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 712 | 3.048
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 713 | 3.053
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 714 | 3.058
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 715 | 3.063
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 716 | 3.069
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 717 | 3.074
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 718 | 3.079
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 719 | 3.085
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 720 | 3.09
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 721 | 3.095
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 722 | 3.1
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 723 | 3.106
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 724 | 3.111
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 725 | 3.116
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 726 | 3.122
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 727 | 3.127
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 728 | 3.132
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 729 | 3.137
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 730 | 3.143
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 731 | 3.148
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 732 | 3.153
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 733 | 3.159
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 734 | 3.164
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 735 | 3.169
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 736 | 3.174
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 737 | 3.18
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 738 | 3.185
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 739 | 3.19
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 740 | 3.196
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 741 | 3.201
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 742 | 3.206
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 743 | 3.212
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 744 | 3.217
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 745 | 3.222
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 746 | 3.227
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 747 | 3.233
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 748 | 3.238
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 749 | 3.243
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 750 | 3.249
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 751 | 3.254
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 752 | 3.259
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 753 | 3.265
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 754 | 3.27
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 755 | 3.275
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 756 | 3.281
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 757 | 3.286
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 758 | 3.291
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 759 | 3.297
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 760 | 3.302
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 761 | 3.307
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 762 | 3.313
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 763 | 3.318
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 764 | 3.323
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 765 | 3.329
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 766 | 3.334
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 767 | 3.339
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 768 | 3.345
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 769 | 3.35
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 770 | 3.355
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 771 | 3.361
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 772 | 3.366
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 773 | 3.371
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 774 | 3.377
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 775 | 3.382
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 776 | 3.387
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 777 | 3.393
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 778 | 3.398
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 779 | 3.403
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 780 | 3.409
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 781 | 3.414
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 782 | 3.419
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 783 | 3.425
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 784 | 3.43
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 785 | 3.435
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 786 | 3.441
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 787 | 3.446
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 788 | 3.451
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 789 | 3.457
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 790 | 3.462
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 791 | 3.468
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 792 | 3.473
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 793 | 3.478
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 794 | 3.484
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 795 | 3.489
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 796 | 3.494
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 797 | 3.5
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 798 | 3.505
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 799 | 3.51
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 800 | 3.516
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 801 | 3.521
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 802 | 3.527
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 803 | 3.532
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 804 | 3.537
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 805 | 3.543
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 806 | 3.548
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 807 | 3.553
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 808 | 3.559
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 809 | 3.564
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 810 | 3.57
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 811 | 3.575
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 812 | 3.58
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 813 | 3.586
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 814 | 3.591
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 815 | 3.596
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 816 | 3.602
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 817 | 3.607
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 818 | 3.613
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 819 | 3.618
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 820 | 3.623
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 821 | 3.629
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 822 | 3.634
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 823 | 3.64
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 824 | 3.645
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 825 | 3.65
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 826 | 3.656
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 827 | 3.661
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 828 | 3.667
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 829 | 3.672
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 830 | 3.677
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 831 | 3.683
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 832 | 3.688
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 833 | 3.694
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 834 | 3.699
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 835 | 3.704
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 836 | 3.71
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 837 | 3.715
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 838 | 3.721
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 839 | 3.726
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 840 | 3.731
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 841 | 3.737
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 842 | 3.742
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 843 | 3.748
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 844 | 3.753
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 845 | 3.758
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 846 | 3.764
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 847 | 3.769
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 848 | 3.775
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 849 | 3.78
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 850 | 3.786
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 851 | 3.791
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 852 | 3.796
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 853 | 3.802
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 854 | 3.807
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 855 | 3.813
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 856 | 3.818
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 857 | 3.823
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 858 | 3.829
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 859 | 3.834
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 860 | 3.84
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 861 | 3.845
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 862 | 3.851
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 863 | 3.856
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 864 | 3.862
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 865 | 3.867
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 866 | 3.872
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 867 | 3.878
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 868 | 3.883
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 869 | 3.889
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 870 | 3.894
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 871 | 3.9
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 872 | 3.905
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 873 | 3.91
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 874 | 3.916
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 875 | 3.921
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 876 | 3.927
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 877 | 3.932
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 878 | 3.938
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 879 | 3.943
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 880 | 3.949
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 881 | 3.954
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 882 | 3.959
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 883 | 3.965
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 884 | 3.97
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 885 | 3.976
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 886 | 3.981
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 887 | 3.987
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 888 | 3.992
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 889 | 3.998
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 890 | 4.003
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 891 | 4.009
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 892 | 4.014
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 893 | 4.02
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 894 | 4.025
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 895 | 4.03
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 896 | 4.036
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 897 | 4.041
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 898 | 4.047
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 899 | 4.052
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 900 | 4.058
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 901 | 4.063
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 902 | 4.069
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 903 | 4.074
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 904 | 4.08
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 905 | 4.085
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 906 | 4.091
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 907 | 4.096
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 908 | 4.102
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 909 | 4.107
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 910 | 4.113
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 911 | 4.118
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 912 | 4.123
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 913 | 4.129
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 914 | 4.134
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 915 | 4.14
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 916 | 4.145
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 917 | 4.151
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 918 | 4.156
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 919 | 4.162
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 920 | 4.167
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 921 | 4.173
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 922 | 4.178
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 923 | 4.184
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 924 | 4.189
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 925 | 4.195
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 926 | 4.2
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 927 | 4.206
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 928 | 4.211
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 929 | 4.217
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 930 | 4.222
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 931 | 4.228
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 932 | 4.233
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 933 | 4.239
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 934 | 4.244
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 935 | 4.25
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 936 | 4.255
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 937 | 4.261
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 938 | 4.266
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 939 | 4.272
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 940 | 4.277
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 941 | 4.283
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 942 | 4.288
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 943 | 4.294
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 944 | 4.299
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 945 | 4.305
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 946 | 4.31
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 947 | 4.316
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 948 | 4.321
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 949 | 4.327
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 950 | 4.332
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 951 | 4.338
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 952 | 4.343
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 953 | 4.349
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 954 | 4.355
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 955 | 4.36
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 956 | 4.366
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 957 | 4.371
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 958 | 4.377
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 959 | 4.382
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 960 | 4.388
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 961 | 4.393
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 962 | 4.399
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 963 | 4.404
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 964 | 4.41
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 965 | 4.415
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 966 | 4.421
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 967 | 4.426
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 968 | 4.432
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 969 | 4.437
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 970 | 4.443
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 971 | 4.449
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 972 | 4.454
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 973 | 4.46
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 974 | 4.465
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 975 | 4.471
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 976 | 4.476
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 977 | 4.482
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 978 | 4.487
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 979 | 4.493
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 980 | 4.498
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 981 | 4.504
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 982 | 4.51
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 983 | 4.515
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 984 | 4.521
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 985 | 4.526
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 986 | 4.532
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 987 | 4.537
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 988 | 4.543
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 989 | 4.548
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 990 | 4.554
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 991 | 4.559
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 992 | 4.565
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 993 | 4.571
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 994 | 4.576
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 995 | 4.582
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 996 | 4.587
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 997 | 4.593
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 998 | 4.598
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 999 | 4.604
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1000 | 4.61
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1001 | 4.615
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1002 | 4.621
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1003 | 4.626
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1004 | 4.632
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1005 | 4.637
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1006 | 4.643
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1007 | 4.648
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1008 | 4.654
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1009 | 4.66
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1010 | 4.665
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1011 | 4.671
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1012 | 4.676
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1013 | 4.682
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1014 | 4.688
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1015 | 4.693
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1016 | 4.699
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1017 | 4.704
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1018 | 4.71
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1019 | 4.715
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1020 | 4.721
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1021 | 4.727
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1022 | 4.732
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1023 | 4.738
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1024 | 4.743
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1025 | 4.749
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1026 | 4.755
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1027 | 4.76
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1028 | 4.766
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1029 | 4.771
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1030 | 4.777
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1031 | 4.782
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1032 | 4.788
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1033 | 4.794
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1034 | 4.799
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1035 | 4.805
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1036 | 4.81
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1037 | 4.816
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1038 | 4.822
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1039 | 4.827
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1040 | 4.833
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1041 | 4.838
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1042 | 4.844
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1043 | 4.85
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1044 | 4.855
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1045 | 4.861
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1046 | 4.866
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1047 | 4.872
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1048 | 4.878
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1049 | 4.883
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1050 | 4.889
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1051 | 4.895
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1052 | 4.9
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1053 | 4.906
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1054 | 4.911
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1055 | 4.917
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1056 | 4.923
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1057 | 4.928
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1058 | 4.934
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1059 | 4.939
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1060 | 4.945
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1061 | 4.951
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1062 | 4.956
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1063 | 4.962
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1064 | 4.968
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1065 | 4.973
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1066 | 4.979
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1067 | 4.984
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1068 | 4.99
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1069 | 4.996
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1070 | 5.001
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1071 | 5.007
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1072 | 5.013
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1073 | 5.018
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1074 | 5.024
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1075 | 5.03
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1076 | 5.035
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1077 | 5.041
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1078 | 5.046
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1079 | 5.052
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1080 | 5.058
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1081 | 5.063
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1082 | 5.069
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1083 | 5.075
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1084 | 5.08
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1085 | 5.086
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1086 | 5.092
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1087 | 5.097
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1088 | 5.103
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1089 | 5.109
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1090 | 5.114
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1091 | 5.12
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1092 | 5.125
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1093 | 5.131
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1094 | 5.137
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1095 | 5.142
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1096 | 5.148
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1097 | 5.154
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1098 | 5.159
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1099 | 5.165
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1100 | 5.171
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1101 | 5.176
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1102 | 5.182
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1103 | 5.188
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1104 | 5.193
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1105 | 5.199
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1106 | 5.205
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1107 | 5.21
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1108 | 5.216
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1109 | 5.222
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1110 | 5.227
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1111 | 5.233
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1112 | 5.239
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1113 | 5.244
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1114 | 5.25
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1115 | 5.256
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1116 | 5.261
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1117 | 5.267
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1118 | 5.273
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1119 | 5.278
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1120 | 5.284
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1121 | 5.29
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1122 | 5.295
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1123 | 5.301
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1124 | 5.307
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1125 | 5.312
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1126 | 5.318
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1127 | 5.324
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1128 | 5.33
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1129 | 5.335
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1130 | 5.341
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1131 | 5.347
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1132 | 5.352
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1133 | 5.358
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1134 | 5.364
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1135 | 5.369
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1136 | 5.375
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1137 | 5.381
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1138 | 5.386
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1139 | 5.392
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1140 | 5.398
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1141 | 5.404
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1142 | 5.409
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1143 | 5.415
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1144 | 5.421
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1145 | 5.426
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1146 | 5.432
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1147 | 5.438
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1148 | 5.443
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1149 | 5.449
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1150 | 5.455
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1151 | 5.461
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1152 | 5.466
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1153 | 5.472
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1154 | 5.478
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1155 | 5.483
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1156 | 5.489
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1157 | 5.495
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1158 | 5.501
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1159 | 5.506
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1160 | 5.512
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1161 | 5.518
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1162 | 5.523
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1163 | 5.529
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1164 | 5.535
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1165 | 5.541
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1166 | 5.546
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1167 | 5.552
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1168 | 5.558
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1169 | 5.563
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1170 | 5.569
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1171 | 5.575
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1172 | 5.581
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1173 | 5.586
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1174 | 5.592
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1175 | 5.598
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1176 | 5.604
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1177 | 5.609
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1178 | 5.615
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1179 | 5.621
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1180 | 5.627
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1181 | 5.632
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1182 | 5.638
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1183 | 5.644
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1184 | 5.649
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1185 | 5.655
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1186 | 5.661
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1187 | 5.667
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1188 | 5.672
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1189 | 5.678
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1190 | 5.684
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1191 | 5.69
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1192 | 5.695
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1193 | 5.701
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1194 | 5.707
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1195 | 5.713
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1196 | 5.718
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1197 | 5.724
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1198 | 5.73
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1199 | 5.736
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1200 | 5.741
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1201 | 5.747
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1202 | 5.753
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1203 | 5.759
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1204 | 5.764
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1205 | 5.77
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1206 | 5.776
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1207 | 5.782
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1208 | 5.788
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1209 | 5.793
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1210 | 5.799
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1211 | 5.805
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1212 | 5.811
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1213 | 5.816
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1214 | 5.822
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1215 | 5.828
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1216 | 5.834
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1217 | 5.839
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1218 | 5.845
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1219 | 5.851
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1220 | 5.857
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1221 | 5.863
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1222 | 5.868
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1223 | 5.874
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1224 | 5.88
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1225 | 5.886
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1226 | 5.891
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1227 | 5.897
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1228 | 5.903
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1229 | 5.909
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1230 | 5.915
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1231 | 5.92
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1232 | 5.926
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1233 | 5.932
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1234 | 5.938
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1235 | 5.944
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1236 | 5.949
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1237 | 5.955
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1238 | 5.961
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1239 | 5.967
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1240 | 5.972
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1241 | 5.978
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1242 | 5.984
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1243 | 5.99
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1244 | 5.996
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1245 | 6.001
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1246 | 6.007
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1247 | 6.013
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1248 | 6.019
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1249 | 6.025
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1250 | 6.03
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1251 | 6.036
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1252 | 6.042
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1253 | 6.048
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1254 | 6.054
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1255 | 6.06
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1256 | 6.065
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1257 | 6.071
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1258 | 6.077
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1259 | 6.083
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1260 | 6.089
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1261 | 6.094
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1262 | 6.1
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1263 | 6.106
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1264 | 6.112
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1265 | 6.118
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1266 | 6.124
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1267 | 6.129
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1268 | 6.135
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1269 | 6.141
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1270 | 6.147
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1271 | 6.153
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1272 | 6.158
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1273 | 6.164
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1274 | 6.17
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1275 | 6.176
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1276 | 6.182
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1277 | 6.188
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1278 | 6.193
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1279 | 6.199
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1280 | 6.205
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1281 | 6.211
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1282 | 6.217
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1283 | 6.223
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1284 | 6.228
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1285 | 6.234
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1286 | 6.24
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1287 | 6.246
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1288 | 6.252
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1289 | 6.258
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1290 | 6.264
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1291 | 6.269
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1292 | 6.275
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1293 | 6.281
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1294 | 6.287
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1295 | 6.293
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1296 | 6.299
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1297 | 6.305
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1298 | 6.31
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1299 | 6.316
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1300 | 6.322
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1301 | 6.328
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1302 | 6.334
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1303 | 6.34
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1304 | 6.346
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1305 | 6.351
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1306 | 6.357
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1307 | 6.363
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1308 | 6.369
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1309 | 6.375
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1310 | 6.381
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1311 | 6.387
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1312 | 6.392
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1313 | 6.398
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1314 | 6.404
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1315 | 6.41
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1316 | 6.416
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1317 | 6.422
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1318 | 6.428
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1319 | 6.434
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1320 | 6.439
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1321 | 6.445
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1322 | 6.451
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1323 | 6.457
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1324 | 6.463
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1325 | 6.469
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1326 | 6.475
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1327 | 6.481
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1328 | 6.486
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1329 | 6.492
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1330 | 6.498
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1331 | 6.504
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1332 | 6.51
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1333 | 6.516
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1334 | 6.522
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1335 | 6.528
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1336 | 6.534
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1337 | 6.539
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1338 | 6.545
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1339 | 6.551
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1340 | 6.557
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1341 | 6.563
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1342 | 6.569
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1343 | 6.575
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1344 | 6.581
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1345 | 6.587
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1346 | 6.593
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1347 | 6.598
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1348 | 6.604
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1349 | 6.61
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1350 | 6.616
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1351 | 6.622
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1352 | 6.628
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1353 | 6.634
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1354 | 6.64
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1355 | 6.646
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1356 | 6.652
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1357 | 6.658
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1358 | 6.664
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1359 | 6.669
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1360 | 6.675
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1361 | 6.681
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1362 | 6.687
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1363 | 6.693
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1364 | 6.699
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1365 | 6.705
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1366 | 6.711
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1367 | 6.717
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1368 | 6.723
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1369 | 6.729
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1370 | 6.735
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1371 | 6.741
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1372 | 6.746
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1373 | 6.752
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1374 | 6.758
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1375 | 6.764
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1376 | 6.77
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1377 | 6.776
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1378 | 6.782
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1379 | 6.788
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1380 | 6.794
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1381 | 6.8
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1382 | 6.806
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1383 | 6.812
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1384 | 6.818
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1385 | 6.824
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1386 | 6.83
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1387 | 6.836
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1388 | 6.842
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1389 | 6.847
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1390 | 6.853
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1391 | 6.859
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1392 | 6.865
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1393 | 6.871
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1394 | 6.877
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1395 | 6.883
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1396 | 6.889
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1397 | 6.895
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1398 | 6.901
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1399 | 6.907
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1400 | 6.913
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1401 | 6.919
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1402 | 6.925
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1403 | 6.931
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1404 | 6.937
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1405 | 6.943
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1406 | 6.949
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1407 | 6.955
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1408 | 6.961
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1409 | 6.967
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1410 | 6.973
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1411 | 6.979
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1412 | 6.985
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1413 | 6.991
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1414 | 6.997
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1415 | 7.003
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1416 | 7.008
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1417 | 7.014
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1418 | 7.02
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1419 | 7.026
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1420 | 7.032
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1421 | 7.038
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1422 | 7.044
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1423 | 7.05
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1424 | 7.056
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1425 | 7.062
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1426 | 7.068
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1427 | 7.074
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1428 | 7.08
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1429 | 7.086
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1430 | 7.092
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1431 | 7.098
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1432 | 7.104
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1433 | 7.11
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1434 | 7.116
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1435 | 7.122
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1436 | 7.128
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1437 | 7.134
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1438 | 7.14
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1439 | 7.146
-50 | #VALUE! | 0 | mV | -0.236 | Temp(F) | -58 |  | SP1 BEGINS HERE | 1440 | 7.152
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1441 | 7.158
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1442 | 7.164
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1443 | 7.17
31.778 | 5464.6598 | 3273 | 18.69 | 0 | 3214 | 0 |  |  | 1444 | 7.176
0.5933 | 199.075 | 258 | 0.593 | 0.597 | 199 | 200 |  |  | 1445 | 7.182
1.4742 | 399.05 | 458 | 1.474 | 1.478 | 399 | 400 |  |  | 1446 | 7.188
2.4642 | 599.64 | 658 | 2.461 | 2.466 | 599 | 600 |  |  | 1447 | 7.194
3.5162 | 800.04 | 859 | 3.516 | 3.521 | 800 | 801 |  |  | 1448 | 7.2
4.611 | 1000.2 | 1059 | 4.61 | 4.615 | 1000 | 1001 |  |  | 1449 | 7.206
5.7377 | 1199.34 | 1258 | 5.736 | 5.741 | 1199 | 1200 |  |  | 1450 | 7.212
0.0001 | 32.0333 | 91 | 0 | 0.003 | 32 | 33 |  | SP2 begins Here | 1451 | 7.218
0.5923 | 198.825 | 257 | 0.589 | 0.593 | 198 | 199 |  |  | 1452 | 7.224
1.4735 | 398.9 | 457 | 1.469 | 1.474 | 398 | 399 |  |  | 1453 | 7.23
2.4638 | 599.56 | 658 | 2.461 | 2.466 | 599 | 600 |  |  | 1454 | 7.236
3.5154 | 799.9 | 858 | 3.51 | 3.516 | 799 | 800 |  |  | 1455 | 7.242
4.6085 | 999.75 | 1058 | 4.604 | 4.61 | 999 | 1000 |  |  | 1456 | 7.249
5.7339 | 1198.65 | 1257 | 5.73 | 5.736 | 1198 | 1199 |  |  | 1457 | 7.255
200 | 34392.7234 | 3273 | 18.69 | 0 | 3214 | 0 | SP1 |  | 1458 | 7.261
-50 | #VALUE! | 0 | mV | -0.236 | Temp(F) | -58 | SP2 |  | 1459 | 7.267
200 | 34392.7234 | 3273 | 18.69 | 0 | 3214 | 0 | SP2 |  | 1460 | 7.273
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1461 | 7.279
4.61 | 1000 | 1058 | 4.604 | 4.61 | 999 | 1000 |  |  | 1462 | 7.285
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1463 | 7.291
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1464 | 7.297
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1465 | 7.303
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1466 | 7.309
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1467 | 7.315
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1468 | 7.321
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1469 | 7.327
450 | 77383.6276 | 3273 | 18.69 | 0 | 3214 | 0 | SP1 |  | 1470 | 7.333
450 | 77383.6276 | 3273 | 18.69 | 0 | 3214 | 0 | SP2 | and so on | 1471 | 7.339
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1472 | 7.345
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1473 | 7.351
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1474 | 7.357
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1475 | 7.363
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1476 | 7.369
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1477 | 7.375
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1478 | 7.381
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1479 | 7.387
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1480 | 7.393
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1481 | 7.399
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1482 | 7.405
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1483 | 7.411
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1484 | 7.418
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1485 | 7.424
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1486 | 7.43
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1487 | 7.436
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1488 | 7.442
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1489 | 7.448
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1490 | 7.454
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1491 | 7.46
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1492 | 7.466
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1493 | 7.472
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1494 | 7.478
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1495 | 7.484
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1496 | 7.49
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1497 | 7.496
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1498 | 7.502
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1499 | 7.508
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1500 | 7.514
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1501 | 7.521
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1502 | 7.527
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1503 | 7.533
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1504 | 7.539
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1505 | 7.545
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1506 | 7.551
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1507 | 7.557
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1508 | 7.563
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1509 | 7.569
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1510 | 7.575
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1511 | 7.581
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1512 | 7.587
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1513 | 7.593
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1514 | 7.6
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1515 | 7.606
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1516 | 7.612
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1517 | 7.618
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1518 | 7.624
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1519 | 7.63
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1520 | 7.636
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1521 | 7.642
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1522 | 7.648
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1523 | 7.654
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1524 | 7.66
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1525 | 7.667
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1526 | 7.673
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1527 | 7.679
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1528 | 7.685
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1529 | 7.691
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1530 | 7.697
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1531 | 7.703
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1532 | 7.709
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1533 | 7.715
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1534 | 7.721
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1535 | 7.728
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1536 | 7.734
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1537 | 7.74
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1538 | 7.746
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1539 | 7.752
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1540 | 7.758
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1541 | 7.764
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1542 | 7.77
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1543 | 7.776
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1544 | 7.783
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1545 | 7.789
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1546 | 7.795
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1547 | 7.801
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1548 | 7.807
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1549 | 7.813
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1550 | 7.819
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1551 | 7.825
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1552 | 7.832
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1553 | 7.838
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1554 | 7.844
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1555 | 7.85
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1556 | 7.856
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1557 | 7.862
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1558 | 7.868
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1559 | 7.874
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1560 | 7.881
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1561 | 7.887
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1562 | 7.893
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1563 | 7.899
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1564 | 7.905
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1565 | 7.911
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1566 | 7.917
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1567 | 7.923
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1568 | 7.93
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1569 | 7.936
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1570 | 7.942
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1571 | 7.948
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1572 | 7.954
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1573 | 7.96
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1574 | 7.966
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1575 | 7.973
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1576 | 7.979
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1577 | 7.985
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1578 | 7.991
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1579 | 7.997
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1580 | 8.003
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1581 | 8.01
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1582 | 8.016
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1583 | 8.022
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1584 | 8.028
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1585 | 8.034
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1586 | 8.04
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1587 | 8.047
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1588 | 8.053
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1589 | 8.059
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1590 | 8.065
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1591 | 8.071
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1592 | 8.077
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1593 | 8.083
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1594 | 8.09
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1595 | 8.096
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1596 | 8.102
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1597 | 8.108
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1598 | 8.114
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1599 | 8.121
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1600 | 8.127
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1601 | 8.133
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1602 | 8.139
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1603 | 8.145
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1604 | 8.151
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1605 | 8.158
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1606 | 8.164
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1607 | 8.17
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1608 | 8.176
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1609 | 8.182
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1610 | 8.189
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1611 | 8.195
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1612 | 8.201
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1613 | 8.207
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1614 | 8.213
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1615 | 8.219
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1616 | 8.226
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1617 | 8.232
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1618 | 8.238
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1619 | 8.244
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1620 | 8.25
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1621 | 8.257
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1622 | 8.263
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1623 | 8.269
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1624 | 8.275
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1625 | 8.281
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1626 | 8.288
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1627 | 8.294
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1628 | 8.3
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1629 | 8.306
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1630 | 8.312
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1631 | 8.319
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1632 | 8.325
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1633 | 8.331
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1634 | 8.337
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1635 | 8.343
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1636 | 8.35
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1637 | 8.356
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1638 | 8.362
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1639 | 8.368
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1640 | 8.375
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1641 | 8.381
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1642 | 8.387
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1643 | 8.393
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1644 | 8.399
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1645 | 8.406
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1646 | 8.412
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1647 | 8.418
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1648 | 8.424
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1649 | 8.431
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1650 | 8.437
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1651 | 8.443
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1652 | 8.449
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1653 | 8.455
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1654 | 8.462
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1655 | 8.468
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1656 | 8.474
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1657 | 8.48
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1658 | 8.487
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1659 | 8.493
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1660 | 8.499
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1661 | 8.505
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1662 | 8.512
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1663 | 8.518
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1664 | 8.524
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1665 | 8.53
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1666 | 8.537
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1667 | 8.543
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1668 | 8.549
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1669 | 8.555
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1670 | 8.562
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1671 | 8.568
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1672 | 8.574
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1673 | 8.58
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1674 | 8.587
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1675 | 8.593
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1676 | 8.599
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1677 | 8.605
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1678 | 8.612
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1679 | 8.618
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1680 | 8.624
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1681 | 8.63
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1682 | 8.637
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1683 | 8.643
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1684 | 8.649
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1685 | 8.655
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1686 | 8.662
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1687 | 8.668
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1688 | 8.674
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1689 | 8.68
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1690 | 8.687
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1691 | 8.693
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1692 | 8.699
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1693 | 8.706
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1694 | 8.712
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1695 | 8.718
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1696 | 8.724
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1697 | 8.731
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1698 | 8.737
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1699 | 8.743
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1700 | 8.749
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1701 | 8.756
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1702 | 8.762
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1703 | 8.768
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1704 | 8.775
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1705 | 8.781
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1706 | 8.787
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1707 | 8.793
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1708 | 8.8
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1709 | 8.806
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1710 | 8.812
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1711 | 8.819
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1712 | 8.825
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1713 | 8.831
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1714 | 8.837
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1715 | 8.844
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1716 | 8.85
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1717 | 8.856
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1718 | 8.863
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1719 | 8.869
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1720 | 8.875
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1721 | 8.882
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1722 | 8.888
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1723 | 8.894
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1724 | 8.9
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1725 | 8.907
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1726 | 8.913
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1727 | 8.919
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1728 | 8.926
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1729 | 8.932
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1730 | 8.938
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1731 | 8.945
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1732 | 8.951
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1733 | 8.957
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1734 | 8.964
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1735 | 8.97
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1736 | 8.976
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1737 | 8.983
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1738 | 8.989
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1739 | 8.995
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1740 | 9.001
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1741 | 9.008
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1742 | 9.014
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1743 | 9.02
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1744 | 9.027
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1745 | 9.033
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1746 | 9.039
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1747 | 9.046
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1748 | 9.052
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1749 | 9.058
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1750 | 9.065
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1751 | 9.071
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1752 | 9.077
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1753 | 9.084
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1754 | 9.09
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1755 | 9.096
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1756 | 9.103
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1757 | 9.109
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1758 | 9.115
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1759 | 9.122
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1760 | 9.128
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1761 | 9.134
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1762 | 9.141
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1763 | 9.147
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1764 | 9.153
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1765 | 9.16
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1766 | 9.166
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1767 | 9.172
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1768 | 9.179
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1769 | 9.185
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1770 | 9.192
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1771 | 9.198
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1772 | 9.204
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1773 | 9.211
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1774 | 9.217
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1775 | 9.223
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1776 | 9.23
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1777 | 9.236
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1778 | 9.242
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1779 | 9.249
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1780 | 9.255
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1781 | 9.261
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1782 | 9.268
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1783 | 9.274
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1784 | 9.281
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1785 | 9.287
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1786 | 9.293
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1787 | 9.3
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1788 | 9.306
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1789 | 9.312
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1790 | 9.319
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1791 | 9.325
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1792 | 9.331
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1793 | 9.338
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1794 | 9.344
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1795 | 9.351
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1796 | 9.357
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1797 | 9.363
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1798 | 9.37
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1799 | 9.376
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1800 | 9.382
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1801 | 9.389
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1802 | 9.395
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1803 | 9.402
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1804 | 9.408
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1805 | 9.414
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1806 | 9.421
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1807 | 9.427
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1808 | 9.434
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1809 | 9.44
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1810 | 9.446
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1811 | 9.453
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1812 | 9.459
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1813 | 9.465
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1814 | 9.472
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1815 | 9.478
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1816 | 9.485
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1817 | 9.491
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1818 | 9.497
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1819 | 9.504
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1820 | 9.51
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1821 | 9.517
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1822 | 9.523
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1823 | 9.529
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1824 | 9.536
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1825 | 9.542
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1826 | 9.549
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1827 | 9.555
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1828 | 9.561
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1829 | 9.568
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1830 | 9.574
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1831 | 9.581
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1832 | 9.587
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1833 | 9.594
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1834 | 9.6
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1835 | 9.606
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1836 | 9.613
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1837 | 9.619
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1838 | 9.626
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1839 | 9.632
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1840 | 9.638
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1841 | 9.645
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1842 | 9.651
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1843 | 9.658
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1844 | 9.664
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1845 | 9.671
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1846 | 9.677
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1847 | 9.683
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1848 | 9.69
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1849 | 9.696
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1850 | 9.703
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1851 | 9.709
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1852 | 9.716
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1853 | 9.722
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1854 | 9.728
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1855 | 9.735
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1856 | 9.741
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1857 | 9.748
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1858 | 9.754
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1859 | 9.761
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1860 | 9.767
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1861 | 9.773
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1862 | 9.78
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1863 | 9.786
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1864 | 9.793
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1865 | 9.799
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1866 | 9.806
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1867 | 9.812
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1868 | 9.819
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1869 | 9.825
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1870 | 9.831
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1871 | 9.838
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1872 | 9.844
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1873 | 9.851
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1874 | 9.857
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1875 | 9.864
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1876 | 9.87
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1877 | 9.877
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1878 | 9.883
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1879 | 9.889
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1880 | 9.896
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1881 | 9.902
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1882 | 9.909
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1883 | 9.915
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1884 | 9.922
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1885 | 9.928
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1886 | 9.935
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1887 | 9.941
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1888 | 9.948
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1889 | 9.954
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1890 | 9.961
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1891 | 9.967
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1892 | 9.973
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1893 | 9.98
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1894 | 9.986
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1895 | 9.993
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1896 | 9.999
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1897 | 10.006
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1898 | 10.012
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1899 | 10.019
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1900 | 10.025
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1901 | 10.032
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1902 | 10.038
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1903 | 10.045
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1904 | 10.051
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1905 | 10.058
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1906 | 10.064
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1907 | 10.071
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1908 | 10.077
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1909 | 10.084
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1910 | 10.09
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1911 | 10.097
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1912 | 10.103
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1913 | 10.11
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1914 | 10.116
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1915 | 10.123
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1916 | 10.129
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1917 | 10.136
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1918 | 10.142
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1919 | 10.149
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1920 | 10.155
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1921 | 10.162
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1922 | 10.168
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1923 | 10.175
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1924 | 10.181
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1925 | 10.188
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1926 | 10.194
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1927 | 10.201
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1928 | 10.207
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1929 | 10.214
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1930 | 10.22
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1931 | 10.227
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1932 | 10.233
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1933 | 10.24
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1934 | 10.246
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1935 | 10.253
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1936 | 10.259
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1937 | 10.266
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1938 | 10.272
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1939 | 10.279
0 | 32 | 90 | -0.003 | 0 | 31 | 32 |  |  | 1940 | 10.285
[9 empty] | 1941 | 10.292
[9 empty] | 1942 | 10.298
[9 empty] | 1943 | 10.305
[9 empty] | 1944 | 10.311
[9 empty] | 1945 | 10.318
[9 empty] | 1946 | 10.324
[9 empty] | 1947 | 10.331
[9 empty] | 1948 | 10.337
[9 empty] | 1949 | 10.344
[9 empty] | 1950 | 10.35
[9 empty] | 1951 | 10.357
[9 empty] | 1952 | 10.363
[9 empty] | 1953 | 10.37
[9 empty] | 1954 | 10.376
[9 empty] | 1955 | 10.383
[9 empty] | 1956 | 10.39
[9 empty] | 1957 | 10.396
[9 empty] | 1958 | 10.403
[9 empty] | 1959 | 10.409
[9 empty] | 1960 | 10.416
[9 empty] | 1961 | 10.422
[9 empty] | 1962 | 10.429
[9 empty] | 1963 | 10.435
[9 empty] | 1964 | 10.442
[9 empty] | 1965 | 10.448
[9 empty] | 1966 | 10.455
[9 empty] | 1967 | 10.461
[9 empty] | 1968 | 10.468
[9 empty] | 1969 | 10.475
[9 empty] | 1970 | 10.481
[9 empty] | 1971 | 10.488
[9 empty] | 1972 | 10.494
[9 empty] | 1973 | 10.501
[9 empty] | 1974 | 10.507
[9 empty] | 1975 | 10.514
[9 empty] | 1976 | 10.52
[9 empty] | 1977 | 10.527
[9 empty] | 1978 | 10.533
[9 empty] | 1979 | 10.54
[9 empty] | 1980 | 10.547
[9 empty] | 1981 | 10.553
[9 empty] | 1982 | 10.56
[9 empty] | 1983 | 10.566
[9 empty] | 1984 | 10.573
[9 empty] | 1985 | 10.579
[9 empty] | 1986 | 10.586
[9 empty] | 1987 | 10.592
[9 empty] | 1988 | 10.599
[9 empty] | 1989 | 10.606
[9 empty] | 1990 | 10.612
[9 empty] | 1991 | 10.619
[9 empty] | 1992 | 10.625
[9 empty] | 1993 | 10.632
[9 empty] | 1994 | 10.638
[9 empty] | 1995 | 10.645
[9 empty] | 1996 | 10.651
[9 empty] | 1997 | 10.658
[9 empty] | 1998 | 10.665
[9 empty] | 1999 | 10.671
[9 empty] | 2000 | 10.678
[9 empty] | 2001 | 10.684
[9 empty] | 2002 | 10.691
[9 empty] | 2003 | 10.697
[9 empty] | 2004 | 10.704
[9 empty] | 2005 | 10.711
[9 empty] | 2006 | 10.717
[9 empty] | 2007 | 10.724
[9 empty] | 2008 | 10.73
[9 empty] | 2009 | 10.737
[9 empty] | 2010 | 10.743
[9 empty] | 2011 | 10.75
[9 empty] | 2012 | 10.757
[9 empty] | 2013 | 10.763
[9 empty] | 2014 | 10.77
[9 empty] | 2015 | 10.776
[9 empty] | 2016 | 10.783
[9 empty] | 2017 | 10.789
[9 empty] | 2018 | 10.796
[9 empty] | 2019 | 10.803
[9 empty] | 2020 | 10.809
[9 empty] | 2021 | 10.816
[9 empty] | 2022 | 10.822
[9 empty] | 2023 | 10.829
[9 empty] | 2024 | 10.836
[9 empty] | 2025 | 10.842
[9 empty] | 2026 | 10.849
[9 empty] | 2027 | 10.855
[9 empty] | 2028 | 10.862
[9 empty] | 2029 | 10.868
[9 empty] | 2030 | 10.875
[9 empty] | 2031 | 10.882
[9 empty] | 2032 | 10.888
[9 empty] | 2033 | 10.895
[9 empty] | 2034 | 10.901
[9 empty] | 2035 | 10.908
[9 empty] | 2036 | 10.915
[9 empty] | 2037 | 10.921
[9 empty] | 2038 | 10.928
[9 empty] | 2039 | 10.934
[9 empty] | 2040 | 10.941
[9 empty] | 2041 | 10.948
[9 empty] | 2042 | 10.954
[9 empty] | 2043 | 10.961
[9 empty] | 2044 | 10.967
[9 empty] | 2045 | 10.974
[9 empty] | 2046 | 10.981
[9 empty] | 2047 | 10.987
[9 empty] | 2048 | 10.994
[9 empty] | 2049 | 11
[9 empty] | 2050 | 11.007
[9 empty] | 2051 | 11.014
[9 empty] | 2052 | 11.02
[9 empty] | 2053 | 11.027
[9 empty] | 2054 | 11.033
[9 empty] | 2055 | 11.04
[9 empty] | 2056 | 11.047
[9 empty] | 2057 | 11.053
[9 empty] | 2058 | 11.06
[9 empty] | 2059 | 11.066
[9 empty] | 2060 | 11.073
[9 empty] | 2061 | 11.08
[9 empty] | 2062 | 11.086
[9 empty] | 2063 | 11.093
[9 empty] | 2064 | 11.099
[9 empty] | 2065 | 11.106
[9 empty] | 2066 | 11.113
[9 empty] | 2067 | 11.119
[9 empty] | 2068 | 11.126
[9 empty] | 2069 | 11.132
[9 empty] | 2070 | 11.139
[9 empty] | 2071 | 11.146
[9 empty] | 2072 | 11.152
[9 empty] | 2073 | 11.159
[9 empty] | 2074 | 11.166
[9 empty] | 2075 | 11.172
[9 empty] | 2076 | 11.179
[9 empty] | 2077 | 11.185
[9 empty] | 2078 | 11.192
[9 empty] | 2079 | 11.199
[9 empty] | 2080 | 11.205
[9 empty] | 2081 | 11.212
[9 empty] | 2082 | 11.219
[9 empty] | 2083 | 11.225
[9 empty] | 2084 | 11.232
[9 empty] | 2085 | 11.238
[9 empty] | 2086 | 11.245
[9 empty] | 2087 | 11.252
[9 empty] | 2088 | 11.258
[9 empty] | 2089 | 11.265
[9 empty] | 2090 | 11.272
[9 empty] | 2091 | 11.278
[9 empty] | 2092 | 11.285
[9 empty] | 2093 | 11.291
[9 empty] | 2094 | 11.298
[9 empty] | 2095 | 11.305
[9 empty] | 2096 | 11.311
[9 empty] | 2097 | 11.318
[9 empty] | 2098 | 11.325
[9 empty] | 2099 | 11.331
[9 empty] | 2100 | 11.338
[9 empty] | 2101 | 11.345
[9 empty] | 2102 | 11.351
[9 empty] | 2103 | 11.358
[9 empty] | 2104 | 11.364
[9 empty] | 2105 | 11.371
[9 empty] | 2106 | 11.378
[9 empty] | 2107 | 11.384
[9 empty] | 2108 | 11.391
[9 empty] | 2109 | 11.398
[9 empty] | 2110 | 11.404
[9 empty] | 2111 | 11.411
[9 empty] | 2112 | 11.418
[9 empty] | 2113 | 11.424
[9 empty] | 2114 | 11.431
[9 empty] | 2115 | 11.437
[9 empty] | 2116 | 11.444
[9 empty] | 2117 | 11.451
[9 empty] | 2118 | 11.457
[9 empty] | 2119 | 11.464
[9 empty] | 2120 | 11.471
[9 empty] | 2121 | 11.477
[9 empty] | 2122 | 11.484
[9 empty] | 2123 | 11.491
[9 empty] | 2124 | 11.497
[9 empty] | 2125 | 11.504
[9 empty] | 2126 | 11.511
[9 empty] | 2127 | 11.517
[9 empty] | 2128 | 11.524
[9 empty] | 2129 | 11.531
[9 empty] | 2130 | 11.537
[9 empty] | 2131 | 11.544
[9 empty] | 2132 | 11.55
[9 empty] | 2133 | 11.557
[9 empty] | 2134 | 11.564
[9 empty] | 2135 | 11.57
[9 empty] | 2136 | 11.577
[9 empty] | 2137 | 11.584
[9 empty] | 2138 | 11.59
[9 empty] | 2139 | 11.597
[9 empty] | 2140 | 11.604
[9 empty] | 2141 | 11.61
[9 empty] | 2142 | 11.617
[9 empty] | 2143 | 11.624
[9 empty] | 2144 | 11.63
[9 empty] | 2145 | 11.637
[9 empty] | 2146 | 11.644
[9 empty] | 2147 | 11.65
[9 empty] | 2148 | 11.657
[9 empty] | 2149 | 11.664
[9 empty] | 2150 | 11.67
[9 empty] | 2151 | 11.677
[9 empty] | 2152 | 11.684
[9 empty] | 2153 | 11.69
[9 empty] | 2154 | 11.697
[9 empty] | 2155 | 11.704
[9 empty] | 2156 | 11.71
[9 empty] | 2157 | 11.717
[9 empty] | 2158 | 11.724
[9 empty] | 2159 | 11.73
[9 empty] | 2160 | 11.737
[9 empty] | 2161 | 11.744
[9 empty] | 2162 | 11.75
[9 empty] | 2163 | 11.757
[9 empty] | 2164 | 11.764
[9 empty] | 2165 | 11.77
[9 empty] | 2166 | 11.777
[9 empty] | 2167 | 11.784
[9 empty] | 2168 | 11.79
[9 empty] | 2169 | 11.797
[9 empty] | 2170 | 11.804
[9 empty] | 2171 | 11.81
[9 empty] | 2172 | 11.817
[9 empty] | 2173 | 11.824
[9 empty] | 2174 | 11.83
[9 empty] | 2175 | 11.837
[9 empty] | 2176 | 11.844
[9 empty] | 2177 | 11.85
[9 empty] | 2178 | 11.857
[9 empty] | 2179 | 11.864
[9 empty] | 2180 | 11.87
[9 empty] | 2181 | 11.877
[9 empty] | 2182 | 11.884
[9 empty] | 2183 | 11.89
[9 empty] | 2184 | 11.897
[9 empty] | 2185 | 11.904
[9 empty] | 2186 | 11.91
[9 empty] | 2187 | 11.917
[9 empty] | 2188 | 11.924
[9 empty] | 2189 | 11.931
[9 empty] | 2190 | 11.937
[9 empty] | 2191 | 11.944
[9 empty] | 2192 | 11.951
[9 empty] | 2193 | 11.957
[9 empty] | 2194 | 11.964
[9 empty] | 2195 | 11.971
[9 empty] | 2196 | 11.977
[9 empty] | 2197 | 11.984
[9 empty] | 2198 | 11.991
[9 empty] | 2199 | 11.997
[9 empty] | 2200 | 12.004
[9 empty] | 2201 | 12.011
[9 empty] | 2202 | 12.017
[9 empty] | 2203 | 12.024
[9 empty] | 2204 | 12.031
[9 empty] | 2205 | 12.037
[9 empty] | 2206 | 12.044
[9 empty] | 2207 | 12.051
[9 empty] | 2208 | 12.058
[9 empty] | 2209 | 12.064
[9 empty] | 2210 | 12.071
[9 empty] | 2211 | 12.078
[9 empty] | 2212 | 12.084
[9 empty] | 2213 | 12.091
[9 empty] | 2214 | 12.098
[9 empty] | 2215 | 12.104
[9 empty] | 2216 | 12.111
[9 empty] | 2217 | 12.118
[9 empty] | 2218 | 12.124
[9 empty] | 2219 | 12.131
[9 empty] | 2220 | 12.138
[9 empty] | 2221 | 12.145
[9 empty] | 2222 | 12.151
[9 empty] | 2223 | 12.158
[9 empty] | 2224 | 12.165
[9 empty] | 2225 | 12.171
[9 empty] | 2226 | 12.178
[9 empty] | 2227 | 12.185
[9 empty] | 2228 | 12.191
[9 empty] | 2229 | 12.198
[9 empty] | 2230 | 12.205
[9 empty] | 2231 | 12.211
[9 empty] | 2232 | 12.218
[9 empty] | 2233 | 12.225
[9 empty] | 2234 | 12.232
[9 empty] | 2235 | 12.238
[9 empty] | 2236 | 12.245
[9 empty] | 2237 | 12.252
[9 empty] | 2238 | 12.258
[9 empty] | 2239 | 12.265
[9 empty] | 2240 | 12.272
[9 empty] | 2241 | 12.278
[9 empty] | 2242 | 12.285
[9 empty] | 2243 | 12.292
[9 empty] | 2244 | 12.299
[9 empty] | 2245 | 12.305
[9 empty] | 2246 | 12.312
[9 empty] | 2247 | 12.319
[9 empty] | 2248 | 12.325
[9 empty] | 2249 | 12.332
[9 empty] | 2250 | 12.339
[9 empty] | 2251 | 12.346
[9 empty] | 2252 | 12.352
[9 empty] | 2253 | 12.359
[9 empty] | 2254 | 12.366
[9 empty] | 2255 | 12.372
[9 empty] | 2256 | 12.379
[9 empty] | 2257 | 12.386
[9 empty] | 2258 | 12.392
[9 empty] | 2259 | 12.399
[9 empty] | 2260 | 12.406
[9 empty] | 2261 | 12.413
[9 empty] | 2262 | 12.419
[9 empty] | 2263 | 12.426
[9 empty] | 2264 | 12.433
[9 empty] | 2265 | 12.439
[9 empty] | 2266 | 12.446
[9 empty] | 2267 | 12.453
[9 empty] | 2268 | 12.46
[9 empty] | 2269 | 12.466
[9 empty] | 2270 | 12.473
[9 empty] | 2271 | 12.48
[9 empty] | 2272 | 12.486
[9 empty] | 2273 | 12.493
[9 empty] | 2274 | 12.5
[9 empty] | 2275 | 12.507
[9 empty] | 2276 | 12.513
[9 empty] | 2277 | 12.52
[9 empty] | 2278 | 12.527
[9 empty] | 2279 | 12.533
[9 empty] | 2280 | 12.54
[9 empty] | 2281 | 12.547
[9 empty] | 2282 | 12.554
[9 empty] | 2283 | 12.56
[9 empty] | 2284 | 12.567
[9 empty] | 2285 | 12.574
[9 empty] | 2286 | 12.58
[9 empty] | 2287 | 12.587
[9 empty] | 2288 | 12.594
[9 empty] | 2289 | 12.601
[9 empty] | 2290 | 12.607
[9 empty] | 2291 | 12.614
[9 empty] | 2292 | 12.621
[9 empty] | 2293 | 12.627
[9 empty] | 2294 | 12.634
[9 empty] | 2295 | 12.641
[9 empty] | 2296 | 12.648
[9 empty] | 2297 | 12.654
[9 empty] | 2298 | 12.661
[9 empty] | 2299 | 12.668
[9 empty] | 2300 | 12.675
[9 empty] | 2301 | 12.681
[9 empty] | 2302 | 12.688
[9 empty] | 2303 | 12.695
[9 empty] | 2304 | 12.701
[9 empty] | 2305 | 12.708
[9 empty] | 2306 | 12.715
[9 empty] | 2307 | 12.722
[9 empty] | 2308 | 12.728
[9 empty] | 2309 | 12.735
[9 empty] | 2310 | 12.742
[9 empty] | 2311 | 12.748
[9 empty] | 2312 | 12.755
[9 empty] | 2313 | 12.762
[9 empty] | 2314 | 12.769
[9 empty] | 2315 | 12.775
[9 empty] | 2316 | 12.782
[9 empty] | 2317 | 12.789
[9 empty] | 2318 | 12.796
[9 empty] | 2319 | 12.802
[9 empty] | 2320 | 12.809
[9 empty] | 2321 | 12.816
[9 empty] | 2322 | 12.822
[9 empty] | 2323 | 12.829
[9 empty] | 2324 | 12.836
[9 empty] | 2325 | 12.843
[9 empty] | 2326 | 12.849
[9 empty] | 2327 | 12.856
[9 empty] | 2328 | 12.863
[9 empty] | 2329 | 12.87
[9 empty] | 2330 | 12.876
[9 empty] | 2331 | 12.883
[9 empty] | 2332 | 12.89
[9 empty] | 2333 | 12.896
[9 empty] | 2334 | 12.903
[9 empty] | 2335 | 12.91
[9 empty] | 2336 | 12.917
[9 empty] | 2337 | 12.923
[9 empty] | 2338 | 12.93
[9 empty] | 2339 | 12.937
[9 empty] | 2340 | 12.944
[9 empty] | 2341 | 12.95
[9 empty] | 2342 | 12.957
[9 empty] | 2343 | 12.964
[9 empty] | 2344 | 12.971
[9 empty] | 2345 | 12.977
[9 empty] | 2346 | 12.984
[9 empty] | 2347 | 12.991
[9 empty] | 2348 | 12.997
[9 empty] | 2349 | 13.004
[9 empty] | 2350 | 13.011
[9 empty] | 2351 | 13.018
[9 empty] | 2352 | 13.024
[9 empty] | 2353 | 13.031
[9 empty] | 2354 | 13.038
[9 empty] | 2355 | 13.045
[9 empty] | 2356 | 13.051
[9 empty] | 2357 | 13.058
[9 empty] | 2358 | 13.065
[9 empty] | 2359 | 13.072
[9 empty] | 2360 | 13.078
[9 empty] | 2361 | 13.085
[9 empty] | 2362 | 13.092
[9 empty] | 2363 | 13.098
[9 empty] | 2364 | 13.105
[9 empty] | 2365 | 13.112
[9 empty] | 2366 | 13.119
[9 empty] | 2367 | 13.125
[9 empty] | 2368 | 13.132
[9 empty] | 2369 | 13.139
[9 empty] | 2370 | 13.146
[9 empty] | 2371 | 13.152
[9 empty] | 2372 | 13.159
[9 empty] | 2373 | 13.166
[9 empty] | 2374 | 13.173
[9 empty] | 2375 | 13.179
[9 empty] | 2376 | 13.186
[9 empty] | 2377 | 13.193
[9 empty] | 2378 | 13.199
[9 empty] | 2379 | 13.206
[9 empty] | 2380 | 13.213
[9 empty] | 2381 | 13.22
[9 empty] | 2382 | 13.226
[9 empty] | 2383 | 13.233
[9 empty] | 2384 | 13.24
[9 empty] | 2385 | 13.247
[9 empty] | 2386 | 13.253
[9 empty] | 2387 | 13.26
[9 empty] | 2388 | 13.267
[9 empty] | 2389 | 13.274
[9 empty] | 2390 | 13.28
[9 empty] | 2391 | 13.287
[9 empty] | 2392 | 13.294
[9 empty] | 2393 | 13.301
[9 empty] | 2394 | 13.307
[9 empty] | 2395 | 13.314
[9 empty] | 2396 | 13.321
[9 empty] | 2397 | 13.328
[9 empty] | 2398 | 13.334
[9 empty] | 2399 | 13.341
[9 empty] | 2400 | 13.348
[9 empty] | 2401 | 13.354
[9 empty] | 2402 | 13.361
[9 empty] | 2403 | 13.368
[9 empty] | 2404 | 13.375
[9 empty] | 2405 | 13.381
[9 empty] | 2406 | 13.388
[9 empty] | 2407 | 13.395
[9 empty] | 2408 | 13.402
[9 empty] | 2409 | 13.408
[9 empty] | 2410 | 13.415
[9 empty] | 2411 | 13.422
[9 empty] | 2412 | 13.429
[9 empty] | 2413 | 13.435
[9 empty] | 2414 | 13.442
[9 empty] | 2415 | 13.449
[9 empty] | 2416 | 13.456
[9 empty] | 2417 | 13.462
[9 empty] | 2418 | 13.469
[9 empty] | 2419 | 13.476
[9 empty] | 2420 | 13.483
[9 empty] | 2421 | 13.489
[9 empty] | 2422 | 13.496
[9 empty] | 2423 | 13.503
[9 empty] | 2424 | 13.51
[9 empty] | 2425 | 13.516
[9 empty] | 2426 | 13.523
[9 empty] | 2427 | 13.53
[9 empty] | 2428 | 13.537
[9 empty] | 2429 | 13.543
[9 empty] | 2430 | 13.55
[9 empty] | 2431 | 13.557
[9 empty] | 2432 | 13.563
[9 empty] | 2433 | 13.57
[9 empty] | 2434 | 13.577
[9 empty] | 2435 | 13.584
[9 empty] | 2436 | 13.59
[9 empty] | 2437 | 13.597
[9 empty] | 2438 | 13.604
[9 empty] | 2439 | 13.611
[9 empty] | 2440 | 13.617
[9 empty] | 2441 | 13.624
[9 empty] | 2442 | 13.631
[9 empty] | 2443 | 13.638
[9 empty] | 2444 | 13.644
[9 empty] | 2445 | 13.651
[9 empty] | 2446 | 13.658
[9 empty] | 2447 | 13.665
[9 empty] | 2448 | 13.671
[9 empty] | 2449 | 13.678
[9 empty] | 2450 | 13.685
[9 empty] | 2451 | 13.692
[9 empty] | 2452 | 13.698
[9 empty] | 2453 | 13.705
[9 empty] | 2454 | 13.712
[9 empty] | 2455 | 13.719
[9 empty] | 2456 | 13.725
[9 empty] | 2457 | 13.732
[9 empty] | 2458 | 13.739
[9 empty] | 2459 | 13.746
[9 empty] | 2460 | 13.752
[9 empty] | 2461 | 13.759
[9 empty] | 2462 | 13.766
[9 empty] | 2463 | 13.773
[9 empty] | 2464 | 13.779
[9 empty] | 2465 | 13.786
[9 empty] | 2466 | 13.793
[9 empty] | 2467 | 13.8
[9 empty] | 2468 | 13.806
[9 empty] | 2469 | 13.813
[9 empty] | 2470 | 13.82
[9 empty] | 2471 | 13.826
[9 empty] | 2472 | 13.833
[9 empty] | 2473 | 13.84
[9 empty] | 2474 | 13.847
[9 empty] | 2475 | 13.853
[9 empty] | 2476 | 13.86
[9 empty] | 2477 | 13.867
[9 empty] | 2478 | 13.874
[9 empty] | 2479 | 13.88
[9 empty] | 2480 | 13.887
[9 empty] | 2481 | 13.894
[9 empty] | 2482 | 13.901
[9 empty] | 2483 | 13.907
[9 empty] | 2484 | 13.914
[9 empty] | 2485 | 13.921
[9 empty] | 2486 | 13.928
[9 empty] | 2487 | 13.934
[9 empty] | 2488 | 13.941
[9 empty] | 2489 | 13.948
[9 empty] | 2490 | 13.955
[9 empty] | 2491 | 13.961
[9 empty] | 2492 | 13.968
[9 empty] | 2493 | 13.975
[9 empty] | 2494 | 13.982
[9 empty] | 2495 | 13.988
[9 empty] | 2496 | 13.995
[9 empty] | 2497 | 14.002
[9 empty] | 2498 | 14.009
[9 empty] | 2499 | 14.015
[9 empty] | 2500 | 14.022
[9 empty] | 2501 | 14.029
[9 empty] | 2502 | 14.036
[9 empty] | 2503 | 14.042
[9 empty] | 2504 | 14.049
[9 empty] | 2505 | 14.056
[9 empty] | 2506 | 14.063
[9 empty] | 2507 | 14.069
[9 empty] | 2508 | 14.076
[9 empty] | 2509 | 14.083
[9 empty] | 2510 | 14.089
[9 empty] | 2511 | 14.096
[9 empty] | 2512 | 14.103
[9 empty] | 2513 | 14.11
[9 empty] | 2514 | 14.116
[9 empty] | 2515 | 14.123
[9 empty] | 2516 | 14.13
[9 empty] | 2517 | 14.137
[9 empty] | 2518 | 14.143
[9 empty] | 2519 | 14.15
[9 empty] | 2520 | 14.157
[9 empty] | 2521 | 14.164
[9 empty] | 2522 | 14.17
[9 empty] | 2523 | 14.177
[9 empty] | 2524 | 14.184
[9 empty] | 2525 | 14.191
[9 empty] | 2526 | 14.197
[9 empty] | 2527 | 14.204
[9 empty] | 2528 | 14.211
[9 empty] | 2529 | 14.218
[9 empty] | 2530 | 14.224
[9 empty] | 2531 | 14.231
[9 empty] | 2532 | 14.238
[9 empty] | 2533 | 14.245
[9 empty] | 2534 | 14.251
[9 empty] | 2535 | 14.258
[9 empty] | 2536 | 14.265
[9 empty] | 2537 | 14.272
[9 empty] | 2538 | 14.278
[9 empty] | 2539 | 14.285
[9 empty] | 2540 | 14.292
[9 empty] | 2541 | 14.298
[9 empty] | 2542 | 14.305
[9 empty] | 2543 | 14.312
[9 empty] | 2544 | 14.319
[9 empty] | 2545 | 14.325
[9 empty] | 2546 | 14.332
[9 empty] | 2547 | 14.339
[9 empty] | 2548 | 14.346
[9 empty] | 2549 | 14.352
[9 empty] | 2550 | 14.359
[9 empty] | 2551 | 14.366
[9 empty] | 2552 | 14.373
[9 empty] | 2553 | 14.379
[9 empty] | 2554 | 14.386
[9 empty] | 2555 | 14.393
[9 empty] | 2556 | 14.4
[9 empty] | 2557 | 14.406
[9 empty] | 2558 | 14.413
[9 empty] | 2559 | 14.42
[9 empty] | 2560 | 14.426
[9 empty] | 2561 | 14.433
[9 empty] | 2562 | 14.44
[9 empty] | 2563 | 14.447
[9 empty] | 2564 | 14.453
[9 empty] | 2565 | 14.46
[9 empty] | 2566 | 14.467
[9 empty] | 2567 | 14.474
[9 empty] | 2568 | 14.48
[9 empty] | 2569 | 14.487
[9 empty] | 2570 | 14.494
[9 empty] | 2571 | 14.501
[9 empty] | 2572 | 14.507
[9 empty] | 2573 | 14.514
[9 empty] | 2574 | 14.521
[9 empty] | 2575 | 14.528
[9 empty] | 2576 | 14.534
[9 empty] | 2577 | 14.541
[9 empty] | 2578 | 14.548
[9 empty] | 2579 | 14.554
[9 empty] | 2580 | 14.561
[9 empty] | 2581 | 14.568
[9 empty] | 2582 | 14.575
[9 empty] | 2583 | 14.581
[9 empty] | 2584 | 14.588
[9 empty] | 2585 | 14.595
[9 empty] | 2586 | 14.602
[9 empty] | 2587 | 14.608
[9 empty] | 2588 | 14.615
[9 empty] | 2589 | 14.622
[9 empty] | 2590 | 14.629
[9 empty] | 2591 | 14.635
[9 empty] | 2592 | 14.642
[9 empty] | 2593 | 14.649
[9 empty] | 2594 | 14.655
[9 empty] | 2595 | 14.662
[9 empty] | 2596 | 14.669
[9 empty] | 2597 | 14.676
[9 empty] | 2598 | 14.682
[9 empty] | 2599 | 14.689
[9 empty] | 2600 | 14.696
[9 empty] | 2601 | 14.703
[9 empty] | 2602 | 14.709
[9 empty] | 2603 | 14.716
[9 empty] | 2604 | 14.723
[9 empty] | 2605 | 14.729
[9 empty] | 2606 | 14.736
[9 empty] | 2607 | 14.743
[9 empty] | 2608 | 14.75
[9 empty] | 2609 | 14.756
[9 empty] | 2610 | 14.763
[9 empty] | 2611 | 14.77
[9 empty] | 2612 | 14.777
[9 empty] | 2613 | 14.783
[9 empty] | 2614 | 14.79
[9 empty] | 2615 | 14.797
[9 empty] | 2616 | 14.803
[9 empty] | 2617 | 14.81
[9 empty] | 2618 | 14.817
[9 empty] | 2619 | 14.824
[9 empty] | 2620 | 14.83
[9 empty] | 2621 | 14.837
[9 empty] | 2622 | 14.844
[9 empty] | 2623 | 14.851
[9 empty] | 2624 | 14.857
[9 empty] | 2625 | 14.864
[9 empty] | 2626 | 14.871
[9 empty] | 2627 | 14.877
[9 empty] | 2628 | 14.884
[9 empty] | 2629 | 14.891
[9 empty] | 2630 | 14.898
[9 empty] | 2631 | 14.904
[9 empty] | 2632 | 14.911
[9 empty] | 2633 | 14.918
[9 empty] | 2634 | 14.925
[9 empty] | 2635 | 14.931
[9 empty] | 2636 | 14.938
[9 empty] | 2637 | 14.945
[9 empty] | 2638 | 14.951
[9 empty] | 2639 | 14.958
[9 empty] | 2640 | 14.965
[9 empty] | 2641 | 14.972
[9 empty] | 2642 | 14.978
[9 empty] | 2643 | 14.985
[9 empty] | 2644 | 14.992
[9 empty] | 2645 | 14.998
[9 empty] | 2646 | 15.005
[9 empty] | 2647 | 15.012
[9 empty] | 2648 | 15.019
[9 empty] | 2649 | 15.025
[9 empty] | 2650 | 15.032
[9 empty] | 2651 | 15.039
[9 empty] | 2652 | 15.045
[9 empty] | 2653 | 15.052
[9 empty] | 2654 | 15.059
[9 empty] | 2655 | 15.066
[9 empty] | 2656 | 15.072
[9 empty] | 2657 | 15.079
[9 empty] | 2658 | 15.086
[9 empty] | 2659 | 15.092
[9 empty] | 2660 | 15.099
[9 empty] | 2661 | 15.106
[9 empty] | 2662 | 15.113
[9 empty] | 2663 | 15.119
[9 empty] | 2664 | 15.126
[9 empty] | 2665 | 15.133
[9 empty] | 2666 | 15.139
[9 empty] | 2667 | 15.146
[9 empty] | 2668 | 15.153
[9 empty] | 2669 | 15.16
[9 empty] | 2670 | 15.166
[9 empty] | 2671 | 15.173
[9 empty] | 2672 | 15.18
[9 empty] | 2673 | 15.186
[9 empty] | 2674 | 15.193
[9 empty] | 2675 | 15.2
[9 empty] | 2676 | 15.207
[9 empty] | 2677 | 15.213
[9 empty] | 2678 | 15.22
[9 empty] | 2679 | 15.227
[9 empty] | 2680 | 15.233
[9 empty] | 2681 | 15.24
[9 empty] | 2682 | 15.247
[9 empty] | 2683 | 15.254
[9 empty] | 2684 | 15.26
[9 empty] | 2685 | 15.267
[9 empty] | 2686 | 15.274
[9 empty] | 2687 | 15.28
[9 empty] | 2688 | 15.287
[9 empty] | 2689 | 15.294
[9 empty] | 2690 | 15.3
[9 empty] | 2691 | 15.307
[9 empty] | 2692 | 15.314
[9 empty] | 2693 | 15.321
[9 empty] | 2694 | 15.327
[9 empty] | 2695 | 15.334
[9 empty] | 2696 | 15.341
[9 empty] | 2697 | 15.347
[9 empty] | 2698 | 15.354
[9 empty] | 2699 | 15.361
[9 empty] | 2700 | 15.367
[9 empty] | 2701 | 15.374
[9 empty] | 2702 | 15.381
[9 empty] | 2703 | 15.388
[9 empty] | 2704 | 15.394
[9 empty] | 2705 | 15.401
[9 empty] | 2706 | 15.408
[9 empty] | 2707 | 15.414
[9 empty] | 2708 | 15.421
[9 empty] | 2709 | 15.428
[9 empty] | 2710 | 15.434
[9 empty] | 2711 | 15.441
[9 empty] | 2712 | 15.448
[9 empty] | 2713 | 15.455
[9 empty] | 2714 | 15.461
[9 empty] | 2715 | 15.468
[9 empty] | 2716 | 15.475
[9 empty] | 2717 | 15.481
[9 empty] | 2718 | 15.488
[9 empty] | 2719 | 15.495
[9 empty] | 2720 | 15.501
[9 empty] | 2721 | 15.508
[9 empty] | 2722 | 15.515
[9 empty] | 2723 | 15.521
[9 empty] | 2724 | 15.528
[9 empty] | 2725 | 15.535
[9 empty] | 2726 | 15.542
[9 empty] | 2727 | 15.548
[9 empty] | 2728 | 15.555
[9 empty] | 2729 | 15.562
[9 empty] | 2730 | 15.568
[9 empty] | 2731 | 15.575
[9 empty] | 2732 | 15.582
[9 empty] | 2733 | 15.588
[9 empty] | 2734 | 15.595
[9 empty] | 2735 | 15.602
[9 empty] | 2736 | 15.608
[9 empty] | 2737 | 15.615
[9 empty] | 2738 | 15.622
[9 empty] | 2739 | 15.628
[9 empty] | 2740 | 15.635
[9 empty] | 2741 | 15.642
[9 empty] | 2742 | 15.649
[9 empty] | 2743 | 15.655
[9 empty] | 2744 | 15.662
[9 empty] | 2745 | 15.669
[9 empty] | 2746 | 15.675
[9 empty] | 2747 | 15.682
[9 empty] | 2748 | 15.689
[9 empty] | 2749 | 15.695
[9 empty] | 2750 | 15.702
[9 empty] | 2751 | 15.709
[9 empty] | 2752 | 15.715
[9 empty] | 2753 | 15.722
[9 empty] | 2754 | 15.729
[9 empty] | 2755 | 15.735
[9 empty] | 2756 | 15.742
[9 empty] | 2757 | 15.749
[9 empty] | 2758 | 15.755
[9 empty] | 2759 | 15.762
[9 empty] | 2760 | 15.769
[9 empty] | 2761 | 15.775
[9 empty] | 2762 | 15.782
[9 empty] | 2763 | 15.789
[9 empty] | 2764 | 15.795
[9 empty] | 2765 | 15.802
[9 empty] | 2766 | 15.809
[9 empty] | 2767 | 15.815
[9 empty] | 2768 | 15.822
[9 empty] | 2769 | 15.829
[9 empty] | 2770 | 15.835
[9 empty] | 2771 | 15.842
[9 empty] | 2772 | 15.849
[9 empty] | 2773 | 15.855
[9 empty] | 2774 | 15.862
[9 empty] | 2775 | 15.869
[9 empty] | 2776 | 15.875
[9 empty] | 2777 | 15.882
[9 empty] | 2778 | 15.889
[9 empty] | 2779 | 15.895
[9 empty] | 2780 | 15.902
[9 empty] | 2781 | 15.909
[9 empty] | 2782 | 15.915
[9 empty] | 2783 | 15.922
[9 empty] | 2784 | 15.929
[9 empty] | 2785 | 15.935
[9 empty] | 2786 | 15.942
[9 empty] | 2787 | 15.949
[9 empty] | 2788 | 15.955
[9 empty] | 2789 | 15.962
[9 empty] | 2790 | 15.969
[9 empty] | 2791 | 15.975
[9 empty] | 2792 | 15.982
[9 empty] | 2793 | 15.989
[9 empty] | 2794 | 15.995
[9 empty] | 2795 | 16.002
[9 empty] | 2796 | 16.009
[9 empty] | 2797 | 16.015
[9 empty] | 2798 | 16.022
[9 empty] | 2799 | 16.029
[9 empty] | 2800 | 16.035
[9 empty] | 2801 | 16.042
[9 empty] | 2802 | 16.049
[9 empty] | 2803 | 16.055
[9 empty] | 2804 | 16.062
[9 empty] | 2805 | 16.069
[9 empty] | 2806 | 16.075
[9 empty] | 2807 | 16.082
[9 empty] | 2808 | 16.089
[9 empty] | 2809 | 16.095
[9 empty] | 2810 | 16.102
[9 empty] | 2811 | 16.108
[9 empty] | 2812 | 16.115
[9 empty] | 2813 | 16.122
[9 empty] | 2814 | 16.128
[9 empty] | 2815 | 16.135
[9 empty] | 2816 | 16.142
[9 empty] | 2817 | 16.148
[9 empty] | 2818 | 16.155
[9 empty] | 2819 | 16.162
[9 empty] | 2820 | 16.168
[9 empty] | 2821 | 16.175
[9 empty] | 2822 | 16.182
[9 empty] | 2823 | 16.188
[9 empty] | 2824 | 16.195
[9 empty] | 2825 | 16.202
[9 empty] | 2826 | 16.208
[9 empty] | 2827 | 16.215
[9 empty] | 2828 | 16.221
[9 empty] | 2829 | 16.228
[9 empty] | 2830 | 16.235
[9 empty] | 2831 | 16.241
[9 empty] | 2832 | 16.248
[9 empty] | 2833 | 16.255
[9 empty] | 2834 | 16.261
[9 empty] | 2835 | 16.268
[9 empty] | 2836 | 16.275
[9 empty] | 2837 | 16.281
[9 empty] | 2838 | 16.288
[9 empty] | 2839 | 16.294
[9 empty] | 2840 | 16.301
[9 empty] | 2841 | 16.308
[9 empty] | 2842 | 16.314
[9 empty] | 2843 | 16.321
[9 empty] | 2844 | 16.328
[9 empty] | 2845 | 16.334
[9 empty] | 2846 | 16.341
[9 empty] | 2847 | 16.347
[9 empty] | 2848 | 16.354
[9 empty] | 2849 | 16.361
[9 empty] | 2850 | 16.367
[9 empty] | 2851 | 16.374
[9 empty] | 2852 | 16.381
[9 empty] | 2853 | 16.387
[9 empty] | 2854 | 16.394
[9 empty] | 2855 | 16.4
[9 empty] | 2856 | 16.407
[9 empty] | 2857 | 16.414
[9 empty] | 2858 | 16.42
[9 empty] | 2859 | 16.427
[9 empty] | 2860 | 16.434
[9 empty] | 2861 | 16.44
[9 empty] | 2862 | 16.447
[9 empty] | 2863 | 16.453
[9 empty] | 2864 | 16.46
[9 empty] | 2865 | 16.467
[9 empty] | 2866 | 16.473
[9 empty] | 2867 | 16.48
[9 empty] | 2868 | 16.486
[9 empty] | 2869 | 16.493
[9 empty] | 2870 | 16.5
[9 empty] | 2871 | 16.506
[9 empty] | 2872 | 16.513
[9 empty] | 2873 | 16.52
[9 empty] | 2874 | 16.526
[9 empty] | 2875 | 16.533
[9 empty] | 2876 | 16.539
[9 empty] | 2877 | 16.546
[9 empty] | 2878 | 16.553
[9 empty] | 2879 | 16.559
[9 empty] | 2880 | 16.566
[9 empty] | 2881 | 16.572
[9 empty] | 2882 | 16.579
[9 empty] | 2883 | 16.586
[9 empty] | 2884 | 16.592
[9 empty] | 2885 | 16.599
[9 empty] | 2886 | 16.605
[9 empty] | 2887 | 16.612
[9 empty] | 2888 | 16.619
[9 empty] | 2889 | 16.625
[9 empty] | 2890 | 16.632
[9 empty] | 2891 | 16.638
[9 empty] | 2892 | 16.645
[9 empty] | 2893 | 16.652
[9 empty] | 2894 | 16.658
[9 empty] | 2895 | 16.665
[9 empty] | 2896 | 16.671
[9 empty] | 2897 | 16.678
[9 empty] | 2898 | 16.685
[9 empty] | 2899 | 16.691
[9 empty] | 2900 | 16.698
[9 empty] | 2901 | 16.704
[9 empty] | 2902 | 16.711
[9 empty] | 2903 | 16.718
[9 empty] | 2904 | 16.724
[9 empty] | 2905 | 16.731
[9 empty] | 2906 | 16.737
[9 empty] | 2907 | 16.744
[9 empty] | 2908 | 16.751
[9 empty] | 2909 | 16.757
[9 empty] | 2910 | 16.764
[9 empty] | 2911 | 16.77
[9 empty] | 2912 | 16.777
[9 empty] | 2913 | 16.783
[9 empty] | 2914 | 16.79
[9 empty] | 2915 | 16.797
[9 empty] | 2916 | 16.803
[9 empty] | 2917 | 16.81
[9 empty] | 2918 | 16.816
[9 empty] | 2919 | 16.823
[9 empty] | 2920 | 16.829
[9 empty] | 2921 | 16.836
[9 empty] | 2922 | 16.843
[9 empty] | 2923 | 16.849
[9 empty] | 2924 | 16.856
[9 empty] | 2925 | 16.862
[9 empty] | 2926 | 16.869
[9 empty] | 2927 | 16.876
[9 empty] | 2928 | 16.882
[9 empty] | 2929 | 16.889
[9 empty] | 2930 | 16.895
[9 empty] | 2931 | 16.902
[9 empty] | 2932 | 16.908
[9 empty] | 2933 | 16.915
[9 empty] | 2934 | 16.922
[9 empty] | 2935 | 16.928
[9 empty] | 2936 | 16.935
[9 empty] | 2937 | 16.941
[9 empty] | 2938 | 16.948
[9 empty] | 2939 | 16.954
[9 empty] | 2940 | 16.961
[9 empty] | 2941 | 16.967
[9 empty] | 2942 | 16.974
[9 empty] | 2943 | 16.981
[9 empty] | 2944 | 16.987
[9 empty] | 2945 | 16.994
[9 empty] | 2946 | 17
[9 empty] | 2947 | 17.007
[9 empty] | 2948 | 17.013
[9 empty] | 2949 | 17.02
[9 empty] | 2950 | 17.026
[9 empty] | 2951 | 17.033
[9 empty] | 2952 | 17.04
[9 empty] | 2953 | 17.046
[9 empty] | 2954 | 17.053
[9 empty] | 2955 | 17.059
[9 empty] | 2956 | 17.066
[9 empty] | 2957 | 17.072
[9 empty] | 2958 | 17.079
[9 empty] | 2959 | 17.085
[9 empty] | 2960 | 17.092
[9 empty] | 2961 | 17.099
[9 empty] | 2962 | 17.105
[9 empty] | 2963 | 17.112
[9 empty] | 2964 | 17.118
[9 empty] | 2965 | 17.125
[9 empty] | 2966 | 17.131
[9 empty] | 2967 | 17.138
[9 empty] | 2968 | 17.144
[9 empty] | 2969 | 17.151
[9 empty] | 2970 | 17.157
[9 empty] | 2971 | 17.164
[9 empty] | 2972 | 17.171
[9 empty] | 2973 | 17.177
[9 empty] | 2974 | 17.184
[9 empty] | 2975 | 17.19
[9 empty] | 2976 | 17.197
[9 empty] | 2977 | 17.203
[9 empty] | 2978 | 17.21
[9 empty] | 2979 | 17.216
[9 empty] | 2980 | 17.223
[9 empty] | 2981 | 17.229
[9 empty] | 2982 | 17.236
[9 empty] | 2983 | 17.242
[9 empty] | 2984 | 17.249
[9 empty] | 2985 | 17.255
[9 empty] | 2986 | 17.262
[9 empty] | 2987 | 17.268
[9 empty] | 2988 | 17.275
[9 empty] | 2989 | 17.282
[9 empty] | 2990 | 17.288
[9 empty] | 2991 | 17.295
[9 empty] | 2992 | 17.301
[9 empty] | 2993 | 17.308
[9 empty] | 2994 | 17.314
[9 empty] | 2995 | 17.321
[9 empty] | 2996 | 17.327
[9 empty] | 2997 | 17.334
[9 empty] | 2998 | 17.34
[9 empty] | 2999 | 17.347
[9 empty] | 3000 | 17.353
[9 empty] | 3001 | 17.36
[9 empty] | 3002 | 17.366
[9 empty] | 3003 | 17.373
[9 empty] | 3004 | 17.379
[9 empty] | 3005 | 17.386
[9 empty] | 3006 | 17.392
[9 empty] | 3007 | 17.399
[9 empty] | 3008 | 17.405
[9 empty] | 3009 | 17.412
[9 empty] | 3010 | 17.418
[9 empty] | 3011 | 17.425
[9 empty] | 3012 | 17.431
[9 empty] | 3013 | 17.438
[9 empty] | 3014 | 17.444
[9 empty] | 3015 | 17.451
[9 empty] | 3016 | 17.457
[9 empty] | 3017 | 17.464
[9 empty] | 3018 | 17.47
[9 empty] | 3019 | 17.477
[9 empty] | 3020 | 17.483
[9 empty] | 3021 | 17.49
[9 empty] | 3022 | 17.496
[9 empty] | 3023 | 17.503
[9 empty] | 3024 | 17.509
[9 empty] | 3025 | 17.516
[9 empty] | 3026 | 17.522
[9 empty] | 3027 | 17.529
[9 empty] | 3028 | 17.535
[9 empty] | 3029 | 17.542
[9 empty] | 3030 | 17.548
[9 empty] | 3031 | 17.555
[9 empty] | 3032 | 17.561
[9 empty] | 3033 | 17.568
[9 empty] | 3034 | 17.574
[9 empty] | 3035 | 17.581
[9 empty] | 3036 | 17.587
[9 empty] | 3037 | 17.594
[9 empty] | 3038 | 17.6
[9 empty] | 3039 | 17.607
[9 empty] | 3040 | 17.613
[9 empty] | 3041 | 17.62
[9 empty] | 3042 | 17.626
[9 empty] | 3043 | 17.633
[9 empty] | 3044 | 17.639
[9 empty] | 3045 | 17.645
[9 empty] | 3046 | 17.652
[9 empty] | 3047 | 17.658
[9 empty] | 3048 | 17.665
[9 empty] | 3049 | 17.671
[9 empty] | 3050 | 17.678
[9 empty] | 3051 | 17.684
[9 empty] | 3052 | 17.691
[9 empty] | 3053 | 17.697
[9 empty] | 3054 | 17.704
[9 empty] | 3055 | 17.71
[9 empty] | 3056 | 17.717
[9 empty] | 3057 | 17.723
[9 empty] | 3058 | 17.729
[9 empty] | 3059 | 17.736
[9 empty] | 3060 | 17.742
[9 empty] | 3061 | 17.749
[9 empty] | 3062 | 17.755
[9 empty] | 3063 | 17.762
[9 empty] | 3064 | 17.768
[9 empty] | 3065 | 17.775
[9 empty] | 3066 | 17.781
[9 empty] | 3067 | 17.787
[9 empty] | 3068 | 17.794
[9 empty] | 3069 | 17.8
[9 empty] | 3070 | 17.807
[9 empty] | 3071 | 17.813
[9 empty] | 3072 | 17.819
[9 empty] | 3073 | 17.826
[9 empty] | 3074 | 17.832
[9 empty] | 3075 | 17.839
[9 empty] | 3076 | 17.845
[9 empty] | 3077 | 17.852
[9 empty] | 3078 | 17.858
[9 empty] | 3079 | 17.864
[9 empty] | 3080 | 17.871
[9 empty] | 3081 | 17.877
[9 empty] | 3082 | 17.884
[9 empty] | 3083 | 17.89
[9 empty] | 3084 | 17.896
[9 empty] | 3085 | 17.903
[9 empty] | 3086 | 17.909
[9 empty] | 3087 | 17.915
[9 empty] | 3088 | 17.922
[9 empty] | 3089 | 17.928
[9 empty] | 3090 | 17.935
[9 empty] | 3091 | 17.941
[9 empty] | 3092 | 17.947
[9 empty] | 3093 | 17.954
[9 empty] | 3094 | 17.96
[9 empty] | 3095 | 17.966
[9 empty] | 3096 | 17.973
[9 empty] | 3097 | 17.979
[9 empty] | 3098 | 17.985
[9 empty] | 3099 | 17.992
[9 empty] | 3100 | 17.998
[9 empty] | 3101 | 18.004
[9 empty] | 3102 | 18.011
[9 empty] | 3103 | 18.017
[9 empty] | 3104 | 18.023
[9 empty] | 3105 | 18.03
[9 empty] | 3106 | 18.036
[9 empty] | 3107 | 18.042
[9 empty] | 3108 | 18.049
[9 empty] | 3109 | 18.055
[9 empty] | 3110 | 18.061
[9 empty] | 3111 | 18.068
[9 empty] | 3112 | 18.074
[9 empty] | 3113 | 18.08
[9 empty] | 3114 | 18.086
[9 empty] | 3115 | 18.093
[9 empty] | 3116 | 18.099
[9 empty] | 3117 | 18.105
[9 empty] | 3118 | 18.112
[9 empty] | 3119 | 18.118
[9 empty] | 3120 | 18.124
[9 empty] | 3121 | 18.13
[9 empty] | 3122 | 18.137
[9 empty] | 3123 | 18.143
[9 empty] | 3124 | 18.149
[9 empty] | 3125 | 18.155
[9 empty] | 3126 | 18.162
[9 empty] | 3127 | 18.168
[9 empty] | 3128 | 18.174
[9 empty] | 3129 | 18.18
[9 empty] | 3130 | 18.187
[9 empty] | 3131 | 18.193
[9 empty] | 3132 | 18.199
[9 empty] | 3133 | 18.205
[9 empty] | 3134 | 18.211
[9 empty] | 3135 | 18.218
[9 empty] | 3136 | 18.224
[9 empty] | 3137 | 18.23
[9 empty] | 3138 | 18.236
[9 empty] | 3139 | 18.242
[9 empty] | 3140 | 18.249
[9 empty] | 3141 | 18.255
[9 empty] | 3142 | 18.261
[9 empty] | 3143 | 18.267
[9 empty] | 3144 | 18.273
[9 empty] | 3145 | 18.279
[9 empty] | 3146 | 18.285
[9 empty] | 3147 | 18.292
[9 empty] | 3148 | 18.298
[9 empty] | 3149 | 18.304
[9 empty] | 3150 | 18.31
[9 empty] | 3151 | 18.316
[9 empty] | 3152 | 18.322
[9 empty] | 3153 | 18.328
[9 empty] | 3154 | 18.334
[9 empty] | 3155 | 18.341
[9 empty] | 3156 | 18.347
[9 empty] | 3157 | 18.353
[9 empty] | 3158 | 18.359
[9 empty] | 3159 | 18.365
[9 empty] | 3160 | 18.371
[9 empty] | 3161 | 18.377
[9 empty] | 3162 | 18.383
[9 empty] | 3163 | 18.389
[9 empty] | 3164 | 18.395
[9 empty] | 3165 | 18.401
[9 empty] | 3166 | 18.407
[9 empty] | 3167 | 18.413
[9 empty] | 3168 | 18.419
[9 empty] | 3169 | 18.425
[9 empty] | 3170 | 18.431
[9 empty] | 3171 | 18.437
[9 empty] | 3172 | 18.443
[9 empty] | 3173 | 18.449
[9 empty] | 3174 | 18.455
[9 empty] | 3175 | 18.461
[9 empty] | 3176 | 18.467
[9 empty] | 3177 | 18.473
[9 empty] | 3178 | 18.479
[9 empty] | 3179 | 18.485
[9 empty] | 3180 | 18.491
[9 empty] | 3181 | 18.497
[9 empty] | 3182 | 18.503
[9 empty] | 3183 | 18.509
[9 empty] | 3184 | 18.515
[9 empty] | 3185 | 18.521
[9 empty] | 3186 | 18.527
[9 empty] | 3187 | 18.533
[9 empty] | 3188 | 18.539
[9 empty] | 3189 | 18.545
[9 empty] | 3190 | 18.551
[9 empty] | 3191 | 18.557
[9 empty] | 3192 | 18.562
[9 empty] | 3193 | 18.568
[9 empty] | 3194 | 18.574
[9 empty] | 3195 | 18.58
[9 empty] | 3196 | 18.586
[9 empty] | 3197 | 18.592
[9 empty] | 3198 | 18.598
[9 empty] | 3199 | 18.603
[9 empty] | 3200 | 18.609
[9 empty] | 3201 | 18.615
[9 empty] | 3202 | 18.621
[9 empty] | 3203 | 18.627
[9 empty] | 3204 | 18.633
[9 empty] | 3205 | 18.638
[9 empty] | 3206 | 18.644
[9 empty] | 3207 | 18.65
[9 empty] | 3208 | 18.656
[9 empty] | 3209 | 18.661
[9 empty] | 3210 | 18.667
[9 empty] | 3211 | 18.673
[9 empty] | 3212 | 18.679
[9 empty] | 3213 | 18.684
[9 empty] | 3214 | 18.69

## ST-7
[7 empty] | Temp(F) | mV | 32 | 1947.524 | a0 | a1 | a2 | a3 | a4 | a5 | a6 | a7 | a8 | -2 | -3 | -6 | -9 | -12 | -15 | -19 | -22 | -26 | a0 * t^0 | a1* t^1 | a2* t^2 | a3 * t^3 | a4* t^4 | a5*t^5 | a6*t^6 | a7*t^7 | a8*t^8 | 1947.524 | 3028.1 | a0 | a1 | a2 | a3 | a4 | 0 | -3 | -6 | -10 | -15 | a0 * t^0 | a1* t^1 | a2* t^2 | a3 * t^3 | a4* t^4
mV | Temp(F) | [5 empty] | 32 | -0.00012105 |  |  | -9.1942 | 2.7364 | 4.2938 | -4.3974 | 3.3601 | -1.903 | 7.9839 | -2.1059 | 2.4632 | -0.091942 | 0.0027364 | 4.2938e-06 | -4.3974e-09 | 3.3601e-12 | -1.903e-15 | 7.9839e-19 | -2.1059e-22 | 2.4632e-26 | -0.091942 | 0.087564 | 0.0043969 | -0.00014409 | 3.5233e-06 | -6.3854e-08 | 8.5726e-10 | -7.2359e-12 | 2.7083e-14 |  |  | 1.2716 | 1.7242 | 2.053 | -2.8283 | 1.2383 | 1.2716 | 0.0017242 | 2.053e-06 | -2.8283e-10 | 1.2383e-15 | 1.2716 | 0.055176 | 0.0021023 | -9.2679e-06 | 1.2984e-09
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 33 | 0.002881 | [20 empty] | -0.091942 | 0.090301 | 0.004676 | -0.00015803 | 3.9848e-06 | -7.4474e-08 | 1.0311e-09 | -8.9752e-12 | 3.4643e-14 | [12 empty] | 1.2716 | 0.0569 | 0.0022357 | -1.0164e-05 | 1.4685e-09
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 34 | 0.0058907 | [20 empty] | -0.091942 | 0.093037 | 0.0049636 | -0.00017283 | 4.4902e-06 | -8.6463e-08 | 1.2333e-09 | -1.1061e-11 | 4.3988e-14 | [12 empty] | 1.2716 | 0.058624 | 0.0023732 | -1.1116e-05 | 1.6548e-09
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 35 | 0.0089082 | [20 empty] | -0.091942 | 0.095774 | 0.0052599 | -0.00018854 | 5.0422e-06 | -9.9949e-08 | 1.4676e-09 | -1.3549e-11 | 5.5468e-14 | [12 empty] | 1.2716 | 0.060348 | 0.0025149 | -1.2126e-05 | 1.8582e-09
31.778 | 5593.3979 | 2619 | 15.0556 | 0 | 2650 | 0 | 36 | 0.011933 | [20 empty] | -0.091942 | 0.09851 | 0.0055648 | -0.00020516 | 5.6437e-06 | -1.1507e-07 | 1.7379e-09 | -1.6503e-11 | 6.9489e-14 | [12 empty] | 1.2716 | 0.062072 | 0.0026607 | -1.3196e-05 | 2.0798e-09
0.5933 | 199.1475 | 168 | 0.59271 | 0.59672 | 199 | 200 | 37 | 0.014966 | [20 empty] | -0.091942 | 0.10125 | 0.0058782 | -0.00022274 | 6.2974e-06 | -1.3196e-07 | 2.0484e-09 | -1.9992e-11 | 8.6519e-14 | [12 empty] | 1.2716 | 0.063797 | 0.0028105 | -1.4326e-05 | 2.3207e-09
1.4742 | 399.2659 | 368 | 1.4729 | 1.4777 | 399 | 400 | 38 | 0.018007 | [20 empty] | -0.091942 | 0.10398 | 0.0062003 | -0.00024129 | 7.0063e-06 | -1.5078e-07 | 2.4039e-09 | -2.4095e-11 | 1.0709e-13 | [12 empty] | 1.2716 | 0.065521 | 0.0029645 | -1.552e-05 | 2.582e-09
2.4642 | 599.826 | 568 | 2.46 | 2.4651 | 599 | 600 | 39 | 0.021055 | [20 empty] | -0.091942 | 0.10672 | 0.0065309 | -0.00026085 | 7.7734e-06 | -1.717e-07 | 2.8093e-09 | -2.89e-11 | 1.3183e-13 | [12 empty] | 1.2716 | 0.067245 | 0.0031226 | -1.6777e-05 | 2.8647e-09
3.5162 | 800.0861 | 769 | 3.5157 | 3.5211 | 800 | 801 | 40 | 0.024111 | [20 empty] | -0.091942 | 0.10946 | 0.0068701 | -0.00028143 | 8.6019e-06 | -1.9487e-07 | 3.2702e-09 | -3.4504e-11 | 1.6143e-13 | [12 empty] | 1.2716 | 0.068969 | 0.0032848 | -1.8101e-05 | 3.17e-09
4.611 | 1000.1046 | 969 | 4.6104 | 4.616 | 1000 | 1001 | 41 | 0.027174 | [20 empty] | -0.091942 | 0.11219 | 0.0072179 | -0.00030307 | 9.4948e-06 | -2.2047e-07 | 3.7924e-09 | -4.1014e-11 | 1.9668e-13 | [12 empty] | 1.2716 | 0.070694 | 0.0034511 | -1.9493e-05 | 3.4991e-09
5.7377 | 1198.961 | 1167 | 5.7322 | 5.7379 | 1198 | 1199 | 42 | 0.030245 | [20 empty] | -0.091942 | 0.11493 | 0.0075743 | -0.00032579 | 1.0456e-05 | -2.487e-07 | 4.3824e-09 | -4.855e-11 | 2.385e-13 | [12 empty] | 1.2716 | 0.072418 | 0.0036215 | -2.0955e-05 | 3.8531e-09
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 43 | 0.033324 | [20 empty] | -0.091942 | 0.11766 | 0.0079393 | -0.00034962 | 1.1488e-05 | -2.7976e-07 | 5.0469e-09 | -5.7243e-11 | 2.879e-13 | [12 empty] | 1.2716 | 0.074142 | 0.003796 | -2.2487e-05 | 4.2334e-09
-50 | -16623.5199 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 44 | 0.03641 | [20 empty] | -0.091942 | 0.1204 | 0.0083128 | -0.00037459 | 1.2594e-05 | -3.1383e-07 | 5.7933e-09 | -6.7238e-11 | 3.4604e-13 | [12 empty] | 1.2716 | 0.075866 | 0.0039746 | -2.4093e-05 | 4.6412e-09
31.841 | 5604.4868 | 2619 | 15.0556 | 0 | 2650 | 0 | 45 | 0.039503 | [20 empty] | -0.091942 | 0.12314 | 0.008695 | -0.00040071 | 1.3779e-05 | -3.5116e-07 | 6.6296e-09 | -7.8692e-11 | 4.1419e-13 | [12 empty] | 1.2716 | 0.077591 | 0.0041573 | -2.5773e-05 | 5.0777e-09
0.5933 | 199.1475 | 168 | 0.59271 | 0.59672 | 199 | 200 | 46 | 0.042604 | [20 empty] | -0.091942 | 0.12587 | 0.0090857 | -0.00042802 | 1.5045e-05 | -3.9195e-07 | 7.5641e-09 | -9.178e-11 | 4.9381e-13 | [12 empty] | 1.2716 | 0.079315 | 0.0043441 | -2.753e-05 | 5.5443e-09
1.4742 | 399.2659 | 368 | 1.4729 | 1.4777 | 399 | 400 | 47 | 0.045713 | [20 empty] | -0.091942 | 0.12861 | 0.009485 | -0.00045655 | 1.6396e-05 | -4.3644e-07 | 8.606e-09 | -1.0669e-10 | 5.8652e-13 | [12 empty] | 1.2716 | 0.081039 | 0.004535 | -2.9365e-05 | 6.0424e-09
2.4642 | 599.826 | 568 | 2.46 | 2.4651 | 599 | 600 | 48 | 0.048829 | [20 empty] | -0.091942 | 0.13135 | 0.0098929 | -0.00048631 | 1.7837e-05 | -4.8489e-07 | 9.7647e-09 | -1.2363e-10 | 6.9411e-13 | [12 empty] | 1.2716 | 0.082763 | 0.0047301 | -3.1279e-05 | 6.5733e-09
3.5162 | 800.0861 | 769 | 3.5157 | 3.5211 | 800 | 801 | 49 | 0.051952 | [20 empty] | -0.091942 | 0.13408 | 0.010309 | -0.00051735 | 1.937e-05 | -5.3755e-07 | 1.1051e-08 | -1.4283e-10 | 8.1859e-13 | [12 empty] | 1.2716 | 0.084488 | 0.0049292 | -3.3275e-05 | 7.1384e-09
4.611 | 1000.1046 | 969 | 4.6104 | 4.616 | 1000 | 1001 | 50 | 0.055083 | [20 empty] | -0.091942 | 0.13682 | 0.010735 | -0.00054967 | 2.1001e-05 | -5.9469e-07 | 1.2475e-08 | -1.6453e-10 | 9.6219e-13 | [12 empty] | 1.2716 | 0.086212 | 0.0051325 | -3.5354e-05 | 7.7392e-09
5.7377 | 1198.961 | 1167 | 5.7322 | 5.7379 | 1198 | 1199 | 51 | 0.058221 | [20 empty] | -0.091942 | 0.13956 | 0.011168 | -0.00058332 | 2.2732e-05 | -6.5658e-07 | 1.4049e-08 | -1.8899e-10 | 1.1274e-12 | [12 empty] | 1.2716 | 0.087936 | 0.0053398 | -3.7518e-05 | 8.3772e-09
6.9019 | 1397.48 | 1366 | 6.899 | 6.905 | 1397 | 1398 | 52 | 0.061366 | [20 empty] | -0.091942 | 0.14229 | 0.01161 | -0.00061831 | 2.4568e-05 | -7.2353e-07 | 1.5785e-08 | -2.165e-10 | 1.3168e-12 | [12 empty] | 1.2716 | 0.08966 | 0.0055513 | -3.9769e-05 | 9.0538e-09
8.1222 | 1598.2896 | 1567 | 8.1204 | 8.1266 | 1598 | 1599 | 53 | 0.064519 | [20 empty] | -0.091942 | 0.14503 | 0.012061 | -0.00065467 | 2.6513e-05 | -7.9582e-07 | 1.7696e-08 | -2.4739e-10 | 1.5336e-12 | [12 empty] | 1.2716 | 0.091384 | 0.0057668 | -4.2107e-05 | 9.7706e-09
9.3802 | 1798.3045 | 1767 | 9.3783 | 9.3846 | 1798 | 1799 | 54 | 0.067679 | [20 empty] | -0.091942 | 0.14776 | 0.012521 | -0.00069243 | 2.8571e-05 | -8.7379e-07 | 1.9796e-08 | -2.8197e-10 | 1.7809e-12 | [12 empty] | 1.2716 | 0.093109 | 0.0059865 | -4.4536e-05 | 1.0529e-08
10.6898 | 2000.0968 | 1969 | 10.6892 | 10.6957 | 2000 | 2001 | 55 | 0.070847 | [20 empty] | -0.091942 | 0.1505 | 0.012989 | -0.00073161 | 3.0747e-05 | -9.5775e-07 | 2.21e-08 | -3.2061e-10 | 2.0625e-12 | [12 empty] | 1.2716 | 0.094833 | 0.0062103 | -4.7056e-05 | 1.1331e-08
12.006 | 2198.0952 | 2167 | 12.0054 | 12.0121 | 2198 | 2199 | 56 | 0.074021 | [20 empty] | -0.091942 | 0.15324 | 0.013465 | -0.00077225 | 3.3045e-05 | -1.048e-06 | 2.4623e-08 | -3.6371e-10 | 2.3823e-12 | [12 empty] | 1.2716 | 0.096557 | 0.0064382 | -4.967e-05 | 1.2178e-08
NO | #VALUE! | 1 | -0.00012105 | 0.002881 | 32 | 33 | 57 | 0.077203 | [20 empty] | -0.091942 | 0.15597 | 0.013951 | -0.00081436 | 3.5469e-05 | -1.145e-06 | 2.7382e-08 | -4.1169e-10 | 2.7447e-12 | [12 empty] | 1.2716 | 0.098281 | 0.0066701 | -5.2379e-05 | 1.3071e-08
0.5933 | 199.1475 | 168 | 0.59271 | 0.59672 | 199 | 200 | 58 | 0.080392 | [20 empty] | -0.091942 | 0.15871 | 0.014444 | -0.00085798 | 3.8025e-05 | -1.249e-06 | 3.0393e-08 | -4.6499e-10 | 3.1544e-12 | [12 empty] | 1.2716 | 0.10001 | 0.0069062 | -5.5184e-05 | 1.4013e-08
1.4742 | 399.2659 | 368 | 1.4729 | 1.4777 | 399 | 400 | 59 | 0.083588 | [20 empty] | -0.091942 | 0.16145 | 0.014947 | -0.00090313 | 4.0716e-05 | -1.3605e-06 | 3.3676e-08 | -5.2409e-10 | 3.6167e-12 | [12 empty] | 1.2716 | 0.10173 | 0.0071464 | -5.8088e-05 | 1.5005e-08
2.4642 | 599.826 | 568 | 2.46 | 2.4651 | 599 | 600 | 60 | 0.086792 | [20 empty] | -0.091942 | 0.16418 | 0.015458 | -0.00094983 | 4.3547e-05 | -1.4798e-06 | 3.7249e-08 | -5.8953e-10 | 4.1372e-12 | [12 empty] | 1.2716 | 0.10345 | 0.0073907 | -6.1092e-05 | 1.6048e-08
3.5162 | 800.0861 | 769 | 3.5157 | 3.5211 | 800 | 801 | 61 | 0.090002 | [20 empty] | -0.091942 | 0.16692 | 0.015977 | -0.00099812 | 4.6523e-05 | -1.6073e-06 | 4.1133e-08 | -6.6184e-10 | 4.7221e-12 | [12 empty] | 1.2716 | 0.10518 | 0.0076392 | -6.4198e-05 | 1.7145e-08
200 | 35202.9573 | 2619 | 15.0556 | 0 | 2650 | 0 | 62 | 0.09322 | [20 empty] | -0.091942 | 0.16966 | 0.016505 | -0.001048 | 4.965e-05 | -1.7434e-06 | 4.5348e-08 | -7.4163e-10 | 5.3781e-12 | [12 empty] | 1.2716 | 0.1069 | 0.0078917 | -6.7407e-05 | 1.8297e-08
4.611 | 1000.1046 | 969 | 4.6104 | 4.616 | 1000 | 1001 | 63 | 0.096444 | [20 empty] | -0.091942 | 0.17239 | 0.017042 | -0.0010995 | 5.2931e-05 | -1.8886e-06 | 4.9918e-08 | -8.2952e-10 | 6.1126e-12 | [12 empty] | 1.2716 | 0.10863 | 0.0081483 | -7.0722e-05 | 1.9507e-08
5.7377 | 1198.961 | 1167 | 5.7322 | 5.7379 | 1198 | 1199 | 64 | 0.099676 | [20 empty] | -0.091942 | 0.17513 | 0.017587 | -0.0011527 | 5.6373e-05 | -2.0433e-06 | 5.4865e-08 | -9.262e-10 | 6.9333e-12 | [12 empty] | 1.2716 | 0.11035 | 0.008409 | -7.4143e-05 | 2.0775e-08
31.778 | 5593.3979 | 2619 | 15.0556 | 0 | 2650 | 0 | 65 | 0.10292 | [20 empty] | -0.091942 | 0.17787 | 0.018141 | -0.0012076 | 5.998e-05 | -2.208e-06 | 6.0213e-08 | -1.0324e-09 | 7.8488e-12 | [12 empty] | 1.2716 | 0.11208 | 0.0086739 | -7.7673e-05 | 2.2104e-08
| 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 66 | 0.10616 | [20 empty] | -0.091942 | 0.1806 | 0.018704 | -0.0012642 | 6.3757e-05 | -2.3832e-06 | 6.599e-08 | -1.1488e-09 | 8.8685e-12 | [12 empty] | 1.2716 | 0.1138 | 0.0089428 | -8.1314e-05 | 2.3496e-08
| 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 67 | 0.10941 | [20 empty] | -0.091942 | 0.18334 | 0.019275 | -0.0013226 | 6.771e-05 | -2.5693e-06 | 7.2221e-08 | -1.2763e-09 | 1.0002e-11 | [12 empty] | 1.2716 | 0.11552 | 0.0092158 | -8.5066e-05 | 2.4953e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 68 | 0.11267 | [20 empty] | -0.091942 | 0.18607 | 0.019855 | -0.0013827 | 7.1844e-05 | -2.7668e-06 | 7.8934e-08 | -1.4158e-09 | 1.1261e-11 | [12 empty] | 1.2716 | 0.11725 | 0.009493 | -8.8932e-05 | 2.6476e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 69 | 0.11594 | [20 empty] | -0.091942 | 0.18881 | 0.020443 | -0.0014446 | 7.6164e-05 | -2.9763e-06 | 8.616e-08 | -1.5682e-09 | 1.2656e-11 | [12 empty] | 1.2716 | 0.11897 | 0.0097743 | -9.2913e-05 | 2.8068e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 70 | 0.11921 | [20 empty] | -0.091942 | 0.19155 | 0.02104 | -0.0015083 | 8.0676e-05 | -3.1984e-06 | 9.3929e-08 | -1.7343e-09 | 1.42e-11 | [12 empty] | 1.2716 | 0.1207 | 0.01006 | -9.7012e-05 | 2.9731e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 71 | 0.1225 | [20 empty] | -0.091942 | 0.19428 | 0.021645 | -0.0015739 | 8.5386e-05 | -3.4334e-06 | 1.0227e-07 | -1.9154e-09 | 1.5906e-11 | [12 empty] | 1.2716 | 0.12242 | 0.010349 | -0.00010123 | 3.1467e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 72 | 0.12578 | [20 empty] | -0.091942 | 0.19702 | 0.022259 | -0.0016413 | 9.0299e-05 | -3.6821e-06 | 1.1123e-07 | -2.1124e-09 | 1.7789e-11 | [12 empty] | 1.2716 | 0.12414 | 0.010643 | -0.00010557 | 3.3277e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 73 | 0.12908 | [20 empty] | -0.091942 | 0.19976 | 0.022882 | -0.0017107 | 9.5421e-05 | -3.945e-06 | 1.2082e-07 | -2.3265e-09 | 1.9865e-11 | [12 empty] | 1.2716 | 0.12587 | 0.01094 | -0.00011003 | 3.5165e-08
450 | 79206.654 | 2619 | 15.0556 | 0 | 2650 | 0 | 74 | 0.13238 | [20 empty] | -0.091942 | 0.20249 | 0.023513 | -0.0017819 | 0.00010076 | -4.2228e-06 | 1.311e-07 | -2.559e-09 | 2.2149e-11 | [12 empty] | 1.2716 | 0.12759 | 0.011242 | -0.00011461 | 3.7132e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 75 | 0.13569 | [20 empty] | -0.091942 | 0.20523 | 0.024153 | -0.0018551 | 0.00010632 | -4.5159e-06 | 1.421e-07 | -2.8111e-09 | 2.466e-11 | [12 empty] | 1.2716 | 0.12932 | 0.011548 | -0.00011932 | 3.918e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 76 | 0.139 | [20 empty] | -0.091942 | 0.20797 | 0.024801 | -0.0019303 | 0.0001121 | -4.8251e-06 | 1.5385e-07 | -3.0842e-09 | 2.7416e-11 | [12 empty] | 1.2716 | 0.13104 | 0.011858 | -0.00012416 | 4.1312e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 77 | 0.14232 | [20 empty] | -0.091942 | 0.2107 | 0.025458 | -0.0020075 | 0.00011812 | -5.151e-06 | 1.664e-07 | -3.3797e-09 | 3.0439e-11 | [12 empty] | 1.2716 | 0.13277 | 0.012172 | -0.00012912 | 4.3529e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 78 | 0.14565 | [20 empty] | -0.091942 | 0.21344 | 0.026124 | -0.0020868 | 0.00012437 | -5.4943e-06 | 1.798e-07 | -3.6992e-09 | 3.3749e-11 | [12 empty] | 1.2716 | 0.13449 | 0.01249 | -0.00013422 | 4.5835e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 79 | 0.14899 | [20 empty] | -0.091942 | 0.21617 | 0.026798 | -0.0021681 | 0.00013088 | -5.8556e-06 | 1.9408e-07 | -4.0442e-09 | 3.7369e-11 | [12 empty] | 1.2716 | 0.13621 | 0.012813 | -0.00013945 | 4.8231e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 80 | 0.15233 | [20 empty] | -0.091942 | 0.21891 | 0.02748 | -0.0022515 | 0.00013763 | -6.2357e-06 | 2.0929e-07 | -4.4165e-09 | 4.1326e-11 | [12 empty] | 1.2716 | 0.13794 | 0.013139 | -0.00014481 | 5.072e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 81 | 0.15568 | [20 empty] | -0.091942 | 0.22165 | 0.028172 | -0.0023369 | 0.00014464 | -6.6353e-06 | 2.2549e-07 | -4.8177e-09 | 4.5644e-11 | [12 empty] | 1.2716 | 0.13966 | 0.01347 | -0.00015031 | 5.3304e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 82 | 0.15903 | [20 empty] | -0.091942 | 0.22438 | 0.028872 | -0.0024246 | 0.00015192 | -7.0551e-06 | 2.4271e-07 | -5.2498e-09 | 5.0351e-11 | [12 empty] | 1.2716 | 0.14139 | 0.013804 | -0.00015595 | 5.5985e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 83 | 0.1624 | [20 empty] | -0.091942 | 0.22712 | 0.02958 | -0.0025144 | 0.00015946 | -7.496e-06 | 2.6102e-07 | -5.7147e-09 | 5.5478e-11 | [12 empty] | 1.2716 | 0.14311 | 0.014143 | -0.00016172 | 5.8767e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 84 | 0.16577 | [20 empty] | -0.091942 | 0.22986 | 0.030297 | -0.0026063 | 0.00016729 | -7.9585e-06 | 2.8047e-07 | -6.2144e-09 | 6.1057e-11 | [12 empty] | 1.2716 | 0.14484 | 0.014486 | -0.00016764 | 6.165e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 85 | 0.16914 | [20 empty] | -0.091942 | 0.23259 | 0.031023 | -0.0027005 | 0.0001754 | -8.4437e-06 | 3.0111e-07 | -6.7511e-09 | 6.712e-11 | [12 empty] | 1.2716 | 0.14656 | 0.014833 | -0.0001737 | 6.4639e-08
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 86 | 0.17252 | [20 empty] | -0.091942 | 0.23533 | 0.031757 | -0.002797 | 0.0001838 | -8.9522e-06 | 3.23e-07 | -7.3271e-09 | 7.3703e-11 | [12 empty] | 1.2716 | 0.14828 | 0.015184 | -0.0001799 | 6.7735e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 87 | 0.17591 | [20 empty] | -0.091942 | 0.23807 | 0.0325 | -0.0028957 | 0.0001925 | -9.4849e-06 | 3.462e-07 | -7.9447e-09 | 8.0845e-11 | [12 empty] | 1.2716 | 0.15001 | 0.015539 | -0.00018625 | 7.0941e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 88 | 0.17931 | [20 empty] | -0.091942 | 0.2408 | 0.033251 | -0.0029967 | 0.0002015 | -1.0043e-05 | 3.7077e-07 | -8.6064e-09 | 8.8585e-11 | [12 empty] | 1.2716 | 0.15173 | 0.015898 | -0.00019274 | 7.4259e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 89 | 0.18271 | [20 empty] | -0.091942 | 0.24354 | 0.034011 | -0.0031 | 0.00021082 | -1.0626e-05 | 3.9678e-07 | -9.3148e-09 | 9.6966e-11 | [12 empty] | 1.2716 | 0.15346 | 0.016262 | -0.00019939 | 7.7692e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 90 | 0.18612 | [20 empty] | -0.091942 | 0.24627 | 0.03478 | -0.0032057 | 0.00022046 | -1.1237e-05 | 4.2429e-07 | -1.0073e-08 | 1.0603e-10 | [12 empty] | 1.2716 | 0.15518 | 0.016629 | -0.00020619 | 8.1243e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 91 | 0.18953 | [20 empty] | -0.091942 | 0.24901 | 0.035557 | -0.0033137 | 0.00023042 | -1.1875e-05 | 4.5338e-07 | -1.0883e-08 | 1.1583e-10 | [12 empty] | 1.2716 | 0.15691 | 0.017001 | -0.00021314 | 8.4915e-08
77 | 13553.1386 | 2619 | 15.0556 | 0 | 2650 | 0 | 92 | 0.19295 | [20 empty] | -0.091942 | 0.25175 | 0.036343 | -0.0034242 | 0.00024072 | -1.2542e-05 | 4.841e-07 | -1.1748e-08 | 1.2642e-10 | [12 empty] | 1.2716 | 0.15863 | 0.017376 | -0.00022024 | 8.8709e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 93 | 0.19638 | [20 empty] | -0.091942 | 0.25448 | 0.037137 | -0.0035371 | 0.00025135 | -1.3239e-05 | 5.1655e-07 | -1.2671e-08 | 1.3784e-10 | [12 empty] | 1.2716 | 0.16035 | 0.017756 | -0.0002275 | 9.263e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 94 | 0.19982 | [20 empty] | -0.091942 | 0.25722 | 0.03794 | -0.0036524 | 0.00026234 | -1.3966e-05 | 5.5078e-07 | -1.3656e-08 | 1.5015e-10 | [12 empty] | 1.2716 | 0.16208 | 0.01814 | -0.00023492 | 9.6679e-08
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 95 | 0.20326 | [20 empty] | -0.091942 | 0.25996 | 0.038752 | -0.0037702 | 0.00027368 | -1.4725e-05 | 5.8689e-07 | -1.4707e-08 | 1.6341e-10 | [12 empty] | 1.2716 | 0.1638 | 0.018528 | -0.00024249 | 1.0086e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 96 | 0.2067 | [20 empty] | -0.091942 | 0.26269 | 0.039572 | -0.0038905 | 0.00028539 | -1.5516e-05 | 6.2494e-07 | -1.5825e-08 | 1.7769e-10 | [12 empty] | 1.2716 | 0.16553 | 0.01892 | -0.00025023 | 1.0517e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 97 | 0.21016 | [20 empty] | -0.091942 | 0.26543 | 0.0404 | -0.0040134 | 0.00029747 | -1.6342e-05 | 6.6503e-07 | -1.7016e-08 | 1.9305e-10 | [12 empty] | 1.2716 | 0.16725 | 0.019317 | -0.00025813 | 1.0962e-07
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 98 | 0.21362 | [20 empty] | -0.091942 | 0.26817 | 0.041238 | -0.0041388 | 0.00030992 | -1.7202e-05 | 7.0724e-07 | -1.8282e-08 | 2.0956e-10 | [12 empty] | 1.2716 | 0.16898 | 0.019717 | -0.0002662 | 1.1421e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 99 | 0.21708 | [20 empty] | -0.091942 | 0.2709 | 0.042084 | -0.0042668 | 0.00032277 | -1.8097e-05 | 7.5166e-07 | -1.9629e-08 | 2.2729e-10 | [12 empty] | 1.2716 | 0.1707 | 0.020121 | -0.00027443 | 1.1895e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 100 | 0.22056 | [20 empty] | -0.091942 | 0.27364 | 0.042938 | -0.0043974 | 0.00033601 | -1.903e-05 | 7.9839e-07 | -2.1059e-08 | 2.4632e-10 | [12 empty] | 1.2716 | 0.17242 | 0.02053 | -0.00028283 | 1.2383e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 101 | 0.22403 | [20 empty] | -0.091942 | 0.27638 | 0.043801 | -0.0045306 | 0.00034965 | -2.0001e-05 | 8.475e-07 | -2.2578e-08 | 2.6673e-10 | [12 empty] | 1.2716 | 0.17415 | 0.020942 | -0.0002914 | 1.2886e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 102 | 0.22752 | [20 empty] | -0.091942 | 0.27911 | 0.044673 | -0.0046665 | 0.00036371 | -2.1011e-05 | 8.9911e-07 | -2.4191e-08 | 2.886e-10 | [12 empty] | 1.2716 | 0.17587 | 0.021359 | -0.00030015 | 1.3404e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 103 | 0.23101 | [20 empty] | -0.091942 | 0.28185 | 0.045553 | -0.0048051 | 0.00037818 | -2.2061e-05 | 9.5331e-07 | -2.59e-08 | 3.1203e-10 | [12 empty] | 1.2716 | 0.1776 | 0.02178 | -0.00030906 | 1.3937e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 104 | 0.23451 | [20 empty] | -0.091942 | 0.28458 | 0.046442 | -0.0049464 | 0.00039308 | -2.3153e-05 | 1.0102e-06 | -2.7713e-08 | 3.3711e-10 | [12 empty] | 1.2716 | 0.17932 | 0.022205 | -0.00031815 | 1.4486e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 105 | 0.23801 | [20 empty] | -0.091942 | 0.28732 | 0.047339 | -0.0050905 | 0.00040842 | -2.4288e-05 | 1.0699e-06 | -2.9633e-08 | 3.6393e-10 | [12 empty] | 1.2716 | 0.18104 | 0.022634 | -0.00032742 | 1.5051e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 106 | 0.24152 | [20 empty] | -0.091942 | 0.29006 | 0.048245 | -0.0052373 | 0.0004242 | -2.5466e-05 | 1.1325e-06 | -3.1665e-08 | 3.926e-10 | [12 empty] | 1.2716 | 0.18277 | 0.023067 | -0.00033686 | 1.5633e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 107 | 0.24504 | [20 empty] | -0.091942 | 0.29279 | 0.04916 | -0.005387 | 0.00044044 | -2.669e-05 | 1.1982e-06 | -3.3817e-08 | 4.2322e-10 | [12 empty] | 1.2716 | 0.18449 | 0.023505 | -0.00034648 | 1.6231e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 108 | 0.24856 | [20 empty] | -0.091942 | 0.29553 | 0.050083 | -0.0055394 | 0.00045714 | -2.7961e-05 | 1.2669e-06 | -3.6092e-08 | 4.5592e-10 | [12 empty] | 1.2716 | 0.18622 | 0.023946 | -0.00035629 | 1.6847e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 109 | 0.25209 | [20 empty] | -0.091942 | 0.29827 | 0.051015 | -0.0056947 | 0.00047431 | -2.928e-05 | 1.339e-06 | -3.8497e-08 | 4.9081e-10 | [12 empty] | 1.2716 | 0.18794 | 0.024392 | -0.00036628 | 1.7479e-07
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 110 | 0.25563 | [20 empty] | -0.091942 | 0.301 | 0.051955 | -0.0058529 | 0.00049195 | -3.0648e-05 | 1.4144e-06 | -4.1039e-08 | 5.2801e-10 | [12 empty] | 1.2716 | 0.18967 | 0.024841 | -0.00037645 | 1.813e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 111 | 0.25917 | [20 empty] | -0.091942 | 0.30374 | 0.052904 | -0.006014 | 0.00051009 | -3.2067e-05 | 1.4933e-06 | -4.3723e-08 | 5.6765e-10 | [12 empty] | 1.2716 | 0.19139 | 0.025295 | -0.00038681 | 1.8798e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 112 | 0.26271 | [20 empty] | -0.091942 | 0.30648 | 0.053862 | -0.006178 | 0.00052872 | -3.3537e-05 | 1.5759e-06 | -4.6555e-08 | 6.0988e-10 | [12 empty] | 1.2716 | 0.19311 | 0.025753 | -0.00039736 | 1.9485e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 113 | 0.26627 | [20 empty] | -0.091942 | 0.30921 | 0.054828 | -0.006345 | 0.00054786 | -3.5061e-05 | 1.6622e-06 | -4.9544e-08 | 6.5483e-10 | [12 empty] | 1.2716 | 0.19484 | 0.026215 | -0.0004081 | 2.019e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 114 | 0.26983 | [20 empty] | -0.091942 | 0.31195 | 0.055802 | -0.0065149 | 0.00056751 | -3.664e-05 | 1.7524e-06 | -5.2696e-08 | 7.0265e-10 | [12 empty] | 1.2716 | 0.19656 | 0.026681 | -0.00041903 | 2.0914e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 115 | 0.27339 | [20 empty] | -0.091942 | 0.31468 | 0.056786 | -0.0066878 | 0.00058768 | -3.8276e-05 | 1.8467e-06 | -5.6018e-08 | 7.535e-10 | [12 empty] | 1.2716 | 0.19829 | 0.027151 | -0.00043015 | 2.1658e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 116 | 0.27696 | [20 empty] | -0.091942 | 0.31742 | 0.057778 | -0.0068638 | 0.00060839 | -3.9969e-05 | 1.9452e-06 | -5.9518e-08 | 8.0754e-10 | [12 empty] | 1.2716 | 0.20001 | 0.027625 | -0.00044147 | 2.2421e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 117 | 0.28054 | [20 empty] | -0.091942 | 0.32016 | 0.058778 | -0.0070429 | 0.00062964 | -4.1722e-05 | 2.048e-06 | -6.3204e-08 | 8.6494e-10 | [12 empty] | 1.2716 | 0.20174 | 0.028103 | -0.00045299 | 2.3204e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 118 | 0.28412 | [20 empty] | -0.091942 | 0.32289 | 0.059787 | -0.007225 | 0.00065145 | -4.3536e-05 | 2.1553e-06 | -6.7084e-08 | 9.2588e-10 | [12 empty] | 1.2716 | 0.20346 | 0.028586 | -0.0004647 | 2.4007e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 119 | 0.28771 | [20 empty] | -0.091942 | 0.32563 | 0.060805 | -0.0074103 | 0.00067381 | -4.5412e-05 | 2.2672e-06 | -7.1166e-08 | 9.9055e-10 | [12 empty] | 1.2716 | 0.20518 | 0.029072 | -0.00047662 | 2.4832e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 120 | 0.29131 | [20 empty] | -0.091942 | 0.32837 | 0.061831 | -0.0075987 | 0.00069675 | -4.7353e-05 | 2.384e-06 | -7.5459e-08 | 1.0591e-09 | [12 empty] | 1.2716 | 0.20691 | 0.029563 | -0.00048874 | 2.5677e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 121 | 0.29491 | [20 empty] | -0.091942 | 0.3311 | 0.062866 | -0.0077902 | 0.00072027 | -4.9359e-05 | 2.5057e-06 | -7.9973e-08 | 1.1318e-09 | [12 empty] | 1.2716 | 0.20863 | 0.030058 | -0.00050106 | 2.6544e-07
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 122 | 0.29852 | [20 empty] | -0.091942 | 0.33384 | 0.063909 | -0.007985 | 0.00074437 | -5.1432e-05 | 2.6325e-06 | -8.4716e-08 | 1.2089e-09 | [12 empty] | 1.2716 | 0.21036 | 0.030557 | -0.00051358 | 2.7432e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 123 | 0.30213 | [20 empty] | -0.091942 | 0.33658 | 0.064961 | -0.0081829 | 0.00076908 | -5.3575e-05 | 2.7647e-06 | -8.9697e-08 | 1.2904e-09 | [12 empty] | 1.2716 | 0.21208 | 0.03106 | -0.00052632 | 2.8343e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 124 | 0.30575 | [20 empty] | -0.091942 | 0.33931 | 0.066022 | -0.0083841 | 0.0007944 | -5.5789e-05 | 2.9023e-06 | -9.4928e-08 | 1.3768e-09 | [12 empty] | 1.2716 | 0.21381 | 0.031567 | -0.00053926 | 2.9276e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 125 | 0.30937 | [20 empty] | -0.091942 | 0.34205 | 0.067091 | -0.0085886 | 0.00082034 | -5.8075e-05 | 3.0456e-06 | -1.0042e-07 | 1.4682e-09 | [12 empty] | 1.2716 | 0.21553 | 0.032078 | -0.00055241 | 3.0231e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 126 | 0.313 | [20 empty] | -0.091942 | 0.34478 | 0.068169 | -0.0087964 | 0.0008469 | -6.0435e-05 | 3.1947e-06 | -1.0618e-07 | 1.5648e-09 | [12 empty] | 1.2716 | 0.21725 | 0.032593 | -0.00056577 | 3.121e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 127 | 0.31664 | [20 empty] | -0.091942 | 0.34752 | 0.069255 | -0.0090075 | 0.00087411 | -6.2872e-05 | 3.3499e-06 | -1.1222e-07 | 1.667e-09 | [12 empty] | 1.2716 | 0.21898 | 0.033113 | -0.00057935 | 3.2213e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 128 | 0.32028 | [20 empty] | -0.091942 | 0.35026 | 0.07035 | -0.009222 | 0.00090197 | -6.5386e-05 | 3.5113e-06 | -1.1855e-07 | 1.7749e-09 | [12 empty] | 1.2716 | 0.2207 | 0.033636 | -0.00059315 | 3.324e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 129 | 0.32393 | [20 empty] | -0.091942 | 0.35299 | 0.071453 | -0.0094398 | 0.00093049 | -6.7981e-05 | 3.6792e-06 | -1.2519e-07 | 1.8889e-09 | [12 empty] | 1.2716 | 0.22243 | 0.034164 | -0.00060716 | 3.4291e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 130 | 0.32759 | [20 empty] | -0.091942 | 0.35573 | 0.072565 | -0.009661 | 0.00095968 | -7.0657e-05 | 3.8537e-06 | -1.3214e-07 | 2.0093e-09 | [12 empty] | 1.2716 | 0.22415 | 0.034695 | -0.00062139 | 3.5366e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 131 | 0.33125 | [20 empty] | -0.091942 | 0.35847 | 0.073686 | -0.0098857 | 0.00098955 | -7.3416e-05 | 4.035e-06 | -1.3943e-07 | 2.1363e-09 | [12 empty] | 1.2716 | 0.22587 | 0.035231 | -0.00063584 | 3.6467e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 132 | 0.33491 | [20 empty] | -0.091942 | 0.3612 | 0.074815 | -0.010114 | 0.0010201 | -7.6262e-05 | 4.2233e-06 | -1.4705e-07 | 2.2703e-09 | [12 empty] | 1.2716 | 0.2276 | 0.035771 | -0.00065051 | 3.7594e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 133 | 0.33858 | [20 empty] | -0.091942 | 0.36394 | 0.075953 | -0.010345 | 0.0010514 | -7.9195e-05 | 4.419e-06 | -1.5503e-07 | 2.4116e-09 | [12 empty] | 1.2716 | 0.22932 | 0.036315 | -0.00066541 | 3.8746e-07
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 134 | 0.34226 | [20 empty] | -0.091942 | 0.36668 | 0.0771 | -0.010581 | 0.0010834 | -8.2217e-05 | 4.6221e-06 | -1.6337e-07 | 2.5606e-09 | [12 empty] | 1.2716 | 0.23105 | 0.036863 | -0.00068053 | 3.9924e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 135 | 0.34594 | [20 empty] | -0.091942 | 0.36941 | 0.078255 | -0.010819 | 0.0011161 | -8.5331e-05 | 4.833e-06 | -1.721e-07 | 2.7175e-09 | [12 empty] | 1.2716 | 0.23277 | 0.037416 | -0.00069588 | 4.113e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 136 | 0.34963 | [20 empty] | -0.091942 | 0.37215 | 0.079418 | -0.011061 | 0.0011495 | -8.8538e-05 | 5.0518e-06 | -1.8122e-07 | 2.8828e-09 | [12 empty] | 1.2716 | 0.2345 | 0.037972 | -0.00071146 | 4.2362e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 137 | 0.35332 | [20 empty] | -0.091942 | 0.37489 | 0.080591 | -0.011307 | 0.0011837 | -9.1842e-05 | 5.2788e-06 | -1.9076e-07 | 3.0568e-09 | [12 empty] | 1.2716 | 0.23622 | 0.038532 | -0.00072727 | 4.3622e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 138 | 0.35702 | [20 empty] | -0.091942 | 0.37762 | 0.081771 | -0.011557 | 0.0012186 | -9.5243e-05 | 5.5143e-06 | -2.0072e-07 | 3.2399e-09 | [12 empty] | 1.2716 | 0.23794 | 0.039097 | -0.00074331 | 4.4909e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 139 | 0.36073 | [20 empty] | -0.091942 | 0.38036 | 0.082961 | -0.01181 | 0.0012543 | -9.8744e-05 | 5.7584e-06 | -2.1113e-07 | 3.4326e-09 | [12 empty] | 1.2716 | 0.23967 | 0.039666 | -0.00075958 | 4.6225e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 140 | 0.36444 | [20 empty] | -0.091942 | 0.38309 | 0.084159 | -0.012066 | 0.0012908 | -0.00010235 | 6.0115e-06 | -2.2199e-07 | 3.6352e-09 | [12 empty] | 1.2716 | 0.24139 | 0.040238 | -0.0007761 | 4.757e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 141 | 0.36816 | [20 empty] | -0.091942 | 0.38583 | 0.085365 | -0.012327 | 0.0013281 | -0.00010606 | 6.2737e-06 | -2.3333e-07 | 3.8481e-09 | [12 empty] | 1.2716 | 0.24312 | 0.040815 | -0.00079285 | 4.8943e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 142 | 0.37188 | [20 empty] | -0.091942 | 0.38857 | 0.08658 | -0.012591 | 0.0013662 | -0.00010987 | 6.5455e-06 | -2.4517e-07 | 4.072e-09 | [12 empty] | 1.2716 | 0.24484 | 0.041396 | -0.00080983 | 5.0347e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 143 | 0.3756 | [20 empty] | -0.091942 | 0.3913 | 0.087804 | -0.012859 | 0.0014051 | -0.00011379 | 6.827e-06 | -2.5751e-07 | 4.3071e-09 | [12 empty] | 1.2716 | 0.24657 | 0.041981 | -0.00082706 | 5.178e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 144 | 0.37934 | [20 empty] | -0.091942 | 0.39404 | 0.089036 | -0.01313 | 0.0014448 | -0.00011783 | 7.1185e-06 | -2.7038e-07 | 4.5541e-09 | [12 empty] | 1.2716 | 0.24829 | 0.042571 | -0.00084454 | 5.3244e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 145 | 0.38308 | [20 empty] | -0.091942 | 0.39678 | 0.090277 | -0.013406 | 0.0014853 | -0.00012198 | 7.4203e-06 | -2.8381e-07 | 4.8133e-09 | [12 empty] | 1.2716 | 0.25001 | 0.043164 | -0.00086225 | 5.4738e-07
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 146 | 0.38682 | [20 empty] | -0.091942 | 0.39951 | 0.091527 | -0.013685 | 0.0015267 | -0.00012624 | 7.7327e-06 | -2.9779e-07 | 5.0854e-09 | [12 empty] | 1.2716 | 0.25174 | 0.043761 | -0.00088022 | 5.6264e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 147 | 0.39057 | [20 empty] | -0.091942 | 0.40225 | 0.092785 | -0.013968 | 0.001569 | -0.00013062 | 8.0559e-06 | -3.1237e-07 | 5.3708e-09 | [12 empty] | 1.2716 | 0.25346 | 0.044363 | -0.00089843 | 5.7821e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 148 | 0.39432 | [20 empty] | -0.091942 | 0.40499 | 0.094052 | -0.014255 | 0.0016121 | -0.00013513 | 8.3904e-06 | -3.2755e-07 | 5.6701e-09 | [12 empty] | 1.2716 | 0.25519 | 0.044969 | -0.00091689 | 5.9411e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 149 | 0.39809 | [20 empty] | -0.091942 | 0.40772 | 0.095327 | -0.014546 | 0.0016561 | -0.00013976 | 8.7364e-06 | -3.4336e-07 | 5.984e-09 | [12 empty] | 1.2716 | 0.25691 | 0.045578 | -0.0009356 | 6.1033e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 150 | 0.40185 | [20 empty] | -0.091942 | 0.41046 | 0.096611 | -0.014841 | 0.001701 | -0.00014451 | 9.0941e-06 | -3.5982e-07 | 6.3129e-09 | [12 empty] | 1.2716 | 0.25864 | 0.046192 | -0.00095456 | 6.2688e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 151 | 0.40562 | [20 empty] | -0.091942 | 0.41319 | 0.097903 | -0.01514 | 0.0017469 | -0.00014939 | 9.464e-06 | -3.7695e-07 | 6.6576e-09 | [12 empty] | 1.2716 | 0.26036 | 0.04681 | -0.00097378 | 6.4376e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 152 | 0.4094 | [20 empty] | -0.091942 | 0.41593 | 0.099204 | -0.015443 | 0.0017936 | -0.0001544 | 9.8463e-06 | -3.9477e-07 | 7.0186e-09 | [12 empty] | 1.2716 | 0.26208 | 0.047432 | -0.00099326 | 6.6099e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 153 | 0.41318 | [20 empty] | -0.091942 | 0.41867 | 0.10051 | -0.01575 | 0.0018413 | -0.00015955 | 1.0241e-05 | -4.1332e-07 | 7.3966e-09 | [12 empty] | 1.2716 | 0.26381 | 0.048058 | -0.001013 | 6.7855e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 154 | 0.41697 | [20 empty] | -0.091942 | 0.4214 | 0.10183 | -0.01606 | 0.0018899 | -0.00016483 | 1.065e-05 | -4.326e-07 | 7.7923e-09 | [12 empty] | 1.2716 | 0.26553 | 0.048689 | -0.001033 | 6.9647e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 155 | 0.42076 | [20 empty] | -0.091942 | 0.42414 | 0.10316 | -0.016375 | 0.0019395 | -0.00017025 | 1.1071e-05 | -4.5265e-07 | 8.2064e-09 | [12 empty] | 1.2716 | 0.26726 | 0.049323 | -0.0010532 | 7.1474e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 156 | 0.42456 | [20 empty] | -0.091942 | 0.42688 | 0.10449 | -0.016694 | 0.00199 | -0.00017582 | 1.1507e-05 | -4.735e-07 | 8.6396e-09 | [12 empty] | 1.2716 | 0.26898 | 0.049961 | -0.0010738 | 7.3336e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 157 | 0.42836 | [20 empty] | -0.091942 | 0.42961 | 0.10584 | -0.017017 | 0.0020415 | -0.00018152 | 1.1957e-05 | -4.9516e-07 | 9.0928e-09 | [12 empty] | 1.2716 | 0.2707 | 0.050604 | -0.0010945 | 7.5235e-07
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 158 | 0.43217 | [20 empty] | -0.091942 | 0.43235 | 0.10719 | -0.017345 | 0.002094 | -0.00018738 | 1.2421e-05 | -5.1766e-07 | 9.5666e-09 | [12 empty] | 1.2716 | 0.27243 | 0.051251 | -0.0011156 | 7.717e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 159 | 0.43599 | [20 empty] | -0.091942 | 0.43509 | 0.10855 | -0.017676 | 0.0021475 | -0.00019338 | 1.29e-05 | -5.4103e-07 | 1.0062e-08 | [12 empty] | 1.2716 | 0.27415 | 0.051901 | -0.0011369 | 7.9142e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 160 | 0.43981 | [20 empty] | -0.091942 | 0.43782 | 0.10992 | -0.018012 | 0.0022021 | -0.00019954 | 1.3395e-05 | -5.6531e-07 | 1.0579e-08 | [12 empty] | 1.2716 | 0.27588 | 0.052556 | -0.0011585 | 8.1152e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 161 | 0.44363 | [20 empty] | -0.091942 | 0.44056 | 0.1113 | -0.018351 | 0.0022576 | -0.00020586 | 1.3905e-05 | -5.9051e-07 | 1.112e-08 | [12 empty] | 1.2716 | 0.2776 | 0.053215 | -0.0011803 | 8.32e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 162 | 0.44746 | [20 empty] | -0.091942 | 0.44329 | 0.11269 | -0.018696 | 0.0023143 | -0.00021233 | 1.4431e-05 | -6.1667e-07 | 1.1685e-08 | [12 empty] | 1.2716 | 0.27933 | 0.053879 | -0.0012025 | 8.5286e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 163 | 0.4513 | [20 empty] | -0.091942 | 0.44603 | 0.11408 | -0.019044 | 0.0023719 | -0.00021897 | 1.4974e-05 | -6.4381e-07 | 1.2274e-08 | [12 empty] | 1.2716 | 0.28105 | 0.054546 | -0.0012249 | 8.7412e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 164 | 0.45514 | [20 empty] | -0.091942 | 0.44877 | 0.11549 | -0.019397 | 0.0024307 | -0.00022576 | 1.5534e-05 | -6.7197e-07 | 1.289e-08 | [12 empty] | 1.2716 | 0.28277 | 0.055217 | -0.0012476 | 8.9576e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 165 | 0.45898 | [20 empty] | -0.091942 | 0.4515 | 0.1169 | -0.019754 | 0.0024905 | -0.00023273 | 1.6111e-05 | -7.0118e-07 | 1.3532e-08 | [12 empty] | 1.2716 | 0.2845 | 0.055892 | -0.0012705 | 9.1781e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 166 | 0.46283 | [20 empty] | -0.091942 | 0.45424 | 0.11832 | -0.020115 | 0.0025514 | -0.00023987 | 1.6706e-05 | -7.3148e-07 | 1.4202e-08 | [12 empty] | 1.2716 | 0.28622 | 0.056572 | -0.0012938 | 9.4027e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 167 | 0.46669 | [20 empty] | -0.091942 | 0.45698 | 0.11975 | -0.020481 | 0.0026135 | -0.00024718 | 1.7319e-05 | -7.6289e-07 | 1.4902e-08 | [12 empty] | 1.2716 | 0.28795 | 0.057256 | -0.0013173 | 9.6313e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 168 | 0.47055 | [20 empty] | -0.091942 | 0.45971 | 0.12119 | -0.020851 | 0.0026766 | -0.00025467 | 1.795e-05 | -7.9544e-07 | 1.5631e-08 | [12 empty] | 1.2716 | 0.28967 | 0.057943 | -0.0013411 | 9.8641e-07
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 169 | 0.47441 | [20 empty] | -0.091942 | 0.46245 | 0.12264 | -0.021225 | 0.0027409 | -0.00026234 | 1.8601e-05 | -8.2918e-07 | 1.6391e-08 | [12 empty] | 1.2716 | 0.2914 | 0.058635 | -0.0013652 | 1.0101e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 170 | 0.47829 | [20 empty] | -0.091942 | 0.46519 | 0.12409 | -0.021604 | 0.0028064 | -0.0002702 | 1.9271e-05 | -8.6415e-07 | 1.7183e-08 | [12 empty] | 1.2716 | 0.29312 | 0.059331 | -0.0013896 | 1.0342e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 171 | 0.48216 | [20 empty] | -0.091942 | 0.46792 | 0.12556 | -0.021988 | 0.002873 | -0.00027824 | 1.9961e-05 | -9.0036e-07 | 1.8008e-08 | [12 empty] | 1.2716 | 0.29484 | 0.060031 | -0.0014142 | 1.0588e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 172 | 0.48604 | [20 empty] | -0.091942 | 0.47066 | 0.12703 | -0.022376 | 0.0029408 | -0.00028647 | 2.0672e-05 | -9.3787e-07 | 1.8868e-08 | [12 empty] | 1.2716 | 0.29657 | 0.060735 | -0.0014392 | 1.0838e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 173 | 0.48993 | [20 empty] | -0.091942 | 0.47339 | 0.12851 | -0.022768 | 0.0030098 | -0.0002949 | 2.1404e-05 | -9.7671e-07 | 1.9764e-08 | [12 empty] | 1.2716 | 0.29829 | 0.061444 | -0.0014644 | 1.1092e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 174 | 0.49382 | [20 empty] | -0.091942 | 0.47613 | 0.13 | -0.023165 | 0.00308 | -0.00030352 | 2.2157e-05 | -1.0169e-06 | 2.0696e-08 | [12 empty] | 1.2716 | 0.30002 | 0.062156 | -0.00149 | 1.1351e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 175 | 0.49772 | [20 empty] | -0.091942 | 0.47887 | 0.1315 | -0.023567 | 0.0031514 | -0.00031234 | 2.2932e-05 | -1.0585e-06 | 2.1667e-08 | [12 empty] | 1.2716 | 0.30174 | 0.062873 | -0.0015158 | 1.1614e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 176 | 0.50162 | [20 empty] | -0.091942 | 0.4816 | 0.13301 | -0.023973 | 0.0032241 | -0.00032137 | 2.3729e-05 | -1.1016e-06 | 2.2678e-08 | [12 empty] | 1.2716 | 0.30347 | 0.063593 | -0.0015419 | 1.1881e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 177 | 0.50553 | [20 empty] | -0.091942 | 0.48434 | 0.13452 | -0.024384 | 0.003298 | -0.0003306 | 2.455e-05 | -1.1462e-06 | 2.3729e-08 | [12 empty] | 1.2716 | 0.30519 | 0.064318 | -0.0015684 | 1.2154e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 178 | 0.50944 | [20 empty] | -0.091942 | 0.48708 | 0.13605 | -0.0248 | 0.0033731 | -0.00034005 | 2.5394e-05 | -1.1923e-06 | 2.4823e-08 | [12 empty] | 1.2716 | 0.30691 | 0.065047 | -0.0015951 | 1.2431e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 179 | 0.51335 | [20 empty] | -0.091942 | 0.48981 | 0.13758 | -0.02522 | 0.0034496 | -0.00034971 | 2.6262e-05 | -1.24e-06 | 2.5961e-08 | [12 empty] | 1.2716 | 0.30864 | 0.06578 | -0.0016221 | 1.2712e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 180 | 0.51728 | [20 empty] | -0.091942 | 0.49255 | 0.13912 | -0.025645 | 0.0035273 | -0.00035958 | 2.7155e-05 | -1.2893e-06 | 2.7144e-08 | [12 empty] | 1.2716 | 0.31036 | 0.066517 | -0.0016495 | 1.2999e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 181 | 0.5212 | [20 empty] | -0.091942 | 0.49529 | 0.14067 | -0.026075 | 0.0036063 | -0.00036968 | 2.8073e-05 | -1.3403e-06 | 2.8374e-08 | [12 empty] | 1.2716 | 0.31209 | 0.067258 | -0.0016771 | 1.329e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 182 | 0.52513 | [20 empty] | -0.091942 | 0.49802 | 0.14223 | -0.02651 | 0.0036867 | -0.00038001 | 2.9016e-05 | -1.393e-06 | 2.9653e-08 | [12 empty] | 1.2716 | 0.31381 | 0.068003 | -0.0017051 | 1.3586e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 183 | 0.52907 | [20 empty] | -0.091942 | 0.50076 | 0.1438 | -0.026949 | 0.0037684 | -0.00039056 | 2.9986e-05 | -1.4474e-06 | 3.0982e-08 | [12 empty] | 1.2716 | 0.31554 | 0.068752 | -0.0017333 | 1.3887e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 184 | 0.53301 | [20 empty] | -0.091942 | 0.5035 | 0.14537 | -0.027393 | 0.0038514 | -0.00040135 | 3.0983e-05 | -1.5037e-06 | 3.2362e-08 | [12 empty] | 1.2716 | 0.31726 | 0.069506 | -0.0017619 | 1.4194e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 185 | 0.53696 | [20 empty] | -0.091942 | 0.50623 | 0.14696 | -0.027842 | 0.0039359 | -0.00041238 | 3.2007e-05 | -1.5619e-06 | 3.3797e-08 | [12 empty] | 1.2716 | 0.31898 | 0.070263 | -0.0017908 | 1.4505e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 186 | 0.54091 | [20 empty] | -0.091942 | 0.50897 | 0.14855 | -0.028296 | 0.0040216 | -0.00042364 | 3.3059e-05 | -1.6219e-06 | 3.5286e-08 | [12 empty] | 1.2716 | 0.32071 | 0.071025 | -0.00182 | 1.4821e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 187 | 0.54486 | [20 empty] | -0.091942 | 0.5117 | 0.15015 | -0.028755 | 0.0041088 | -0.00043516 | 3.414e-05 | -1.684e-06 | 3.6833e-08 | [12 empty] | 1.2716 | 0.32243 | 0.071791 | -0.0018495 | 1.5142e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 188 | 0.54882 | [20 empty] | -0.091942 | 0.51444 | 0.15176 | -0.029219 | 0.0041974 | -0.00044692 | 3.525e-05 | -1.748e-06 | 3.8438e-08 | [12 empty] | 1.2716 | 0.32416 | 0.072561 | -0.0018793 | 1.5469e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 189 | 0.55279 | [20 empty] | -0.091942 | 0.51718 | 0.15338 | -0.029688 | 0.0042875 | -0.00045893 | 3.639e-05 | -1.8142e-06 | 4.0105e-08 | [12 empty] | 1.2716 | 0.32588 | 0.073335 | -0.0019095 | 1.58e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 190 | 0.55676 | [20 empty] | -0.091942 | 0.51991 | 0.15501 | -0.030162 | 0.0043789 | -0.0004712 | 3.7561e-05 | -1.8824e-06 | 4.1834e-08 | [12 empty] | 1.2716 | 0.3276 | 0.074113 | -0.00194 | 1.6137e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 191 | 0.56074 | [20 empty] | -0.091942 | 0.52265 | 0.15664 | -0.03064 | 0.0044718 | -0.00048373 | 3.8763e-05 | -1.9529e-06 | 4.3628e-08 | [12 empty] | 1.2716 | 0.32933 | 0.074895 | -0.0019707 | 1.648e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 192 | 0.56472 | [20 empty] | -0.091942 | 0.52539 | 0.15829 | -0.031124 | 0.0045662 | -0.00049653 | 3.9996e-05 | -2.0256e-06 | 4.5489e-08 | [12 empty] | 1.2716 | 0.33105 | 0.075681 | -0.0020019 | 1.6828e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 193 | 0.5687 | [20 empty] | -0.091942 | 0.52812 | 0.15994 | -0.031613 | 0.0046621 | -0.00050959 | 4.1263e-05 | -2.1006e-06 | 4.742e-08 | [12 empty] | 1.2716 | 0.33278 | 0.076472 | -0.0020333 | 1.7181e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 194 | 0.57269 | [20 empty] | -0.091942 | 0.53086 | 0.1616 | -0.032107 | 0.0047595 | -0.00052293 | 4.2562e-05 | -2.178e-06 | 4.9421e-08 | [12 empty] | 1.2716 | 0.3345 | 0.077266 | -0.0020651 | 1.754e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 195 | 0.57668 | [20 empty] | -0.091942 | 0.5336 | 0.16327 | -0.032606 | 0.0048584 | -0.00053655 | 4.3896e-05 | -2.2578e-06 | 5.1496e-08 | [12 empty] | 1.2716 | 0.33623 | 0.078065 | -0.0020972 | 1.7904e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 196 | 0.58068 | [20 empty] | -0.091942 | 0.53633 | 0.16495 | -0.03311 | 0.0049588 | -0.00055045 | 4.5264e-05 | -2.3401e-06 | 5.3647e-08 | [12 empty] | 1.2716 | 0.33795 | 0.078867 | -0.0021296 | 1.8274e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 197 | 0.58469 | [20 empty] | -0.091942 | 0.53907 | 0.16664 | -0.03362 | 0.0050608 | -0.00056464 | 4.6667e-05 | -2.425e-06 | 5.5876e-08 | [12 empty] | 1.2716 | 0.33967 | 0.079674 | -0.0021624 | 1.865e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 198 | 0.58869 | [20 empty] | -0.091942 | 0.5418 | 0.16833 | -0.034134 | 0.0051643 | -0.00057911 | 4.8106e-05 | -2.5125e-06 | 5.8186e-08 | [12 empty] | 1.2716 | 0.3414 | 0.080485 | -0.0021955 | 1.9032e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 199 | 0.59271 | [20 empty] | -0.091942 | 0.54454 | 0.17004 | -0.034654 | 0.0052694 | -0.00059388 | 4.9583e-05 | -2.6027e-06 | 6.0579e-08 | [12 empty] | 1.2716 | 0.34312 | 0.0813 | -0.0022289 | 1.9419e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 200 | 0.59672 | [20 empty] | -0.091942 | 0.54728 | 0.17175 | -0.035179 | 0.0053762 | -0.00060896 | 5.1097e-05 | -2.6956e-06 | 6.3058e-08 | [12 empty] | 1.2716 | 0.34485 | 0.082119 | -0.0022627 | 1.9812e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 201 | 0.60075 | [20 empty] | -0.091942 | 0.55001 | 0.17347 | -0.035709 | 0.0054845 | -0.00062433 | 5.2649e-05 | -2.7914e-06 | 6.5625e-08 | [12 empty] | 1.2716 | 0.34657 | 0.082943 | -0.0022968 | 2.0212e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 202 | 0.60477 | [20 empty] | -0.091942 | 0.55275 | 0.1752 | -0.036245 | 0.0055945 | -0.00064002 | 5.424e-05 | -2.89e-06 | 6.8283e-08 | [12 empty] | 1.2716 | 0.3483 | 0.08377 | -0.0023312 | 2.0617e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 203 | 0.60881 | [20 empty] | -0.091942 | 0.55549 | 0.17694 | -0.036786 | 0.0057061 | -0.00065602 | 5.5871e-05 | -2.9917e-06 | 7.1034e-08 | [12 empty] | 1.2716 | 0.35002 | 0.084601 | -0.002366 | 2.1028e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 204 | 0.61284 | [20 empty] | -0.091942 | 0.55822 | 0.17869 | -0.037332 | 0.0058193 | -0.00067234 | 5.7543e-05 | -3.0964e-06 | 7.3882e-08 | [12 empty] | 1.2716 | 0.35174 | 0.085437 | -0.0024012 | 2.1446e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 205 | 0.61688 | [20 empty] | -0.091942 | 0.56096 | 0.18045 | -0.037884 | 0.0059343 | -0.00068898 | 5.9256e-05 | -3.2042e-06 | 7.683e-08 | [12 empty] | 1.2716 | 0.35347 | 0.086277 | -0.0024366 | 2.1869e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 206 | 0.62093 | [20 empty] | -0.091942 | 0.5637 | 0.18221 | -0.038441 | 0.0060509 | -0.00070595 | 6.1012e-05 | -3.3152e-06 | 7.988e-08 | [12 empty] | 1.2716 | 0.35519 | 0.08712 | -0.0024725 | 2.2299e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 207 | 0.62498 | [20 empty] | -0.091942 | 0.56643 | 0.18399 | -0.039004 | 0.0061693 | -0.00072325 | 6.2811e-05 | -3.4295e-06 | 8.3035e-08 | [12 empty] | 1.2716 | 0.35692 | 0.087968 | -0.0025087 | 2.2735e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 208 | 0.62903 | [20 empty] | -0.091942 | 0.56917 | 0.18577 | -0.039572 | 0.0062893 | -0.00074089 | 6.4654e-05 | -3.5472e-06 | 8.6299e-08 | [12 empty] | 1.2716 | 0.35864 | 0.08882 | -0.0025452 | 2.3178e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 209 | 0.63309 | [20 empty] | -0.091942 | 0.5719 | 0.18756 | -0.040145 | 0.0064112 | -0.00075887 | 6.6541e-05 | -3.6683e-06 | 8.9675e-08 | [12 empty] | 1.2716 | 0.36037 | 0.089676 | -0.0025821 | 2.3627e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 210 | 0.63715 | [20 empty] | -0.091942 | 0.57464 | 0.18936 | -0.040724 | 0.0065348 | -0.0007772 | 6.8474e-05 | -3.793e-06 | 9.3165e-08 | [12 empty] | 1.2716 | 0.36209 | 0.090537 | -0.0026193 | 2.4082e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 211 | 0.64122 | [20 empty] | -0.091942 | 0.57738 | 0.19116 | -0.041309 | 0.0066601 | -0.00079588 | 7.0454e-05 | -3.9212e-06 | 9.6774e-08 | [12 empty] | 1.2716 | 0.36381 | 0.091401 | -0.0026569 | 2.4544e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 212 | 0.6453 | [20 empty] | -0.091942 | 0.58011 | 0.19298 | -0.041899 | 0.0067873 | -0.00081492 | 7.2482e-05 | -4.0532e-06 | 1.005e-07 | [12 empty] | 1.2716 | 0.36554 | 0.092269 | -0.0026949 | 2.5013e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 213 | 0.64937 | [20 empty] | -0.091942 | 0.58285 | 0.19481 | -0.042494 | 0.0069162 | -0.00083432 | 7.4557e-05 | -4.1889e-06 | 1.0436e-07 | [12 empty] | 1.2716 | 0.36726 | 0.093142 | -0.0027332 | 2.5488e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 214 | 0.65345 | [20 empty] | -0.091942 | 0.58559 | 0.19664 | -0.043096 | 0.007047 | -0.00085409 | 7.6682e-05 | -4.3285e-06 | 1.0835e-07 | [12 empty] | 1.2716 | 0.36899 | 0.094018 | -0.0027719 | 2.597e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 215 | 0.65754 | [20 empty] | -0.091942 | 0.58832 | 0.19848 | -0.043703 | 0.0071797 | -0.00087424 | 7.8858e-05 | -4.4721e-06 | 1.1246e-07 | [12 empty] | 1.2716 | 0.37071 | 0.094899 | -0.0028109 | 2.6459e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 216 | 0.66163 | [20 empty] | -0.091942 | 0.59106 | 0.20033 | -0.044315 | 0.0073142 | -0.00089476 | 8.1084e-05 | -4.6198e-06 | 1.1672e-07 | [12 empty] | 1.2716 | 0.37243 | 0.095784 | -0.0028503 | 2.6955e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 217 | 0.66573 | [20 empty] | -0.091942 | 0.5938 | 0.20219 | -0.044934 | 0.0074506 | -0.00091566 | 8.3363e-05 | -4.7716e-06 | 1.2111e-07 | [12 empty] | 1.2716 | 0.37416 | 0.096673 | -0.0028901 | 2.7457e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 218 | 0.66982 | [20 empty] | -0.091942 | 0.59653 | 0.20406 | -0.045558 | 0.0075889 | -0.00093696 | 8.5694e-05 | -4.9276e-06 | 1.2565e-07 | [12 empty] | 1.2716 | 0.37588 | 0.097566 | -0.0029302 | 2.7967e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 219 | 0.67393 | [20 empty] | -0.091942 | 0.59927 | 0.20594 | -0.046188 | 0.0077291 | -0.00095864 | 8.808e-05 | -5.0881e-06 | 1.3033e-07 | [12 empty] | 1.2716 | 0.37761 | 0.098463 | -0.0029707 | 2.8484e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 220 | 0.67804 | [20 empty] | -0.091942 | 0.60201 | 0.20782 | -0.046823 | 0.0078712 | -0.00098073 | 9.0521e-05 | -5.2529e-06 | 1.3517e-07 | [12 empty] | 1.2716 | 0.37933 | 0.099364 | -0.0030116 | 2.9007e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 221 | 0.68215 | [20 empty] | -0.091942 | 0.60474 | 0.20971 | -0.047465 | 0.0080153 | -0.0010032 | 9.3018e-05 | -5.4224e-06 | 1.4016e-07 | [12 empty] | 1.2716 | 0.38106 | 0.10027 | -0.0030529 | 2.9538e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 222 | 0.68627 | [20 empty] | -0.091942 | 0.60748 | 0.21162 | -0.048112 | 0.0081614 | -0.0010261 | 9.5572e-05 | -5.5965e-06 | 1.4532e-07 | [12 empty] | 1.2716 | 0.38278 | 0.10118 | -0.0030945 | 3.0077e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 223 | 0.69039 | [20 empty] | -0.091942 | 0.61021 | 0.21353 | -0.048765 | 0.0083094 | -0.0010494 | 9.8184e-05 | -5.7754e-06 | 1.5064e-07 | [12 empty] | 1.2716 | 0.3845 | 0.10209 | -0.0031365 | 3.0622e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 224 | 0.69451 | [20 empty] | -0.091942 | 0.61295 | 0.21545 | -0.049424 | 0.0084595 | -0.0010732 | 0.00010086 | -5.9591e-06 | 1.5613e-07 | [12 empty] | 1.2716 | 0.38623 | 0.10301 | -0.0031789 | 3.1175e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 225 | 0.69864 | [20 empty] | -0.091942 | 0.61569 | 0.21737 | -0.050089 | 0.0086116 | -0.0010974 | 0.00010359 | -6.1478e-06 | 1.6179e-07 | [12 empty] | 1.2716 | 0.38795 | 0.10393 | -0.0032217 | 3.1736e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 226 | 0.70278 | [20 empty] | -0.091942 | 0.61842 | 0.21931 | -0.05076 | 0.0087657 | -0.001122 | 0.00010638 | -6.3417e-06 | 1.6764e-07 | [12 empty] | 1.2716 | 0.38968 | 0.10486 | -0.0032648 | 3.2304e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 227 | 0.70692 | [20 empty] | -0.091942 | 0.62116 | 0.22126 | -0.051436 | 0.0089219 | -0.001147 | 0.00010924 | -6.5407e-06 | 1.7366e-07 | [12 empty] | 1.2716 | 0.3914 | 0.10579 | -0.0033083 | 3.2879e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 228 | 0.71106 | [20 empty] | -0.091942 | 0.6239 | 0.22321 | -0.052119 | 0.0090801 | -0.0011725 | 0.00011216 | -6.7451e-06 | 1.7988e-07 | [12 empty] | 1.2716 | 0.39313 | 0.10672 | -0.0033522 | 3.3462e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 229 | 0.71521 | [20 empty] | -0.091942 | 0.62663 | 0.22517 | -0.052808 | 0.0092405 | -0.0011984 | 0.00011514 | -6.9549e-06 | 1.8629e-07 | [12 empty] | 1.2716 | 0.39485 | 0.10766 | -0.0033965 | 3.4053e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 230 | 0.71936 | [20 empty] | -0.091942 | 0.62937 | 0.22714 | -0.053503 | 0.0094029 | -0.0012248 | 0.00011819 | -7.1703e-06 | 1.929e-07 | [12 empty] | 1.2716 | 0.39657 | 0.1086 | -0.0034412 | 3.4652e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 231 | 0.72351 | [20 empty] | -0.091942 | 0.63211 | 0.22912 | -0.054204 | 0.0095675 | -0.0012517 | 0.00012131 | -7.3914e-06 | 1.9971e-07 | [12 empty] | 1.2716 | 0.3983 | 0.10955 | -0.0034863 | 3.5259e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 232 | 0.72767 | [20 empty] | -0.091942 | 0.63484 | 0.23111 | -0.054911 | 0.0097343 | -0.001279 | 0.00012449 | -7.6183e-06 | 2.0673e-07 | [12 empty] | 1.2716 | 0.40002 | 0.1105 | -0.0035318 | 3.5873e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 233 | 0.73184 | [20 empty] | -0.091942 | 0.63758 | 0.23311 | -0.055624 | 0.0099032 | -0.0013068 | 0.00012775 | -7.8512e-06 | 2.1397e-07 | [12 empty] | 1.2716 | 0.40175 | 0.11145 | -0.0035777 | 3.6496e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 234 | 0.736 | [20 empty] | -0.091942 | 0.64031 | 0.23511 | -0.056343 | 0.010074 | -0.0013351 | 0.00013107 | -8.0901e-06 | 2.2142e-07 | [12 empty] | 1.2716 | 0.40347 | 0.11241 | -0.0036239 | 3.7126e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 235 | 0.74018 | [20 empty] | -0.091942 | 0.64305 | 0.23713 | -0.057069 | 0.010248 | -0.0013639 | 0.00013447 | -8.3353e-06 | 2.2911e-07 | [12 empty] | 1.2716 | 0.4052 | 0.11338 | -0.0036706 | 3.7765e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 236 | 0.74435 | [20 empty] | -0.091942 | 0.64579 | 0.23915 | -0.0578 | 0.010423 | -0.0013931 | 0.00013794 | -8.5867e-06 | 2.3703e-07 | [12 empty] | 1.2716 | 0.40692 | 0.11434 | -0.0037176 | 3.8412e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 237 | 0.74853 | [20 empty] | -0.091942 | 0.64852 | 0.24118 | -0.058538 | 0.010601 | -0.0014229 | 0.00014148 | -8.8447e-06 | 2.4518e-07 | [12 empty] | 1.2716 | 0.40864 | 0.11531 | -0.0037651 | 3.9067e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 238 | 0.75272 | [20 empty] | -0.091942 | 0.65126 | 0.24322 | -0.059282 | 0.010781 | -0.0014532 | 0.0001451 | -9.1093e-06 | 2.5358e-07 | [12 empty] | 1.2716 | 0.41037 | 0.11629 | -0.003813 | 3.9731e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 239 | 0.75691 | [20 empty] | -0.091942 | 0.654 | 0.24527 | -0.060033 | 0.010963 | -0.001484 | 0.0001488 | -9.3806e-06 | 2.6223e-07 | [12 empty] | 1.2716 | 0.41209 | 0.11727 | -0.0038612 | 4.0403e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 240 | 0.7611 | [20 empty] | -0.091942 | 0.65673 | 0.24732 | -0.060789 | 0.011148 | -0.0015153 | 0.00015257 | -9.6588e-06 | 2.7114e-07 | [12 empty] | 1.2716 | 0.41382 | 0.11825 | -0.0039099 | 4.1083e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 241 | 0.7653 | [20 empty] | -0.091942 | 0.65947 | 0.24939 | -0.061552 | 0.011335 | -0.0015471 | 0.00015643 | -9.9441e-06 | 2.8031e-07 | [12 empty] | 1.2716 | 0.41554 | 0.11924 | -0.003959 | 4.1772e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 242 | 0.7695 | [20 empty] | -0.091942 | 0.66221 | 0.25146 | -0.062322 | 0.011524 | -0.0015795 | 0.00016036 | -1.0237e-05 | 2.8975e-07 | [12 empty] | 1.2716 | 0.41726 | 0.12023 | -0.0040085 | 4.247e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 243 | 0.77371 | [20 empty] | -0.091942 | 0.66494 | 0.25355 | -0.063097 | 0.011716 | -0.0016124 | 0.00016438 | -1.0536e-05 | 2.9947e-07 | [12 empty] | 1.2716 | 0.41899 | 0.12123 | -0.0040584 | 4.3176e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 244 | 0.77792 | [20 empty] | -0.091942 | 0.66768 | 0.25564 | -0.06388 | 0.01191 | -0.0016458 | 0.00016848 | -1.0844e-05 | 3.0947e-07 | [12 empty] | 1.2716 | 0.42071 | 0.12223 | -0.0041087 | 4.3891e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 245 | 0.78213 | [20 empty] | -0.091942 | 0.67041 | 0.25774 | -0.064668 | 0.012106 | -0.0016798 | 0.00017267 | -1.1159e-05 | 3.1976e-07 | [12 empty] | 1.2716 | 0.42244 | 0.12323 | -0.0041594 | 4.4615e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 246 | 0.78635 | [20 empty] | -0.091942 | 0.67315 | 0.25984 | -0.065463 | 0.012305 | -0.0017144 | 0.00017694 | -1.1481e-05 | 3.3035e-07 | [12 empty] | 1.2716 | 0.42416 | 0.12424 | -0.0042105 | 4.5348e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 247 | 0.79057 | [20 empty] | -0.091942 | 0.67589 | 0.26196 | -0.066265 | 0.012507 | -0.0017495 | 0.0001813 | -1.1812e-05 | 3.4125e-07 | [12 empty] | 1.2716 | 0.42589 | 0.12525 | -0.0042621 | 4.609e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 248 | 0.79479 | [20 empty] | -0.091942 | 0.67862 | 0.26409 | -0.067073 | 0.01271 | -0.0017852 | 0.00018575 | -1.2151e-05 | 3.5246e-07 | [12 empty] | 1.2716 | 0.42761 | 0.12627 | -0.0043141 | 4.6841e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 249 | 0.79902 | [20 empty] | -0.091942 | 0.68136 | 0.26622 | -0.067888 | 0.012917 | -0.0018215 | 0.00019029 | -1.2498e-05 | 3.6399e-07 | [12 empty] | 1.2716 | 0.42933 | 0.12729 | -0.0043665 | 4.7601e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 250 | 0.80326 | [20 empty] | -0.091942 | 0.6841 | 0.26836 | -0.068709 | 0.013125 | -0.0018584 | 0.00019492 | -1.2854e-05 | 3.7585e-07 | [12 empty] | 1.2716 | 0.43106 | 0.12831 | -0.0044193 | 4.837e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 251 | 0.8075 | [20 empty] | -0.091942 | 0.68683 | 0.27051 | -0.069537 | 0.013337 | -0.0018959 | 0.00019964 | -1.3218e-05 | 3.8805e-07 | [12 empty] | 1.2716 | 0.43278 | 0.12934 | -0.0044725 | 4.9149e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 252 | 0.81174 | [20 empty] | -0.091942 | 0.68957 | 0.27267 | -0.070371 | 0.01355 | -0.0019339 | 0.00020446 | -1.3591e-05 | 4.0059e-07 | [12 empty] | 1.2716 | 0.43451 | 0.13037 | -0.0045262 | 4.9937e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 253 | 0.81598 | [20 empty] | -0.091942 | 0.69231 | 0.27484 | -0.071212 | 0.013767 | -0.0019726 | 0.00020938 | -1.3973e-05 | 4.1349e-07 | [12 empty] | 1.2716 | 0.43623 | 0.13141 | -0.0045803 | 5.0734e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 254 | 0.82023 | [20 empty] | -0.091942 | 0.69504 | 0.27702 | -0.07206 | 0.013986 | -0.0020119 | 0.0002144 | -1.4364e-05 | 4.2675e-07 | [12 empty] | 1.2716 | 0.43796 | 0.13245 | -0.0046348 | 5.1541e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 255 | 0.82449 | [20 empty] | -0.091942 | 0.69778 | 0.27921 | -0.072914 | 0.014207 | -0.0020518 | 0.00021951 | -1.4765e-05 | 4.4037e-07 | [12 empty] | 1.2716 | 0.43968 | 0.1335 | -0.0046898 | 5.2358e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 256 | 0.82875 | [20 empty] | -0.091942 | 0.70052 | 0.2814 | -0.073776 | 0.014432 | -0.0020924 | 0.00022473 | -1.5175e-05 | 4.5438e-07 | [12 empty] | 1.2716 | 0.4414 | 0.13454 | -0.0047452 | 5.3184e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 257 | 0.83301 | [20 empty] | -0.091942 | 0.70325 | 0.2836 | -0.074644 | 0.014658 | -0.0021335 | 0.00023004 | -1.5595e-05 | 4.6877e-07 | [12 empty] | 1.2716 | 0.44313 | 0.1356 | -0.004801 | 5.402e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 258 | 0.83727 | [20 empty] | -0.091942 | 0.70599 | 0.28581 | -0.075518 | 0.014888 | -0.0021754 | 0.00023547 | -1.6024e-05 | 4.8357e-07 | [12 empty] | 1.2716 | 0.44485 | 0.13665 | -0.0048572 | 5.4865e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 259 | 0.84154 | [20 empty] | -0.091942 | 0.70872 | 0.28803 | -0.0764 | 0.01512 | -0.0022179 | 0.000241 | -1.6464e-05 | 4.9877e-07 | [12 empty] | 1.2716 | 0.44658 | 0.13772 | -0.0049139 | 5.5721e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 260 | 0.84582 | [20 empty] | -0.091942 | 0.71146 | 0.29026 | -0.077288 | 0.015355 | -0.002261 | 0.00024663 | -1.6914e-05 | 5.1438e-07 | [12 empty] | 1.2716 | 0.4483 | 0.13878 | -0.0049711 | 5.6586e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 261 | 0.85009 | [20 empty] | -0.091942 | 0.7142 | 0.2925 | -0.078183 | 0.015592 | -0.0023048 | 0.00025238 | -1.7375e-05 | 5.3042e-07 | [12 empty] | 1.2716 | 0.45003 | 0.13985 | -0.0050287 | 5.7462e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 262 | 0.85438 | [20 empty] | -0.091942 | 0.71693 | 0.29474 | -0.079085 | 0.015833 | -0.0023493 | 0.00025824 | -1.7847e-05 | 5.469e-07 | [12 empty] | 1.2716 | 0.45175 | 0.14093 | -0.0050867 | 5.8348e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 263 | 0.85866 | [20 empty] | -0.091942 | 0.71967 | 0.297 | -0.079995 | 0.016076 | -0.0023945 | 0.00026421 | -1.8329e-05 | 5.6383e-07 | [12 empty] | 1.2716 | 0.45347 | 0.142 | -0.0051452 | 5.9244e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 264 | 0.86295 | [20 empty] | -0.091942 | 0.72241 | 0.29926 | -0.08091 | 0.016322 | -0.0024404 | 0.00027029 | -1.8822e-05 | 5.8121e-07 | [12 empty] | 1.2716 | 0.4552 | 0.14308 | -0.0052041 | 6.015e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 265 | 0.86724 | [20 empty] | -0.091942 | 0.72514 | 0.30153 | -0.081833 | 0.01657 | -0.0024869 | 0.0002765 | -1.9327e-05 | 5.9905e-07 | [12 empty] | 1.2716 | 0.45692 | 0.14417 | -0.0052634 | 6.1066e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 266 | 0.87154 | [20 empty] | -0.091942 | 0.72788 | 0.30381 | -0.082763 | 0.016822 | -0.0025342 | 0.00028281 | -1.9843e-05 | 6.1738e-07 | [12 empty] | 1.2716 | 0.45865 | 0.14526 | -0.0053232 | 6.1993e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 267 | 0.87584 | [20 empty] | -0.091942 | 0.73062 | 0.3061 | -0.0837 | 0.017076 | -0.0025822 | 0.00028925 | -2.0372e-05 | 6.3619e-07 | [12 empty] | 1.2716 | 0.46037 | 0.14636 | -0.0053835 | 6.2931e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 268 | 0.88014 | [20 empty] | -0.091942 | 0.73335 | 0.3084 | -0.084644 | 0.017334 | -0.0026309 | 0.00029582 | -2.0912e-05 | 6.5551e-07 | [12 empty] | 1.2716 | 0.46209 | 0.14745 | -0.0054442 | 6.3879e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 269 | 0.88445 | [20 empty] | -0.091942 | 0.73609 | 0.3107 | -0.085595 | 0.017594 | -0.0026804 | 0.0003025 | -2.1464e-05 | 6.7533e-07 | [12 empty] | 1.2716 | 0.46382 | 0.14856 | -0.0055054 | 6.4838e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 270 | 0.88876 | [20 empty] | -0.091942 | 0.73882 | 0.31302 | -0.086553 | 0.017857 | -0.0027306 | 0.00030931 | -2.2029e-05 | 6.9568e-07 | [12 empty] | 1.2716 | 0.46554 | 0.14966 | -0.005567 | 6.5807e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 271 | 0.89308 | [20 empty] | -0.091942 | 0.74156 | 0.31534 | -0.087519 | 0.018123 | -0.0027815 | 0.00031625 | -2.2606e-05 | 7.1656e-07 | [12 empty] | 1.2716 | 0.46727 | 0.15077 | -0.0056291 | 6.6788e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 272 | 0.8974 | [20 empty] | -0.091942 | 0.7443 | 0.31767 | -0.088491 | 0.018392 | -0.0028332 | 0.00032332 | -2.3197e-05 | 7.3799e-07 | [12 empty] | 1.2716 | 0.46899 | 0.15189 | -0.0056916 | 6.7779e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 273 | 0.90172 | [20 empty] | -0.091942 | 0.74703 | 0.32001 | -0.089471 | 0.018664 | -0.0028857 | 0.00033051 | -2.38e-05 | 7.5998e-07 | [12 empty] | 1.2716 | 0.47072 | 0.15301 | -0.0057547 | 6.8781e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 274 | 0.90605 | [20 empty] | -0.091942 | 0.74977 | 0.32236 | -0.090458 | 0.018939 | -0.0029389 | 0.00033784 | -2.4417e-05 | 7.8253e-07 | [12 empty] | 1.2716 | 0.47244 | 0.15413 | -0.0058181 | 6.9794e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 275 | 0.91038 | [20 empty] | -0.091942 | 0.75251 | 0.32472 | -0.091452 | 0.019217 | -0.002993 | 0.00034531 | -2.5048e-05 | 8.0568e-07 | [12 empty] | 1.2716 | 0.47416 | 0.15526 | -0.0058821 | 7.0819e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 276 | 0.91471 | [20 empty] | -0.091942 | 0.75524 | 0.32709 | -0.092453 | 0.019498 | -0.0030478 | 0.00035291 | -2.5693e-05 | 8.2941e-07 | [12 empty] | 1.2716 | 0.47589 | 0.15639 | -0.0059465 | 7.1855e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 277 | 0.91905 | [20 empty] | -0.091942 | 0.75798 | 0.32946 | -0.093461 | 0.019782 | -0.0031034 | 0.00036065 | -2.6351e-05 | 8.5376e-07 | [12 empty] | 1.2716 | 0.47761 | 0.15752 | -0.0060113 | 7.2902e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 278 | 0.92339 | [20 empty] | -0.091942 | 0.76072 | 0.33184 | -0.094477 | 0.020069 | -0.0031598 | 0.00036854 | -2.7025e-05 | 8.7873e-07 | [12 empty] | 1.2716 | 0.47934 | 0.15866 | -0.0060767 | 7.396e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 279 | 0.92774 | [20 empty] | -0.091942 | 0.76345 | 0.33423 | -0.0955 | 0.02036 | -0.0032171 | 0.00037656 | -2.7712e-05 | 9.0434e-07 | [12 empty] | 1.2716 | 0.48106 | 0.15981 | -0.0061425 | 7.503e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 280 | 0.93209 | [20 empty] | -0.091942 | 0.76619 | 0.33663 | -0.096531 | 0.020653 | -0.0032751 | 0.00038473 | -2.8415e-05 | 9.306e-07 | [12 empty] | 1.2716 | 0.48279 | 0.16095 | -0.0062088 | 7.6112e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 281 | 0.93644 | [20 empty] | -0.091942 | 0.76892 | 0.33904 | -0.097569 | 0.02095 | -0.003334 | 0.00039305 | -2.9133e-05 | 9.5752e-07 | [12 empty] | 1.2716 | 0.48451 | 0.16211 | -0.0062755 | 7.7205e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 282 | 0.94079 | [20 empty] | -0.091942 | 0.77166 | 0.34146 | -0.098614 | 0.021249 | -0.0033938 | 0.00040152 | -2.9867e-05 | 9.8513e-07 | [12 empty] | 1.2716 | 0.48623 | 0.16326 | -0.0063428 | 7.831e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 283 | 0.94515 | [20 empty] | -0.091942 | 0.7744 | 0.34389 | -0.099667 | 0.021553 | -0.0034544 | 0.00041014 | -3.0616e-05 | 1.0134e-06 | [12 empty] | 1.2716 | 0.48796 | 0.16442 | -0.0064105 | 7.9426e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 284 | 0.94952 | [20 empty] | -0.091942 | 0.77713 | 0.34632 | -0.10073 | 0.021859 | -0.0035158 | 0.00041891 | -3.1381e-05 | 1.0424e-06 | [12 empty] | 1.2716 | 0.48968 | 0.16559 | -0.0064787 | 8.0555e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 285 | 0.95389 | [20 empty] | -0.091942 | 0.77987 | 0.34876 | -0.1018 | 0.022168 | -0.0035782 | 0.00042784 | -3.2163e-05 | 1.0722e-06 | [12 empty] | 1.2716 | 0.49141 | 0.16675 | -0.0065474 | 8.1695e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 286 | 0.95826 | [20 empty] | -0.091942 | 0.78261 | 0.35122 | -0.10287 | 0.022481 | -0.0036414 | 0.00043693 | -3.2961e-05 | 1.1026e-06 | [12 empty] | 1.2716 | 0.49313 | 0.16793 | -0.0066165 | 8.2848e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 287 | 0.96263 | [20 empty] | -0.091942 | 0.78534 | 0.35368 | -0.10395 | 0.022797 | -0.0037055 | 0.00044617 | -3.3777e-05 | 1.1338e-06 | [12 empty] | 1.2716 | 0.49486 | 0.1691 | -0.0066862 | 8.4013e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 288 | 0.96701 | [20 empty] | -0.091942 | 0.78808 | 0.35615 | -0.10504 | 0.023116 | -0.0037705 | 0.00045558 | -3.4609e-05 | 1.1658e-06 | [12 empty] | 1.2716 | 0.49658 | 0.17028 | -0.0067563 | 8.519e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 289 | 0.97139 | [20 empty] | -0.091942 | 0.79082 | 0.35862 | -0.10614 | 0.023439 | -0.0038364 | 0.00046516 | -3.5459e-05 | 1.1986e-06 | [12 empty] | 1.2716 | 0.4983 | 0.17147 | -0.0068269 | 8.6379e-06
0 | 32.0403 | 1 | -0.00012105 | 0.002881 | 32 | 33 | 290 | 0.97577 | [20 empty] | -0.091942 | 0.79355 | 0.36111 | -0.10725 | 0.023765 | -0.0039033 | 0.0004749 | -3.6327e-05 | 1.2322e-06 | [12 empty] | 1.2716 | 0.50003 | 0.17266 | -0.006898 | 8.7581e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 291 | 0.98016 | [20 empty] | -0.091942 | 0.79629 | 0.3636 | -0.10836 | 0.024095 | -0.003971 | 0.00048481 | -3.7213e-05 | 1.2666e-06 | [12 empty] | 1.2716 | 0.50175 | 0.17385 | -0.0069696 | 8.8795e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 292 | 0.98456 | [20 empty] | -0.091942 | 0.79903 | 0.36611 | -0.10948 | 0.024428 | -0.0040397 | 0.00049489 | -3.8117e-05 | 1.3019e-06 | [12 empty] | 1.2716 | 0.50348 | 0.17505 | -0.0070417 | 9.0022e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 293 | 0.98895 | [20 empty] | -0.091942 | 0.80176 | 0.36862 | -0.11061 | 0.024764 | -0.0041094 | 0.00050515 | -3.9041e-05 | 1.338e-06 | [12 empty] | 1.2716 | 0.5052 | 0.17625 | -0.0071143 | 9.1262e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 294 | 0.99335 | [20 empty] | -0.091942 | 0.8045 | 0.37114 | -0.11175 | 0.025104 | -0.00418 | 0.00051558 | -3.9983e-05 | 1.3749e-06 | [12 empty] | 1.2716 | 0.50693 | 0.17745 | -0.0071874 | 9.2514e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 295 | 0.99775 | [20 empty] | -0.091942 | 0.80723 | 0.37367 | -0.11289 | 0.025447 | -0.0042515 | 0.00052619 | -4.0945e-05 | 1.4128e-06 | [12 empty] | 1.2716 | 0.50865 | 0.17866 | -0.007261 | 9.3779e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 296 | 1.0022 | [20 empty] | -0.091942 | 0.80997 | 0.37621 | -0.11404 | 0.025794 | -0.0043241 | 0.00053699 | -4.1926e-05 | 1.4516e-06 | [12 empty] | 1.2716 | 0.51037 | 0.17987 | -0.0073351 | 9.5057e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 297 | 1.0066 | [20 empty] | -0.091942 | 0.81271 | 0.37875 | -0.1152 | 0.026144 | -0.0043976 | 0.00054796 | -4.2928e-05 | 1.4913e-06 | [12 empty] | 1.2716 | 0.5121 | 0.18109 | -0.0074097 | 9.6348e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 298 | 1.011 | [20 empty] | -0.091942 | 0.81544 | 0.38131 | -0.11637 | 0.026498 | -0.0044722 | 0.00055913 | -4.395e-05 | 1.5319e-06 | [12 empty] | 1.2716 | 0.51382 | 0.18231 | -0.0074848 | 9.7653e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 299 | 1.0154 | [20 empty] | -0.091942 | 0.81818 | 0.38387 | -0.11755 | 0.026856 | -0.0045477 | 0.00057048 | -4.4993e-05 | 1.5735e-06 | [12 empty] | 1.2716 | 0.51555 | 0.18354 | -0.0075604 | 9.897e-06
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 300 | 1.0198 | [20 empty] | -0.091942 | 0.82092 | 0.38644 | -0.11873 | 0.027217 | -0.0046243 | 0.00058202 | -4.6057e-05 | 1.6161e-06 | [12 empty] | 1.2716 | 0.51727 | 0.18477 | -0.0076365 | 1.003e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 301 | 1.0242 | [20 empty] | -0.091942 | 0.82365 | 0.38902 | -0.11992 | 0.027582 | -0.0047019 | 0.00059376 | -4.7142e-05 | 1.6597e-06 | [12 empty] | 1.2716 | 0.51899 | 0.186 | -0.0077131 | 1.0164e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 302 | 1.0287 | [20 empty] | -0.091942 | 0.82639 | 0.39161 | -0.12112 | 0.02795 | -0.0047805 | 0.0006057 | -4.825e-05 | 1.7043e-06 | [12 empty] | 1.2716 | 0.52072 | 0.18724 | -0.0077903 | 1.03e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 303 | 1.0331 | [20 empty] | -0.091942 | 0.82913 | 0.39421 | -0.12233 | 0.028322 | -0.0048602 | 0.00061783 | -4.9379e-05 | 1.75e-06 | [12 empty] | 1.2716 | 0.52244 | 0.18848 | -0.0078679 | 1.0437e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 304 | 1.0375 | [20 empty] | -0.091942 | 0.83186 | 0.39682 | -0.12354 | 0.028698 | -0.0049409 | 0.00063016 | -5.0531e-05 | 1.7967e-06 | [12 empty] | 1.2716 | 0.52417 | 0.18973 | -0.0079461 | 1.0576e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 305 | 1.042 | [20 empty] | -0.091942 | 0.8346 | 0.39943 | -0.12476 | 0.029077 | -0.0050227 | 0.0006427 | -5.1706e-05 | 1.8446e-06 | [12 empty] | 1.2716 | 0.52589 | 0.19098 | -0.0080247 | 1.0716e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 306 | 1.0464 | [20 empty] | -0.091942 | 0.83733 | 0.40206 | -0.126 | 0.02946 | -0.0051056 | 0.00065545 | -5.2905e-05 | 1.8935e-06 | [12 empty] | 1.2716 | 0.52762 | 0.19223 | -0.0081039 | 1.0857e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 307 | 1.0509 | [20 empty] | -0.091942 | 0.84007 | 0.40469 | -0.12724 | 0.029847 | -0.0051895 | 0.00066841 | -5.4127e-05 | 1.9436e-06 | [12 empty] | 1.2716 | 0.52934 | 0.19349 | -0.0081836 | 1.0999e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 308 | 1.0553 | [20 empty] | -0.091942 | 0.84281 | 0.40733 | -0.12848 | 0.030238 | -0.0052746 | 0.00068158 | -5.5373e-05 | 1.9948e-06 | [12 empty] | 1.2716 | 0.53106 | 0.19475 | -0.0082639 | 1.1143e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 309 | 1.0598 | [20 empty] | -0.091942 | 0.84554 | 0.40998 | -0.12974 | 0.030633 | -0.0053608 | 0.00069497 | -5.6644e-05 | 2.0472e-06 | [12 empty] | 1.2716 | 0.53279 | 0.19602 | -0.0083446 | 1.1289e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 310 | 1.0642 | [20 empty] | -0.091942 | 0.84828 | 0.41264 | -0.131 | 0.031031 | -0.0054481 | 0.00070857 | -5.794e-05 | 2.1008e-06 | [12 empty] | 1.2716 | 0.53451 | 0.19729 | -0.0084259 | 1.1436e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 311 | 1.0687 | [20 empty] | -0.091942 | 0.85102 | 0.4153 | -0.13227 | 0.031434 | -0.0055365 | 0.0007224 | -5.9261e-05 | 2.1557e-06 | [12 empty] | 1.2716 | 0.53624 | 0.19857 | -0.0085077 | 1.1584e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 312 | 1.0731 | [20 empty] | -0.091942 | 0.85375 | 0.41798 | -0.13355 | 0.03184 | -0.0056261 | 0.00073644 | -6.0608e-05 | 2.2117e-06 | [12 empty] | 1.2716 | 0.53796 | 0.19985 | -0.00859 | 1.1734e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 313 | 1.0776 | [20 empty] | -0.091942 | 0.85649 | 0.42066 | -0.13484 | 0.03225 | -0.0057169 | 0.00075072 | -6.198e-05 | 2.2691e-06 | [12 empty] | 1.2716 | 0.53969 | 0.20113 | -0.0086729 | 1.1885e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 314 | 1.0821 | [20 empty] | -0.091942 | 0.85923 | 0.42335 | -0.13614 | 0.032664 | -0.0058088 | 0.00076523 | -6.338e-05 | 2.3277e-06 | [12 empty] | 1.2716 | 0.54141 | 0.20242 | -0.0087563 | 1.2038e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 315 | 1.0865 | [20 empty] | -0.091942 | 0.86196 | 0.42605 | -0.13744 | 0.033082 | -0.0059019 | 0.00077997 | -6.4806e-05 | 2.3877e-06 | [12 empty] | 1.2716 | 0.54313 | 0.20371 | -0.0088402 | 1.2192e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 316 | 1.091 | [20 empty] | -0.091942 | 0.8647 | 0.42876 | -0.13876 | 0.033504 | -0.0059961 | 0.00079494 | -6.626e-05 | 2.449e-06 | [12 empty] | 1.2716 | 0.54486 | 0.205 | -0.0089247 | 1.2347e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 317 | 1.0955 | [20 empty] | -0.091942 | 0.86743 | 0.43148 | -0.14008 | 0.03393 | -0.0060916 | 0.00081015 | -6.7742e-05 | 2.5117e-06 | [12 empty] | 1.2716 | 0.54658 | 0.2063 | -0.0090097 | 1.2504e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 318 | 1.1 | [20 empty] | -0.091942 | 0.87017 | 0.43421 | -0.14141 | 0.034361 | -0.0061883 | 0.00082561 | -6.9252e-05 | 2.5758e-06 | [12 empty] | 1.2716 | 0.54831 | 0.20761 | -0.0090952 | 1.2663e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 319 | 1.1044 | [20 empty] | -0.091942 | 0.87291 | 0.43694 | -0.14275 | 0.034795 | -0.0062862 | 0.00084131 | -7.0791e-05 | 2.6413e-06 | [12 empty] | 1.2716 | 0.55003 | 0.20891 | -0.0091813 | 1.2823e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 320 | 1.1089 | [20 empty] | -0.091942 | 0.87564 | 0.43969 | -0.14409 | 0.035233 | -0.0063854 | 0.00085726 | -7.2359e-05 | 2.7083e-06 | [12 empty] | 1.2716 | 0.55176 | 0.21023 | -0.0092679 | 1.2984e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 321 | 1.1134 | [20 empty] | -0.091942 | 0.87838 | 0.44244 | -0.14545 | 0.035676 | -0.0064858 | 0.00087346 | -7.3957e-05 | 2.7768e-06 | [12 empty] | 1.2716 | 0.55348 | 0.21154 | -0.0093551 | 1.3147e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 322 | 1.1179 | [20 empty] | -0.091942 | 0.88112 | 0.4452 | -0.14681 | 0.036122 | -0.0065874 | 0.00088991 | -7.5585e-05 | 2.8467e-06 | [12 empty] | 1.2716 | 0.5552 | 0.21286 | -0.0094428 | 1.3312e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 323 | 1.1224 | [20 empty] | -0.091942 | 0.88385 | 0.44797 | -0.14818 | 0.036573 | -0.0066904 | 0.00090662 | -7.7244e-05 | 2.9182e-06 | [12 empty] | 1.2716 | 0.55693 | 0.21419 | -0.009531 | 1.3478e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 324 | 1.1269 | [20 empty] | -0.091942 | 0.88659 | 0.45075 | -0.14956 | 0.037028 | -0.0067946 | 0.0009236 | -7.8933e-05 | 2.9913e-06 | [12 empty] | 1.2716 | 0.55865 | 0.21551 | -0.0096198 | 1.3646e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 325 | 1.1314 | [20 empty] | -0.091942 | 0.88933 | 0.45353 | -0.15095 | 0.037487 | -0.0069001 | 0.00094083 | -8.0654e-05 | 3.066e-06 | [12 empty] | 1.2716 | 0.56038 | 0.21685 | -0.0097092 | 1.3815e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 326 | 1.1359 | [20 empty] | -0.091942 | 0.89206 | 0.45633 | -0.15235 | 0.037951 | -0.0070069 | 0.00095834 | -8.2408e-05 | 3.1422e-06 | [12 empty] | 1.2716 | 0.5621 | 0.21818 | -0.0097991 | 1.3986e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 327 | 1.1404 | [20 empty] | -0.091942 | 0.8948 | 0.45913 | -0.15376 | 0.038419 | -0.007115 | 0.00097611 | -8.4194e-05 | 3.2202e-06 | [12 empty] | 1.2716 | 0.56382 | 0.21952 | -0.0098895 | 1.4158e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 328 | 1.1449 | [20 empty] | -0.091942 | 0.89753 | 0.46195 | -0.15517 | 0.038891 | -0.0072245 | 0.00099416 | -8.6012e-05 | 3.2998e-06 | [12 empty] | 1.2716 | 0.56555 | 0.22087 | -0.0099805 | 1.4332e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 329 | 1.1495 | [20 empty] | -0.091942 | 0.90027 | 0.46477 | -0.1566 | 0.039367 | -0.0073353 | 0.0010125 | -8.7865e-05 | 3.3812e-06 | [12 empty] | 1.2716 | 0.56727 | 0.22222 | -0.010072 | 1.4508e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 330 | 1.154 | [20 empty] | -0.091942 | 0.90301 | 0.4676 | -0.15803 | 0.039848 | -0.0074474 | 0.0010311 | -8.9752e-05 | 3.4643e-06 | [12 empty] | 1.2716 | 0.569 | 0.22357 | -0.010164 | 1.4685e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 331 | 1.1585 | [20 empty] | -0.091942 | 0.90574 | 0.47043 | -0.15947 | 0.040333 | -0.007561 | 0.00105 | -9.1673e-05 | 3.5491e-06 | [12 empty] | 1.2716 | 0.57072 | 0.22493 | -0.010257 | 1.4864e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 332 | 1.163 | [20 empty] | -0.091942 | 0.90848 | 0.47328 | -0.16092 | 0.040823 | -0.0076759 | 0.0010692 | -9.3629e-05 | 3.6358e-06 | [12 empty] | 1.2716 | 0.57245 | 0.22629 | -0.01035 | 1.5044e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 333 | 1.1676 | [20 empty] | -0.091942 | 0.91122 | 0.47614 | -0.16238 | 0.041317 | -0.0077922 | 0.0010886 | -9.5621e-05 | 3.7244e-06 | [12 empty] | 1.2716 | 0.57417 | 0.22765 | -0.010444 | 1.5226e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 334 | 1.1721 | [20 empty] | -0.091942 | 0.91395 | 0.479 | -0.16384 | 0.041816 | -0.0079099 | 0.0011084 | -9.7649e-05 | 3.8148e-06 | [12 empty] | 1.2716 | 0.57589 | 0.22902 | -0.010538 | 1.541e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 335 | 1.1766 | [20 empty] | -0.091942 | 0.91669 | 0.48187 | -0.16532 | 0.042319 | -0.008029 | 0.0011284 | -9.9714e-05 | 3.9071e-06 | [12 empty] | 1.2716 | 0.57762 | 0.2304 | -0.010633 | 1.5595e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 336 | 1.1812 | [20 empty] | -0.091942 | 0.91943 | 0.48475 | -0.16681 | 0.042826 | -0.0081495 | 0.0011488 | -0.00010182 | 4.0014e-06 | [12 empty] | 1.2716 | 0.57934 | 0.23177 | -0.010729 | 1.5782e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 337 | 1.1857 | [20 empty] | -0.091942 | 0.92216 | 0.48764 | -0.1683 | 0.043338 | -0.0082715 | 0.0011695 | -0.00010396 | 4.0977e-06 | [12 empty] | 1.2716 | 0.58107 | 0.23316 | -0.010825 | 1.5971e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 338 | 1.1902 | [20 empty] | -0.091942 | 0.9249 | 0.49054 | -0.1698 | 0.043855 | -0.008395 | 0.0011905 | -0.00010614 | 4.196e-06 | [12 empty] | 1.2716 | 0.58279 | 0.23454 | -0.010921 | 1.6162e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 339 | 1.1948 | [20 empty] | -0.091942 | 0.92764 | 0.49345 | -0.17131 | 0.044376 | -0.0085199 | 0.0012117 | -0.00010835 | 4.2963e-06 | [12 empty] | 1.2716 | 0.58452 | 0.23593 | -0.011019 | 1.6354e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 340 | 1.1993 | [20 empty] | -0.091942 | 0.93037 | 0.49636 | -0.17283 | 0.044902 | -0.0086463 | 0.0012333 | -0.00011061 | 4.3988e-06 | [12 empty] | 1.2716 | 0.58624 | 0.23732 | -0.011116 | 1.6548e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 341 | 1.2039 | [20 empty] | -0.091942 | 0.93311 | 0.49929 | -0.17436 | 0.045433 | -0.0087742 | 0.0012553 | -0.00011291 | 4.5033e-06 | [12 empty] | 1.2716 | 0.58796 | 0.23872 | -0.011215 | 1.6743e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 342 | 1.2085 | [20 empty] | -0.091942 | 0.93584 | 0.50222 | -0.1759 | 0.045968 | -0.0089036 | 0.0012775 | -0.00011525 | 4.6101e-06 | [12 empty] | 1.2716 | 0.58969 | 0.24013 | -0.011314 | 1.694e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 343 | 1.213 | [20 empty] | -0.091942 | 0.93858 | 0.50516 | -0.17745 | 0.046508 | -0.0090346 | 0.0013001 | -0.00011763 | 4.719e-06 | [12 empty] | 1.2716 | 0.59141 | 0.24153 | -0.011413 | 1.7139e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 344 | 1.2176 | [20 empty] | -0.091942 | 0.94132 | 0.50811 | -0.17901 | 0.047053 | -0.009167 | 0.001323 | -0.00012005 | 4.8302e-06 | [12 empty] | 1.2716 | 0.59314 | 0.24294 | -0.011513 | 1.734e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 345 | 1.2221 | [20 empty] | -0.091942 | 0.94405 | 0.51107 | -0.18057 | 0.047602 | -0.0093011 | 0.0013463 | -0.00012251 | 4.9437e-06 | [12 empty] | 1.2716 | 0.59486 | 0.24436 | -0.011614 | 1.7543e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 346 | 1.2267 | [20 empty] | -0.091942 | 0.94679 | 0.51404 | -0.18215 | 0.048157 | -0.0094366 | 0.0013698 | -0.00012502 | 5.0595e-06 | [12 empty] | 1.2716 | 0.59659 | 0.24578 | -0.011715 | 1.7747e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 347 | 1.2313 | [20 empty] | -0.091942 | 0.94953 | 0.51701 | -0.18373 | 0.048716 | -0.0095738 | 0.0013938 | -0.00012757 | 5.1777e-06 | [12 empty] | 1.2716 | 0.59831 | 0.2472 | -0.011817 | 1.7953e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 348 | 1.2359 | [20 empty] | -0.091942 | 0.95226 | 0.52 | -0.18532 | 0.04928 | -0.0097125 | 0.001418 | -0.00013017 | 5.2983e-06 | [12 empty] | 1.2716 | 0.60003 | 0.24862 | -0.01192 | 1.8161e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 349 | 1.2404 | [20 empty] | -0.091942 | 0.955 | 0.52299 | -0.18693 | 0.049849 | -0.0098529 | 0.0014427 | -0.00013281 | 5.4213e-06 | [12 empty] | 1.2716 | 0.60176 | 0.25006 | -0.012023 | 1.837e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 350 | 1.245 | [20 empty] | -0.091942 | 0.95774 | 0.52599 | -0.18854 | 0.050422 | -0.0099949 | 0.0014676 | -0.00013549 | 5.5468e-06 | [12 empty] | 1.2716 | 0.60348 | 0.25149 | -0.012126 | 1.8582e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 351 | 1.2496 | [20 empty] | -0.091942 | 0.96047 | 0.529 | -0.19016 | 0.051001 | -0.010138 | 0.001493 | -0.00013823 | 5.6749e-06 | [12 empty] | 1.2716 | 0.60521 | 0.25293 | -0.012231 | 1.8795e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 352 | 1.2542 | [20 empty] | -0.091942 | 0.96321 | 0.53202 | -0.19179 | 0.051585 | -0.010284 | 0.0015187 | -0.00014101 | 5.8055e-06 | [12 empty] | 1.2716 | 0.60693 | 0.25437 | -0.012336 | 1.901e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 353 | 1.2588 | [20 empty] | -0.091942 | 0.96594 | 0.53505 | -0.19343 | 0.052174 | -0.010431 | 0.0015448 | -0.00014384 | 5.9388e-06 | [12 empty] | 1.2716 | 0.60865 | 0.25582 | -0.012441 | 1.9227e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 354 | 1.2634 | [20 empty] | -0.091942 | 0.96868 | 0.53808 | -0.19508 | 0.052767 | -0.010579 | 0.0015712 | -0.00014671 | 6.0747e-06 | [12 empty] | 1.2716 | 0.61038 | 0.25727 | -0.012547 | 1.9446e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 355 | 1.268 | [20 empty] | -0.091942 | 0.97142 | 0.54113 | -0.19673 | 0.053366 | -0.010729 | 0.001598 | -0.00014964 | 6.2133e-06 | [12 empty] | 1.2716 | 0.6121 | 0.25873 | -0.012654 | 1.9667e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 356 | 1.2726 | [20 empty] | -0.091942 | 0.97415 | 0.54418 | -0.1984 | 0.05397 | -0.010881 | 0.0016252 | -0.00015261 | 6.3548e-06 | [12 empty] | 1.2716 | 0.61383 | 0.26019 | -0.012761 | 1.9889e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 357 | 1.2772 | [20 empty] | -0.091942 | 0.97689 | 0.54724 | -0.20008 | 0.054579 | -0.011035 | 0.0016528 | -0.00015564 | 6.499e-06 | [12 empty] | 1.2716 | 0.61555 | 0.26165 | -0.012869 | 2.0114e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 358 | 1.2818 | [20 empty] | -0.091942 | 0.97963 | 0.55031 | -0.20176 | 0.055193 | -0.011191 | 0.0016808 | -0.00015872 | 6.646e-06 | [12 empty] | 1.2716 | 0.61728 | 0.26312 | -0.012977 | 2.034e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 359 | 1.2864 | [20 empty] | -0.091942 | 0.98236 | 0.55339 | -0.20346 | 0.055812 | -0.011348 | 0.0017091 | -0.00016185 | 6.796e-06 | [12 empty] | 1.2716 | 0.619 | 0.26459 | -0.013086 | 2.0568e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 360 | 1.291 | [20 empty] | -0.091942 | 0.9851 | 0.55648 | -0.20516 | 0.056437 | -0.011507 | 0.0017379 | -0.00016503 | 6.9489e-06 | [12 empty] | 1.2716 | 0.62072 | 0.26607 | -0.013196 | 2.0798e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 361 | 1.2956 | [20 empty] | -0.091942 | 0.98784 | 0.55957 | -0.20688 | 0.057066 | -0.011667 | 0.0017671 | -0.00016827 | 7.1049e-06 | [12 empty] | 1.2716 | 0.62245 | 0.26755 | -0.013306 | 2.103e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 362 | 1.3002 | [20 empty] | -0.091942 | 0.99057 | 0.56268 | -0.2086 | 0.057701 | -0.01183 | 0.0017967 | -0.00017156 | 7.2639e-06 | [12 empty] | 1.2716 | 0.62417 | 0.26903 | -0.013417 | 2.1264e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 363 | 1.3048 | [20 empty] | -0.091942 | 0.99331 | 0.56579 | -0.21034 | 0.058342 | -0.011994 | 0.0018266 | -0.0001749 | 7.4259e-06 | [12 empty] | 1.2716 | 0.6259 | 0.27052 | -0.013529 | 2.15e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 364 | 1.3095 | [20 empty] | -0.091942 | 0.99604 | 0.56891 | -0.21208 | 0.058987 | -0.01216 | 0.001857 | -0.0001783 | 7.5912e-06 | [12 empty] | 1.2716 | 0.62762 | 0.27201 | -0.013641 | 2.1738e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 365 | 1.3141 | [20 empty] | -0.091942 | 0.99878 | 0.57204 | -0.21383 | 0.059638 | -0.012328 | 0.0018879 | -0.00018176 | 7.7596e-06 | [12 empty] | 1.2716 | 0.62935 | 0.27351 | -0.013753 | 2.1978e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 366 | 1.3187 | [20 empty] | -0.091942 | 1.0015 | 0.57518 | -0.21559 | 0.060294 | -0.012498 | 0.0019191 | -0.00018527 | 7.9314e-06 | [12 empty] | 1.2716 | 0.63107 | 0.27501 | -0.013867 | 2.222e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 367 | 1.3233 | [20 empty] | -0.091942 | 1.0043 | 0.57833 | -0.21737 | 0.060956 | -0.01267 | 0.0019508 | -0.00018885 | 8.1064e-06 | [12 empty] | 1.2716 | 0.63279 | 0.27651 | -0.013981 | 2.2464e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 368 | 1.328 | [20 empty] | -0.091942 | 1.007 | 0.58149 | -0.21915 | 0.061623 | -0.012843 | 0.0019829 | -0.00019248 | 8.2848e-06 | [12 empty] | 1.2716 | 0.63452 | 0.27802 | -0.014095 | 2.271e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 369 | 1.3326 | [20 empty] | -0.091942 | 1.0097 | 0.58465 | -0.22094 | 0.062296 | -0.013019 | 0.0020154 | -0.00019617 | 8.4666e-06 | [12 empty] | 1.2716 | 0.63624 | 0.27954 | -0.014211 | 2.2957e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 370 | 1.3372 | [20 empty] | -0.091942 | 1.0125 | 0.58782 | -0.22274 | 0.062974 | -0.013196 | 0.0020484 | -0.00019992 | 8.6519e-06 | [12 empty] | 1.2716 | 0.63797 | 0.28105 | -0.014326 | 2.3207e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 371 | 1.3419 | [20 empty] | -0.091942 | 1.0152 | 0.591 | -0.22455 | 0.063657 | -0.013375 | 0.0020819 | -0.00020373 | 8.8408e-06 | [12 empty] | 1.2716 | 0.63969 | 0.28257 | -0.014443 | 2.3459e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 372 | 1.3465 | [20 empty] | -0.091942 | 1.0179 | 0.59419 | -0.22637 | 0.064346 | -0.013557 | 0.0021158 | -0.00020761 | 9.0332e-06 | [12 empty] | 1.2716 | 0.64142 | 0.2841 | -0.01456 | 2.3713e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 373 | 1.3512 | [20 empty] | -0.091942 | 1.0207 | 0.59739 | -0.2282 | 0.065041 | -0.01374 | 0.0021501 | -0.00021155 | 9.2293e-06 | [12 empty] | 1.2716 | 0.64314 | 0.28563 | -0.014678 | 2.3969e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 374 | 1.3558 | [20 empty] | -0.091942 | 1.0234 | 0.6006 | -0.23004 | 0.065741 | -0.013925 | 0.002185 | -0.00021555 | 9.4291e-06 | [12 empty] | 1.2716 | 0.64486 | 0.28716 | -0.014796 | 2.4227e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 375 | 1.3605 | [20 empty] | -0.091942 | 1.0261 | 0.60382 | -0.23189 | 0.066447 | -0.014112 | 0.0022202 | -0.00021962 | 9.6327e-06 | [12 empty] | 1.2716 | 0.64659 | 0.2887 | -0.014915 | 2.4487e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 376 | 1.3651 | [20 empty] | -0.091942 | 1.0289 | 0.60704 | -0.23375 | 0.067159 | -0.014301 | 0.002256 | -0.00022375 | 9.8402e-06 | [12 empty] | 1.2716 | 0.64831 | 0.29024 | -0.015035 | 2.475e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 377 | 1.3698 | [20 empty] | -0.091942 | 1.0316 | 0.61028 | -0.23562 | 0.067876 | -0.014493 | 0.0022922 | -0.00022795 | 1.0051e-05 | [12 empty] | 1.2716 | 0.65004 | 0.29179 | -0.015155 | 2.5014e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 378 | 1.3744 | [20 empty] | -0.091942 | 1.0344 | 0.61352 | -0.2375 | 0.068599 | -0.014686 | 0.002329 | -0.00023221 | 1.0267e-05 | [12 empty] | 1.2716 | 0.65176 | 0.29334 | -0.015276 | 2.5281e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 379 | 1.3791 | [20 empty] | -0.091942 | 1.0371 | 0.61677 | -0.23939 | 0.069328 | -0.014881 | 0.0023662 | -0.00023655 | 1.0486e-05 | [12 empty] | 1.2716 | 0.65349 | 0.29489 | -0.015397 | 2.5549e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 380 | 1.3838 | [20 empty] | -0.091942 | 1.0398 | 0.62003 | -0.24129 | 0.070063 | -0.015078 | 0.0024039 | -0.00024095 | 1.0709e-05 | [12 empty] | 1.2716 | 0.65521 | 0.29645 | -0.01552 | 2.582e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 381 | 1.3884 | [20 empty] | -0.091942 | 1.0426 | 0.62329 | -0.2432 | 0.070803 | -0.015278 | 0.0024421 | -0.00024543 | 1.0937e-05 | [12 empty] | 1.2716 | 0.65693 | 0.29801 | -0.015643 | 2.6093e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 382 | 1.3931 | [20 empty] | -0.091942 | 1.0453 | 0.62657 | -0.24512 | 0.071549 | -0.015479 | 0.0024808 | -0.00024997 | 1.1169e-05 | [12 empty] | 1.2716 | 0.65866 | 0.29958 | -0.015766 | 2.6368e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 383 | 1.3978 | [20 empty] | -0.091942 | 1.048 | 0.62985 | -0.24705 | 0.072301 | -0.015683 | 0.00252 | -0.00025459 | 1.1405e-05 | [12 empty] | 1.2716 | 0.66038 | 0.30115 | -0.01589 | 2.6645e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 384 | 1.4025 | [20 empty] | -0.091942 | 1.0508 | 0.63315 | -0.24899 | 0.07306 | -0.015889 | 0.0025598 | -0.00025928 | 1.1645e-05 | [12 empty] | 1.2716 | 0.66211 | 0.30272 | -0.016015 | 2.6924e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 385 | 1.4071 | [20 empty] | -0.091942 | 1.0535 | 0.63645 | -0.25094 | 0.073824 | -0.016097 | 0.0026 | -0.00026404 | 1.189e-05 | [12 empty] | 1.2716 | 0.66383 | 0.3043 | -0.01614 | 2.7206e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 386 | 1.4118 | [20 empty] | -0.091942 | 1.0562 | 0.63976 | -0.2529 | 0.074594 | -0.016307 | 0.0026408 | -0.00026888 | 1.2139e-05 | [12 empty] | 1.2716 | 0.66555 | 0.30589 | -0.016266 | 2.749e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 387 | 1.4165 | [20 empty] | -0.091942 | 1.059 | 0.64308 | -0.25487 | 0.07537 | -0.016519 | 0.0026821 | -0.00027379 | 1.2393e-05 | [12 empty] | 1.2716 | 0.66728 | 0.30747 | -0.016393 | 2.7776e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 388 | 1.4212 | [20 empty] | -0.091942 | 1.0617 | 0.64641 | -0.25686 | 0.076152 | -0.016734 | 0.002724 | -0.00027878 | 1.2652e-05 | [12 empty] | 1.2716 | 0.669 | 0.30906 | -0.016521 | 2.8064e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 389 | 1.4259 | [20 empty] | -0.091942 | 1.0645 | 0.64974 | -0.25885 | 0.07694 | -0.016951 | 0.0027664 | -0.00028385 | 1.2915e-05 | [12 empty] | 1.2716 | 0.67073 | 0.31066 | -0.016649 | 2.8354e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 390 | 1.4306 | [20 empty] | -0.091942 | 1.0672 | 0.65309 | -0.26085 | 0.077734 | -0.01717 | 0.0028093 | -0.000289 | 1.3183e-05 | [12 empty] | 1.2716 | 0.67245 | 0.31226 | -0.016777 | 2.8647e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 391 | 1.4353 | [20 empty] | -0.091942 | 1.0699 | 0.65644 | -0.26286 | 0.078534 | -0.017391 | 0.0028528 | -0.00029423 | 1.3456e-05 | [12 empty] | 1.2716 | 0.67418 | 0.31386 | -0.016907 | 2.8942e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 392 | 1.44 | [20 empty] | -0.091942 | 1.0727 | 0.6598 | -0.26488 | 0.079341 | -0.017614 | 0.0028969 | -0.00029953 | 1.3734e-05 | [12 empty] | 1.2716 | 0.6759 | 0.31547 | -0.017037 | 2.9239e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 393 | 1.4447 | [20 empty] | -0.091942 | 1.0754 | 0.66317 | -0.26691 | 0.080153 | -0.01784 | 0.0029415 | -0.00030492 | 1.4016e-05 | [12 empty] | 1.2716 | 0.67762 | 0.31708 | -0.017168 | 2.9539e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 394 | 1.4494 | [20 empty] | -0.091942 | 1.0781 | 0.66655 | -0.26896 | 0.080972 | -0.018068 | 0.0029867 | -0.0003104 | 1.4304e-05 | [12 empty] | 1.2716 | 0.67935 | 0.3187 | -0.017299 | 2.984e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 395 | 1.4541 | [20 empty] | -0.091942 | 1.0809 | 0.66994 | -0.27101 | 0.081798 | -0.018299 | 0.0030325 | -0.00031595 | 1.4597e-05 | [12 empty] | 1.2716 | 0.68107 | 0.32032 | -0.017431 | 3.0144e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 396 | 1.4588 | [20 empty] | -0.091942 | 1.0836 | 0.67334 | -0.27307 | 0.082629 | -0.018532 | 0.0030788 | -0.0003216 | 1.4896e-05 | [12 empty] | 1.2716 | 0.6828 | 0.32194 | -0.017564 | 3.0451e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 397 | 1.4635 | [20 empty] | -0.091942 | 1.0863 | 0.67674 | -0.27515 | 0.083467 | -0.018767 | 0.0031258 | -0.00032732 | 1.5199e-05 | [12 empty] | 1.2716 | 0.68452 | 0.32357 | -0.017697 | 3.076e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 398 | 1.4682 | [20 empty] | -0.091942 | 1.0891 | 0.68016 | -0.27723 | 0.084311 | -0.019004 | 0.0031733 | -0.00033314 | 1.5508e-05 | [12 empty] | 1.2716 | 0.68625 | 0.3252 | -0.017831 | 3.1071e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 399 | 1.4729 | [20 empty] | -0.091942 | 1.0918 | 0.68358 | -0.27933 | 0.085162 | -0.019244 | 0.0032214 | -0.00033904 | 1.5823e-05 | [12 empty] | 1.2716 | 0.68797 | 0.32684 | -0.017966 | 3.1384e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 400 | 1.4777 | [20 empty] | -0.091942 | 1.0946 | 0.68701 | -0.28143 | 0.086019 | -0.019487 | 0.0032702 | -0.00034504 | 1.6143e-05 | [12 empty] | 1.2716 | 0.68969 | 0.32848 | -0.018101 | 3.17e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 401 | 1.4824 | [20 empty] | -0.091942 | 1.0973 | 0.69045 | -0.28355 | 0.086882 | -0.019731 | 0.0033195 | -0.00035112 | 1.6468e-05 | [12 empty] | 1.2716 | 0.69142 | 0.33012 | -0.018237 | 3.2018e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 402 | 1.4871 | [20 empty] | -0.091942 | 1.1 | 0.6939 | -0.28567 | 0.087752 | -0.019979 | 0.0033695 | -0.00035729 | 1.68e-05 | [12 empty] | 1.2716 | 0.69314 | 0.33177 | -0.018374 | 3.2339e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 403 | 1.4918 | [20 empty] | -0.091942 | 1.1028 | 0.69735 | -0.28781 | 0.088628 | -0.020228 | 0.0034201 | -0.00036356 | 1.7137e-05 | [12 empty] | 1.2716 | 0.69487 | 0.33342 | -0.018512 | 3.2662e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 404 | 1.4966 | [20 empty] | -0.091942 | 1.1055 | 0.70082 | -0.28996 | 0.089511 | -0.020481 | 0.0034714 | -0.00036993 | 1.748e-05 | [12 empty] | 1.2716 | 0.69659 | 0.33508 | -0.01865 | 3.2987e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 405 | 1.5013 | [20 empty] | -0.091942 | 1.1082 | 0.70429 | -0.29212 | 0.090401 | -0.020735 | 0.0035232 | -0.00037638 | 1.7829e-05 | [12 empty] | 1.2716 | 0.69832 | 0.33674 | -0.018789 | 3.3315e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 406 | 1.506 | [20 empty] | -0.091942 | 1.111 | 0.70777 | -0.29429 | 0.091297 | -0.020993 | 0.0035758 | -0.00038294 | 1.8185e-05 | [12 empty] | 1.2716 | 0.70004 | 0.33841 | -0.018928 | 3.3645e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 407 | 1.5108 | [20 empty] | -0.091942 | 1.1137 | 0.71127 | -0.29647 | 0.0922 | -0.021252 | 0.0036289 | -0.00038959 | 1.8546e-05 | [12 empty] | 1.2716 | 0.70176 | 0.34007 | -0.019068 | 3.3978e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 408 | 1.5155 | [20 empty] | -0.091942 | 1.1164 | 0.71476 | -0.29866 | 0.093109 | -0.021515 | 0.0036828 | -0.00039634 | 1.8914e-05 | [12 empty] | 1.2716 | 0.70349 | 0.34175 | -0.019209 | 3.4313e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 409 | 1.5203 | [20 empty] | -0.091942 | 1.1192 | 0.71827 | -0.30086 | 0.094025 | -0.02178 | 0.0037373 | -0.00040319 | 1.9288e-05 | [12 empty] | 1.2716 | 0.70521 | 0.34343 | -0.019351 | 3.4651e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 410 | 1.525 | [20 empty] | -0.091942 | 1.1219 | 0.72179 | -0.30307 | 0.094948 | -0.022047 | 0.0037924 | -0.00041014 | 1.9668e-05 | [12 empty] | 1.2716 | 0.70694 | 0.34511 | -0.019493 | 3.4991e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 411 | 1.5297 | [20 empty] | -0.091942 | 1.1247 | 0.72531 | -0.30529 | 0.095878 | -0.022318 | 0.0038483 | -0.00041719 | 2.0056e-05 | [12 empty] | 1.2716 | 0.70866 | 0.34679 | -0.019636 | 3.5333e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 412 | 1.5345 | [20 empty] | -0.091942 | 1.1274 | 0.72885 | -0.30753 | 0.096815 | -0.02259 | 0.0039048 | -0.00042435 | 2.0449e-05 | [12 empty] | 1.2716 | 0.71038 | 0.34848 | -0.01978 | 3.5679e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 413 | 1.5392 | [20 empty] | -0.091942 | 1.1301 | 0.73239 | -0.30977 | 0.097758 | -0.022866 | 0.003962 | -0.00043161 | 2.085e-05 | [12 empty] | 1.2716 | 0.71211 | 0.35018 | -0.019924 | 3.6026e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 414 | 1.544 | [20 empty] | -0.091942 | 1.1329 | 0.73594 | -0.31203 | 0.098708 | -0.023144 | 0.0040199 | -0.00043898 | 2.1257e-05 | [12 empty] | 1.2716 | 0.71383 | 0.35187 | -0.020069 | 3.6376e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 415 | 1.5488 | [20 empty] | -0.091942 | 1.1356 | 0.7395 | -0.31429 | 0.099665 | -0.023425 | 0.0040785 | -0.00044646 | 2.1671e-05 | [12 empty] | 1.2716 | 0.71556 | 0.35358 | -0.020215 | 3.6729e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 416 | 1.5535 | [20 empty] | -0.091942 | 1.1383 | 0.74307 | -0.31657 | 0.10063 | -0.023708 | 0.0041378 | -0.00045404 | 2.2093e-05 | [12 empty] | 1.2716 | 0.71728 | 0.35528 | -0.020362 | 3.7084e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 417 | 1.5583 | [20 empty] | -0.091942 | 1.1411 | 0.74665 | -0.31886 | 0.1016 | -0.023995 | 0.0041979 | -0.00046174 | 2.2521e-05 | [12 empty] | 1.2716 | 0.71901 | 0.35699 | -0.020509 | 3.7442e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 418 | 1.563 | [20 empty] | -0.091942 | 1.1438 | 0.75023 | -0.32116 | 0.10258 | -0.024284 | 0.0042586 | -0.00046955 | 2.2957e-05 | [12 empty] | 1.2716 | 0.72073 | 0.35871 | -0.020657 | 3.7803e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 419 | 1.5678 | [20 empty] | -0.091942 | 1.1465 | 0.75383 | -0.32347 | 0.10356 | -0.024576 | 0.0043201 | -0.00047747 | 2.34e-05 | [12 empty] | 1.2716 | 0.72245 | 0.36042 | -0.020805 | 3.8166e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 420 | 1.5726 | [20 empty] | -0.091942 | 1.1493 | 0.75743 | -0.32579 | 0.10456 | -0.02487 | 0.0043824 | -0.0004855 | 2.385e-05 | [12 empty] | 1.2716 | 0.72418 | 0.36215 | -0.020955 | 3.8531e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 421 | 1.5774 | [20 empty] | -0.091942 | 1.152 | 0.76104 | -0.32812 | 0.10556 | -0.025168 | 0.0044453 | -0.00049365 | 2.4308e-05 | [12 empty] | 1.2716 | 0.7259 | 0.36387 | -0.021105 | 3.89e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 422 | 1.5821 | [20 empty] | -0.091942 | 1.1548 | 0.76466 | -0.33047 | 0.10656 | -0.025468 | 0.0045091 | -0.00050192 | 2.4774e-05 | [12 empty] | 1.2716 | 0.72763 | 0.3656 | -0.021255 | 3.9271e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 423 | 1.5869 | [20 empty] | -0.091942 | 1.1575 | 0.76829 | -0.33282 | 0.10758 | -0.025771 | 0.0045736 | -0.0005103 | 2.5248e-05 | [12 empty] | 1.2716 | 0.72935 | 0.36734 | -0.021407 | 3.9644e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 424 | 1.5917 | [20 empty] | -0.091942 | 1.1602 | 0.77192 | -0.33519 | 0.1086 | -0.026078 | 0.0046388 | -0.00051881 | 2.5729e-05 | [12 empty] | 1.2716 | 0.73108 | 0.36908 | -0.021559 | 4.002e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 425 | 1.5965 | [20 empty] | -0.091942 | 1.163 | 0.77557 | -0.33757 | 0.10962 | -0.026386 | 0.0047049 | -0.00052743 | 2.6219e-05 | [12 empty] | 1.2716 | 0.7328 | 0.37082 | -0.021712 | 4.0399e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 426 | 1.6013 | [20 empty] | -0.091942 | 1.1657 | 0.77922 | -0.33996 | 0.11066 | -0.026698 | 0.0047717 | -0.00053618 | 2.6716e-05 | [12 empty] | 1.2716 | 0.73452 | 0.37257 | -0.021866 | 4.0781e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 427 | 1.606 | [20 empty] | -0.091942 | 1.1684 | 0.78289 | -0.34235 | 0.1117 | -0.027013 | 0.0048393 | -0.00054505 | 2.7222e-05 | [12 empty] | 1.2716 | 0.73625 | 0.37432 | -0.02202 | 4.1165e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 428 | 1.6108 | [20 empty] | -0.091942 | 1.1712 | 0.78656 | -0.34477 | 0.11275 | -0.027331 | 0.0049077 | -0.00055405 | 2.7736e-05 | [12 empty] | 1.2716 | 0.73797 | 0.37607 | -0.022175 | 4.1552e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 429 | 1.6156 | [20 empty] | -0.091942 | 1.1739 | 0.79024 | -0.34719 | 0.11381 | -0.027652 | 0.0049769 | -0.00056318 | 2.8259e-05 | [12 empty] | 1.2716 | 0.7397 | 0.37783 | -0.022331 | 4.1942e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 430 | 1.6204 | [20 empty] | -0.091942 | 1.1766 | 0.79393 | -0.34962 | 0.11488 | -0.027976 | 0.0050469 | -0.00057243 | 2.879e-05 | [12 empty] | 1.2716 | 0.74142 | 0.3796 | -0.022487 | 4.2334e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 431 | 1.6252 | [20 empty] | -0.091942 | 1.1794 | 0.79762 | -0.35207 | 0.11595 | -0.028302 | 0.0051177 | -0.00058182 | 2.933e-05 | [12 empty] | 1.2716 | 0.74315 | 0.38136 | -0.022645 | 4.2729e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 432 | 1.63 | [20 empty] | -0.091942 | 1.1821 | 0.80133 | -0.35452 | 0.11703 | -0.028632 | 0.0051894 | -0.00059133 | 2.9879e-05 | [12 empty] | 1.2716 | 0.74487 | 0.38314 | -0.022803 | 4.3127e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 433 | 1.6348 | [20 empty] | -0.091942 | 1.1849 | 0.80504 | -0.35699 | 0.11811 | -0.028965 | 0.0052619 | -0.00060098 | 3.0437e-05 | [12 empty] | 1.2716 | 0.74659 | 0.38491 | -0.022961 | 4.3528e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 434 | 1.6396 | [20 empty] | -0.091942 | 1.1876 | 0.80877 | -0.35947 | 0.11921 | -0.029301 | 0.0053352 | -0.00061076 | 3.1004e-05 | [12 empty] | 1.2716 | 0.74832 | 0.38669 | -0.023121 | 4.3932e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 435 | 1.6444 | [20 empty] | -0.091942 | 1.1903 | 0.8125 | -0.36196 | 0.12031 | -0.02964 | 0.0054094 | -0.00062068 | 3.158e-05 | [12 empty] | 1.2716 | 0.75004 | 0.38848 | -0.023281 | 4.4338e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 436 | 1.6492 | [20 empty] | -0.091942 | 1.1931 | 0.81624 | -0.36446 | 0.12142 | -0.029983 | 0.0054844 | -0.00063074 | 3.2166e-05 | [12 empty] | 1.2716 | 0.75177 | 0.39026 | -0.023442 | 4.4747e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 437 | 1.654 | [20 empty] | -0.091942 | 1.1958 | 0.81998 | -0.36698 | 0.12254 | -0.030328 | 0.0055603 | -0.00064094 | 3.276e-05 | [12 empty] | 1.2716 | 0.75349 | 0.39206 | -0.023603 | 4.5159e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 438 | 1.6588 | [20 empty] | -0.091942 | 1.1985 | 0.82374 | -0.3695 | 0.12367 | -0.030677 | 0.0056371 | -0.00065127 | 3.3365e-05 | [12 empty] | 1.2716 | 0.75521 | 0.39385 | -0.023766 | 4.5574e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 439 | 1.6637 | [20 empty] | -0.091942 | 1.2013 | 0.82751 | -0.37204 | 0.1248 | -0.031028 | 0.0057148 | -0.00066175 | 3.3979e-05 | [12 empty] | 1.2716 | 0.75694 | 0.39565 | -0.023929 | 4.5991e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 440 | 1.6685 | [20 empty] | -0.091942 | 1.204 | 0.83128 | -0.37459 | 0.12594 | -0.031383 | 0.0057933 | -0.00067238 | 3.4604e-05 | [12 empty] | 1.2716 | 0.75866 | 0.39746 | -0.024093 | 4.6412e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 441 | 1.6733 | [20 empty] | -0.091942 | 1.2067 | 0.83506 | -0.37715 | 0.12709 | -0.031742 | 0.0058728 | -0.00068315 | 3.5238e-05 | [12 empty] | 1.2716 | 0.76039 | 0.39927 | -0.024258 | 4.6835e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 442 | 1.6781 | [20 empty] | -0.091942 | 1.2095 | 0.83886 | -0.37972 | 0.12825 | -0.032103 | 0.0059531 | -0.00069407 | 3.5882e-05 | [12 empty] | 1.2716 | 0.76211 | 0.40108 | -0.024423 | 4.7262e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 443 | 1.6829 | [20 empty] | -0.091942 | 1.2122 | 0.84266 | -0.3823 | 0.12941 | -0.032468 | 0.0060344 | -0.00070513 | 3.6537e-05 | [12 empty] | 1.2716 | 0.76384 | 0.4029 | -0.024589 | 4.7691e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 444 | 1.6878 | [20 empty] | -0.091942 | 1.215 | 0.84646 | -0.38489 | 0.13058 | -0.032836 | 0.0061166 | -0.00071635 | 3.7202e-05 | [12 empty] | 1.2716 | 0.76556 | 0.40472 | -0.024756 | 4.8123e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 445 | 1.6926 | [20 empty] | -0.091942 | 1.2177 | 0.85028 | -0.3875 | 0.13176 | -0.033208 | 0.0061997 | -0.00072772 | 3.7877e-05 | [12 empty] | 1.2716 | 0.76728 | 0.40654 | -0.024924 | 4.8558e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 446 | 1.6974 | [20 empty] | -0.091942 | 1.2204 | 0.85411 | -0.39012 | 0.13295 | -0.033582 | 0.0062838 | -0.00073924 | 3.8564e-05 | [12 empty] | 1.2716 | 0.76901 | 0.40837 | -0.025092 | 4.8996e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 447 | 1.7023 | [20 empty] | -0.091942 | 1.2232 | 0.85794 | -0.39275 | 0.13415 | -0.033961 | 0.0063688 | -0.00075093 | 3.9261e-05 | [12 empty] | 1.2716 | 0.77073 | 0.4102 | -0.025261 | 4.9437e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 448 | 1.7071 | [20 empty] | -0.091942 | 1.2259 | 0.86179 | -0.39539 | 0.13535 | -0.034342 | 0.0064548 | -0.00076276 | 3.9969e-05 | [12 empty] | 1.2716 | 0.77246 | 0.41204 | -0.025431 | 4.988e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 449 | 1.7119 | [20 empty] | -0.091942 | 1.2286 | 0.86564 | -0.39804 | 0.13656 | -0.034727 | 0.0065417 | -0.00077476 | 4.0688e-05 | [12 empty] | 1.2716 | 0.77418 | 0.41388 | -0.025602 | 5.0327e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 450 | 1.7168 | [20 empty] | -0.091942 | 1.2314 | 0.8695 | -0.40071 | 0.13779 | -0.035116 | 0.0066296 | -0.00078692 | 4.1419e-05 | [12 empty] | 1.2716 | 0.77591 | 0.41573 | -0.025773 | 5.0777e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 451 | 1.7216 | [20 empty] | -0.091942 | 1.2341 | 0.87337 | -0.40339 | 0.13901 | -0.035507 | 0.0067185 | -0.00079925 | 4.2161e-05 | [12 empty] | 1.2716 | 0.77763 | 0.41758 | -0.025945 | 5.123e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 452 | 1.7265 | [20 empty] | -0.091942 | 1.2368 | 0.87724 | -0.40608 | 0.14025 | -0.035903 | 0.0068084 | -0.00081173 | 4.2915e-05 | [12 empty] | 1.2716 | 0.77935 | 0.41943 | -0.026118 | 5.1686e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 453 | 1.7313 | [20 empty] | -0.091942 | 1.2396 | 0.88113 | -0.40878 | 0.1415 | -0.036302 | 0.0068992 | -0.00082439 | 4.368e-05 | [12 empty] | 1.2716 | 0.78108 | 0.42129 | -0.026292 | 5.2145e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 454 | 1.7362 | [20 empty] | -0.091942 | 1.2423 | 0.88502 | -0.41149 | 0.14275 | -0.036704 | 0.0069911 | -0.00083721 | 4.4458e-05 | [12 empty] | 1.2716 | 0.7828 | 0.42315 | -0.026467 | 5.2607e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 455 | 1.741 | [20 empty] | -0.091942 | 1.2451 | 0.88893 | -0.41422 | 0.14401 | -0.03711 | 0.007084 | -0.00085021 | 4.5247e-05 | [12 empty] | 1.2716 | 0.78453 | 0.42502 | -0.026642 | 5.3072e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 456 | 1.7459 | [20 empty] | -0.091942 | 1.2478 | 0.89284 | -0.41695 | 0.14528 | -0.03752 | 0.007178 | -0.00086337 | 4.6049e-05 | [12 empty] | 1.2716 | 0.78625 | 0.42689 | -0.026818 | 5.354e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 457 | 1.7507 | [20 empty] | -0.091942 | 1.2505 | 0.89676 | -0.4197 | 0.14656 | -0.037933 | 0.0072729 | -0.00087671 | 4.6863e-05 | [12 empty] | 1.2716 | 0.78798 | 0.42876 | -0.026995 | 5.4011e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 458 | 1.7556 | [20 empty] | -0.091942 | 1.2533 | 0.90069 | -0.42246 | 0.14785 | -0.03835 | 0.0073689 | -0.00089023 | 4.769e-05 | [12 empty] | 1.2716 | 0.7897 | 0.43064 | -0.027172 | 5.4485e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 459 | 1.7604 | [20 empty] | -0.091942 | 1.256 | 0.90462 | -0.42524 | 0.14914 | -0.03877 | 0.007466 | -0.00090393 | 4.8529e-05 | [12 empty] | 1.2716 | 0.79142 | 0.43252 | -0.027351 | 5.4963e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 460 | 1.7653 | [20 empty] | -0.091942 | 1.2587 | 0.90857 | -0.42802 | 0.15045 | -0.039195 | 0.0075641 | -0.0009178 | 4.9381e-05 | [12 empty] | 1.2716 | 0.79315 | 0.43441 | -0.02753 | 5.5443e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 461 | 1.7702 | [20 empty] | -0.091942 | 1.2615 | 0.91253 | -0.43082 | 0.15176 | -0.039622 | 0.0076633 | -0.00093186 | 5.0247e-05 | [12 empty] | 1.2716 | 0.79487 | 0.4363 | -0.02771 | 5.5927e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 462 | 1.775 | [20 empty] | -0.091942 | 1.2642 | 0.91649 | -0.43363 | 0.15308 | -0.040054 | 0.0077636 | -0.0009461 | 5.1125e-05 | [12 empty] | 1.2716 | 0.7966 | 0.4382 | -0.027891 | 5.6414e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 463 | 1.7799 | [20 empty] | -0.091942 | 1.2669 | 0.92046 | -0.43645 | 0.15441 | -0.040489 | 0.007865 | -0.00096053 | 5.2017e-05 | [12 empty] | 1.2716 | 0.79832 | 0.4401 | -0.028072 | 5.6904e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 464 | 1.7848 | [20 empty] | -0.091942 | 1.2697 | 0.92444 | -0.43929 | 0.15575 | -0.040929 | 0.0079675 | -0.00097515 | 5.2923e-05 | [12 empty] | 1.2716 | 0.80005 | 0.442 | -0.028254 | 5.7397e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 465 | 1.7896 | [20 empty] | -0.091942 | 1.2724 | 0.92843 | -0.44213 | 0.1571 | -0.041372 | 0.0080711 | -0.00098995 | 5.3842e-05 | [12 empty] | 1.2716 | 0.80177 | 0.44391 | -0.028437 | 5.7894e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 466 | 1.7945 | [20 empty] | -0.091942 | 1.2752 | 0.93243 | -0.44499 | 0.15845 | -0.041818 | 0.0081758 | -0.001005 | 5.4775e-05 | [12 empty] | 1.2716 | 0.80349 | 0.44582 | -0.028621 | 5.8393e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 467 | 1.7994 | [20 empty] | -0.091942 | 1.2779 | 0.93643 | -0.44786 | 0.15982 | -0.042269 | 0.0082816 | -0.0010201 | 5.5723e-05 | [12 empty] | 1.2716 | 0.80522 | 0.44773 | -0.028806 | 5.8896e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 468 | 1.8043 | [20 empty] | -0.091942 | 1.2806 | 0.94045 | -0.45074 | 0.16119 | -0.042723 | 0.0083886 | -0.0010355 | 5.6685e-05 | [12 empty] | 1.2716 | 0.80694 | 0.44965 | -0.028991 | 5.9402e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 469 | 1.8091 | [20 empty] | -0.091942 | 1.2834 | 0.94447 | -0.45364 | 0.16257 | -0.043182 | 0.0084967 | -0.0010511 | 5.7661e-05 | [12 empty] | 1.2716 | 0.80867 | 0.45158 | -0.029178 | 5.9911e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 470 | 1.814 | [20 empty] | -0.091942 | 1.2861 | 0.9485 | -0.45655 | 0.16396 | -0.043644 | 0.008606 | -0.0010669 | 5.8652e-05 | [12 empty] | 1.2716 | 0.81039 | 0.4535 | -0.029365 | 6.0424e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 471 | 1.8189 | [20 empty] | -0.091942 | 1.2888 | 0.95254 | -0.45947 | 0.16536 | -0.04411 | 0.0087164 | -0.0010829 | 5.9658e-05 | [12 empty] | 1.2716 | 0.81211 | 0.45544 | -0.029552 | 6.094e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 472 | 1.8238 | [20 empty] | -0.091942 | 1.2916 | 0.95659 | -0.4624 | 0.16677 | -0.044581 | 0.008828 | -0.0010991 | 6.0679e-05 | [12 empty] | 1.2716 | 0.81384 | 0.45737 | -0.029741 | 6.1459e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 473 | 1.8287 | [20 empty] | -0.091942 | 1.2943 | 0.96065 | -0.46535 | 0.16819 | -0.045055 | 0.0089409 | -0.0011155 | 6.1715e-05 | [12 empty] | 1.2716 | 0.81556 | 0.45931 | -0.029931 | 6.1982e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 474 | 1.8336 | [20 empty] | -0.091942 | 1.297 | 0.96472 | -0.4683 | 0.16962 | -0.045533 | 0.0090549 | -0.0011321 | 6.2766e-05 | [12 empty] | 1.2716 | 0.81729 | 0.46126 | -0.030121 | 6.2507e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 475 | 1.8385 | [20 empty] | -0.091942 | 1.2998 | 0.96879 | -0.47127 | 0.17105 | -0.046016 | 0.0091701 | -0.0011489 | 6.3833e-05 | [12 empty] | 1.2716 | 0.81901 | 0.4632 | -0.030312 | 6.3037e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 476 | 1.8434 | [20 empty] | -0.091942 | 1.3025 | 0.97287 | -0.47426 | 0.1725 | -0.046502 | 0.0092865 | -0.001166 | 6.4916e-05 | [12 empty] | 1.2716 | 0.82074 | 0.46516 | -0.030504 | 6.3569e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 477 | 1.8483 | [20 empty] | -0.091942 | 1.3053 | 0.97697 | -0.47725 | 0.17395 | -0.046993 | 0.0094042 | -0.0011832 | 6.6016e-05 | [12 empty] | 1.2716 | 0.82246 | 0.46711 | -0.030696 | 6.4105e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 478 | 1.8532 | [20 empty] | -0.091942 | 1.308 | 0.98107 | -0.48026 | 0.17541 | -0.047487 | 0.0095231 | -0.0012007 | 6.7131e-05 | [12 empty] | 1.2716 | 0.82418 | 0.46907 | -0.03089 | 6.4644e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 479 | 1.8581 | [20 empty] | -0.091942 | 1.3107 | 0.98518 | -0.48328 | 0.17689 | -0.047986 | 0.0096433 | -0.0012184 | 6.8263e-05 | [12 empty] | 1.2716 | 0.82591 | 0.47104 | -0.031084 | 6.5187e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 480 | 1.863 | [20 empty] | -0.091942 | 1.3135 | 0.98929 | -0.48631 | 0.17837 | -0.048489 | 0.0097647 | -0.0012363 | 6.9411e-05 | [12 empty] | 1.2716 | 0.82763 | 0.47301 | -0.031279 | 6.5733e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 481 | 1.8679 | [20 empty] | -0.091942 | 1.3162 | 0.99342 | -0.48936 | 0.17986 | -0.048996 | 0.0098874 | -0.0012545 | 7.0576e-05 | [12 empty] | 1.2716 | 0.82936 | 0.47498 | -0.031475 | 6.6282e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 482 | 1.8728 | [20 empty] | -0.091942 | 1.3189 | 0.99756 | -0.49242 | 0.18136 | -0.049508 | 0.010011 | -0.0012728 | 7.1759e-05 | [12 empty] | 1.2716 | 0.83108 | 0.47696 | -0.031672 | 6.6835e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 483 | 1.8777 | [20 empty] | -0.091942 | 1.3217 | 1.0017 | -0.49549 | 0.18287 | -0.050023 | 0.010137 | -0.0012914 | 7.2959e-05 | [12 empty] | 1.2716 | 0.83281 | 0.47894 | -0.031869 | 6.7392e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 484 | 1.8826 | [20 empty] | -0.091942 | 1.3244 | 1.0059 | -0.49857 | 0.18439 | -0.050543 | 0.010263 | -0.0013103 | 7.4176e-05 | [12 empty] | 1.2716 | 0.83453 | 0.48092 | -0.032068 | 6.7952e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 485 | 1.8875 | [20 empty] | -0.091942 | 1.3271 | 1.01 | -0.50167 | 0.18592 | -0.051068 | 0.010391 | -0.0013293 | 7.5411e-05 | [12 empty] | 1.2716 | 0.83625 | 0.48291 | -0.032267 | 6.8515e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 486 | 1.8925 | [20 empty] | -0.091942 | 1.3299 | 1.0142 | -0.50478 | 0.18745 | -0.051596 | 0.01052 | -0.0013486 | 7.6664e-05 | [12 empty] | 1.2716 | 0.83798 | 0.48491 | -0.032467 | 6.9082e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 487 | 1.8974 | [20 empty] | -0.091942 | 1.3326 | 1.0184 | -0.5079 | 0.189 | -0.052129 | 0.010651 | -0.0013682 | 7.7935e-05 | [12 empty] | 1.2716 | 0.8397 | 0.4869 | -0.032668 | 6.9652e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 488 | 1.9023 | [20 empty] | -0.091942 | 1.3354 | 1.0225 | -0.51104 | 0.19056 | -0.052667 | 0.010783 | -0.001388 | 7.9224e-05 | [12 empty] | 1.2716 | 0.84143 | 0.48891 | -0.032869 | 7.0226e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 489 | 1.9072 | [20 empty] | -0.091942 | 1.3381 | 1.0267 | -0.51419 | 0.19213 | -0.053209 | 0.010916 | -0.001408 | 8.0532e-05 | [12 empty] | 1.2716 | 0.84315 | 0.49091 | -0.033072 | 7.0803e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 490 | 1.9121 | [20 empty] | -0.091942 | 1.3408 | 1.0309 | -0.51735 | 0.1937 | -0.053755 | 0.011051 | -0.0014283 | 8.1859e-05 | [12 empty] | 1.2716 | 0.84488 | 0.49292 | -0.033275 | 7.1384e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 491 | 1.9171 | [20 empty] | -0.091942 | 1.3436 | 1.0352 | -0.52052 | 0.19529 | -0.054306 | 0.011187 | -0.0014488 | 8.3205e-05 | [12 empty] | 1.2716 | 0.8466 | 0.49494 | -0.033479 | 7.1969e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 492 | 1.922 | [20 empty] | -0.091942 | 1.3463 | 1.0394 | -0.52371 | 0.19688 | -0.054861 | 0.011324 | -0.0014696 | 8.4571e-05 | [12 empty] | 1.2716 | 0.84832 | 0.49695 | -0.033684 | 7.2557e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 493 | 1.9269 | [20 empty] | -0.091942 | 1.349 | 1.0436 | -0.52691 | 0.19849 | -0.055421 | 0.011463 | -0.0014906 | 8.5956e-05 | [12 empty] | 1.2716 | 0.85005 | 0.49898 | -0.03389 | 7.3149e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 494 | 1.9319 | [20 empty] | -0.091942 | 1.3518 | 1.0478 | -0.53012 | 0.20011 | -0.055985 | 0.011603 | -0.0015119 | 8.736e-05 | [12 empty] | 1.2716 | 0.85177 | 0.501 | -0.034097 | 7.3744e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 495 | 1.9368 | [20 empty] | -0.091942 | 1.3545 | 1.0521 | -0.53335 | 0.20173 | -0.056554 | 0.011745 | -0.0015335 | 8.8785e-05 | [12 empty] | 1.2716 | 0.8535 | 0.50303 | -0.034304 | 7.4343e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 496 | 1.9417 | [20 empty] | -0.091942 | 1.3572 | 1.0563 | -0.53658 | 0.20337 | -0.057128 | 0.011888 | -0.0015553 | 9.023e-05 | [12 empty] | 1.2716 | 0.85522 | 0.50507 | -0.034512 | 7.4945e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 497 | 1.9467 | [20 empty] | -0.091942 | 1.36 | 1.0606 | -0.53984 | 0.20501 | -0.057706 | 0.012032 | -0.0015774 | 9.1696e-05 | [12 empty] | 1.2716 | 0.85694 | 0.50711 | -0.034722 | 7.5552e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 498 | 1.9516 | [20 empty] | -0.091942 | 1.3627 | 1.0649 | -0.5431 | 0.20667 | -0.058289 | 0.012178 | -0.0015997 | 9.3182e-05 | [12 empty] | 1.2716 | 0.85867 | 0.50915 | -0.034932 | 7.6162e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 499 | 1.9566 | [20 empty] | -0.091942 | 1.3655 | 1.0692 | -0.54638 | 0.20833 | -0.058876 | 0.012326 | -0.0016224 | 9.469e-05 | [12 empty] | 1.2716 | 0.86039 | 0.5112 | -0.035143 | 7.6775e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 500 | 1.9615 | [20 empty] | -0.091942 | 1.3682 | 1.0735 | -0.54967 | 0.21001 | -0.059469 | 0.012475 | -0.0016453 | 9.6219e-05 | [12 empty] | 1.2716 | 0.86212 | 0.51325 | -0.035354 | 7.7392e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 501 | 1.9664 | [20 empty] | -0.091942 | 1.3709 | 1.0778 | -0.55298 | 0.21169 | -0.060066 | 0.012625 | -0.0016684 | 9.7769e-05 | [12 empty] | 1.2716 | 0.86384 | 0.5153 | -0.035567 | 7.8013e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 502 | 1.9714 | [20 empty] | -0.091942 | 1.3737 | 1.0821 | -0.55629 | 0.21339 | -0.060667 | 0.012777 | -0.0016919 | 9.9341e-05 | [12 empty] | 1.2716 | 0.86557 | 0.51736 | -0.03578 | 7.8638e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 503 | 1.9763 | [20 empty] | -0.091942 | 1.3764 | 1.0864 | -0.55962 | 0.21509 | -0.061274 | 0.012931 | -0.0017156 | 0.00010094 | [12 empty] | 1.2716 | 0.86729 | 0.51942 | -0.035994 | 7.9267e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 504 | 1.9813 | [20 empty] | -0.091942 | 1.3791 | 1.0907 | -0.56297 | 0.21681 | -0.061886 | 0.013086 | -0.0017396 | 0.00010255 | [12 empty] | 1.2716 | 0.86901 | 0.52149 | -0.03621 | 7.9899e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 505 | 1.9862 | [20 empty] | -0.091942 | 1.3819 | 1.095 | -0.56633 | 0.21853 | -0.062502 | 0.013242 | -0.0017639 | 0.00010419 | [12 empty] | 1.2716 | 0.87074 | 0.52356 | -0.036425 | 8.0535e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 506 | 1.9912 | [20 empty] | -0.091942 | 1.3846 | 1.0994 | -0.5697 | 0.22027 | -0.063123 | 0.0134 | -0.0017885 | 0.00010585 | [12 empty] | 1.2716 | 0.87246 | 0.52564 | -0.036642 | 8.1175e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 507 | 1.9962 | [20 empty] | -0.091942 | 1.3873 | 1.1037 | -0.57308 | 0.22202 | -0.063749 | 0.01356 | -0.0018134 | 0.00010754 | [12 empty] | 1.2716 | 0.87419 | 0.52772 | -0.03686 | 8.1818e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 508 | 2.0011 | [20 empty] | -0.091942 | 1.3901 | 1.1081 | -0.57648 | 0.22377 | -0.064381 | 0.013721 | -0.0018386 | 0.00010925 | [12 empty] | 1.2716 | 0.87591 | 0.5298 | -0.037079 | 8.2466e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 509 | 2.0061 | [20 empty] | -0.091942 | 1.3928 | 1.1124 | -0.57989 | 0.22554 | -0.065017 | 0.013884 | -0.0018641 | 0.00011098 | [12 empty] | 1.2716 | 0.87764 | 0.53189 | -0.037298 | 8.3117e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 510 | 2.0111 | [20 empty] | -0.091942 | 1.3956 | 1.1168 | -0.58332 | 0.22732 | -0.065658 | 0.014049 | -0.0018899 | 0.00011274 | [12 empty] | 1.2716 | 0.87936 | 0.53398 | -0.037518 | 8.3772e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 511 | 2.016 | [20 empty] | -0.091942 | 1.3983 | 1.1212 | -0.58675 | 0.22911 | -0.066304 | 0.014215 | -0.001916 | 0.00011452 | [12 empty] | 1.2716 | 0.88108 | 0.53608 | -0.037739 | 8.4431e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 512 | 2.021 | [20 empty] | -0.091942 | 1.401 | 1.1256 | -0.5902 | 0.2309 | -0.066956 | 0.014382 | -0.0019424 | 0.00011632 | [12 empty] | 1.2716 | 0.88281 | 0.53818 | -0.037961 | 8.5094e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 513 | 2.026 | [20 empty] | -0.091942 | 1.4038 | 1.13 | -0.59367 | 0.23271 | -0.067612 | 0.014552 | -0.0019691 | 0.00011815 | [12 empty] | 1.2716 | 0.88453 | 0.54028 | -0.038184 | 8.5761e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 514 | 2.0309 | [20 empty] | -0.091942 | 1.4065 | 1.1344 | -0.59715 | 0.23453 | -0.068274 | 0.014723 | -0.0019961 | 0.00012001 | [12 empty] | 1.2716 | 0.88626 | 0.54239 | -0.038408 | 8.6431e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 515 | 2.0359 | [20 empty] | -0.091942 | 1.4092 | 1.1388 | -0.60064 | 0.23636 | -0.06894 | 0.014896 | -0.0020235 | 0.00012189 | [12 empty] | 1.2716 | 0.88798 | 0.5445 | -0.038633 | 8.7106e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 516 | 2.0409 | [20 empty] | -0.091942 | 1.412 | 1.1433 | -0.60415 | 0.2382 | -0.069612 | 0.01507 | -0.0020511 | 0.00012379 | [12 empty] | 1.2716 | 0.88971 | 0.54662 | -0.038858 | 8.7784e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 517 | 2.0459 | [20 empty] | -0.091942 | 1.4147 | 1.1477 | -0.60767 | 0.24006 | -0.070289 | 0.015246 | -0.0020791 | 0.00012573 | [12 empty] | 1.2716 | 0.89143 | 0.54874 | -0.039084 | 8.8467e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 518 | 2.0508 | [20 empty] | -0.091942 | 1.4174 | 1.1521 | -0.6112 | 0.24192 | -0.070972 | 0.015424 | -0.0021074 | 0.00012768 | [12 empty] | 1.2716 | 0.89315 | 0.55086 | -0.039312 | 8.9153e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 519 | 2.0558 | [20 empty] | -0.091942 | 1.4202 | 1.1566 | -0.61474 | 0.24379 | -0.071659 | 0.015603 | -0.0021361 | 0.00012967 | [12 empty] | 1.2716 | 0.89488 | 0.55299 | -0.03954 | 8.9844e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 520 | 2.0608 | [20 empty] | -0.091942 | 1.4229 | 1.161 | -0.61831 | 0.24568 | -0.072353 | 0.015785 | -0.002165 | 0.00013168 | [12 empty] | 1.2716 | 0.8966 | 0.55513 | -0.039769 | 9.0538e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 521 | 2.0658 | [20 empty] | -0.091942 | 1.4257 | 1.1655 | -0.62188 | 0.24757 | -0.073051 | 0.015968 | -0.0021944 | 0.00013372 | [12 empty] | 1.2716 | 0.89833 | 0.55726 | -0.039999 | 9.1237e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 522 | 2.0708 | [20 empty] | -0.091942 | 1.4284 | 1.17 | -0.62547 | 0.24948 | -0.073755 | 0.016152 | -0.002224 | 0.00013579 | [12 empty] | 1.2716 | 0.90005 | 0.55941 | -0.040229 | 9.1939e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 523 | 2.0758 | [20 empty] | -0.091942 | 1.4311 | 1.1745 | -0.62907 | 0.2514 | -0.074464 | 0.016339 | -0.002254 | 0.00013788 | [12 empty] | 1.2716 | 0.90177 | 0.56155 | -0.040461 | 9.2646e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 524 | 2.0807 | [20 empty] | -0.091942 | 1.4339 | 1.179 | -0.63268 | 0.25332 | -0.075178 | 0.016527 | -0.0022844 | 0.00014001 | [12 empty] | 1.2716 | 0.9035 | 0.5637 | -0.040694 | 9.3356e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 525 | 2.0857 | [20 empty] | -0.091942 | 1.4366 | 1.1835 | -0.63631 | 0.25526 | -0.075899 | 0.016717 | -0.002315 | 0.00014216 | [12 empty] | 1.2716 | 0.90522 | 0.56585 | -0.040927 | 9.4071e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 526 | 2.0907 | [20 empty] | -0.091942 | 1.4393 | 1.188 | -0.63996 | 0.25721 | -0.076624 | 0.016909 | -0.0023461 | 0.00014434 | [12 empty] | 1.2716 | 0.90695 | 0.56801 | -0.041161 | 9.479e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 527 | 2.0957 | [20 empty] | -0.091942 | 1.4421 | 1.1925 | -0.64361 | 0.25918 | -0.077355 | 0.017103 | -0.0023775 | 0.00014655 | [12 empty] | 1.2716 | 0.90867 | 0.57017 | -0.041396 | 9.5513e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 528 | 2.1007 | [20 empty] | -0.091942 | 1.4448 | 1.197 | -0.64728 | 0.26115 | -0.078092 | 0.017299 | -0.0024092 | 0.00014879 | [12 empty] | 1.2716 | 0.9104 | 0.57234 | -0.041633 | 9.624e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 529 | 2.1057 | [20 empty] | -0.091942 | 1.4475 | 1.2016 | -0.65097 | 0.26313 | -0.078834 | 0.017496 | -0.0024414 | 0.00015106 | [12 empty] | 1.2716 | 0.91212 | 0.57451 | -0.04187 | 9.6971e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 530 | 2.1107 | [20 empty] | -0.091942 | 1.4503 | 1.2061 | -0.65467 | 0.26513 | -0.079582 | 0.017696 | -0.0024739 | 0.00015336 | [12 empty] | 1.2716 | 0.91384 | 0.57668 | -0.042107 | 9.7706e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 531 | 2.1157 | [20 empty] | -0.091942 | 1.453 | 1.2107 | -0.65838 | 0.26713 | -0.080336 | 0.017897 | -0.0025067 | 0.00015569 | [12 empty] | 1.2716 | 0.91557 | 0.57886 | -0.042346 | 9.8446e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 532 | 2.1207 | [20 empty] | -0.091942 | 1.4558 | 1.2153 | -0.66211 | 0.26915 | -0.081095 | 0.0181 | -0.00254 | 0.00015805 | [12 empty] | 1.2716 | 0.91729 | 0.58104 | -0.042586 | 9.9189e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 533 | 2.1257 | [20 empty] | -0.091942 | 1.4585 | 1.2198 | -0.66585 | 0.27118 | -0.08186 | 0.018305 | -0.0025736 | 0.00016044 | [12 empty] | 1.2716 | 0.91902 | 0.58323 | -0.042827 | 9.9937e-05
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 534 | 2.1308 | [20 empty] | -0.091942 | 1.4612 | 1.2244 | -0.6696 | 0.27322 | -0.082631 | 0.018512 | -0.0026076 | 0.00016287 | [12 empty] | 1.2716 | 0.92074 | 0.58542 | -0.043068 | 0.00010069
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 535 | 2.1358 | [20 empty] | -0.091942 | 1.464 | 1.229 | -0.67337 | 0.27528 | -0.083408 | 0.018721 | -0.0026419 | 0.00016532 | [12 empty] | 1.2716 | 0.92247 | 0.58762 | -0.04331 | 0.00010145
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 536 | 2.1408 | [20 empty] | -0.091942 | 1.4667 | 1.2336 | -0.67715 | 0.27734 | -0.08419 | 0.018932 | -0.0026767 | 0.00016781 | [12 empty] | 1.2716 | 0.92419 | 0.58981 | -0.043554 | 0.00010221
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 537 | 2.1458 | [20 empty] | -0.091942 | 1.4694 | 1.2382 | -0.68095 | 0.27941 | -0.084978 | 0.019145 | -0.0027118 | 0.00017033 | [12 empty] | 1.2716 | 0.92591 | 0.59202 | -0.043798 | 0.00010297
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 538 | 2.1508 | [20 empty] | -0.091942 | 1.4722 | 1.2428 | -0.68476 | 0.2815 | -0.085773 | 0.01936 | -0.0027474 | 0.00017288 | [12 empty] | 1.2716 | 0.92764 | 0.59422 | -0.044043 | 0.00010374
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 539 | 2.1558 | [20 empty] | -0.091942 | 1.4749 | 1.2474 | -0.68859 | 0.2836 | -0.086573 | 0.019577 | -0.0027833 | 0.00017547 | [12 empty] | 1.2716 | 0.92936 | 0.59643 | -0.044289 | 0.00010451
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 540 | 2.1608 | [20 empty] | -0.091942 | 1.4776 | 1.2521 | -0.69243 | 0.28571 | -0.087379 | 0.019796 | -0.0028197 | 0.00017809 | [12 empty] | 1.2716 | 0.93109 | 0.59865 | -0.044536 | 0.00010529
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 541 | 2.1659 | [20 empty] | -0.091942 | 1.4804 | 1.2567 | -0.69628 | 0.28783 | -0.088191 | 0.020017 | -0.0028564 | 0.00018075 | [12 empty] | 1.2716 | 0.93281 | 0.60087 | -0.044784 | 0.00010607
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 542 | 2.1709 | [20 empty] | -0.091942 | 1.4831 | 1.2614 | -0.70015 | 0.28997 | -0.089009 | 0.02024 | -0.0028936 | 0.00018344 | [12 empty] | 1.2716 | 0.93454 | 0.60309 | -0.045033 | 0.00010686
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 543 | 2.1759 | [20 empty] | -0.091942 | 1.4859 | 1.266 | -0.70403 | 0.29211 | -0.089833 | 0.020465 | -0.0029312 | 0.00018616 | [12 empty] | 1.2716 | 0.93626 | 0.60532 | -0.045283 | 0.00010765
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 544 | 2.1809 | [20 empty] | -0.091942 | 1.4886 | 1.2707 | -0.70793 | 0.29427 | -0.090663 | 0.020692 | -0.0029692 | 0.00018893 | [12 empty] | 1.2716 | 0.93798 | 0.60755 | -0.045533 | 0.00010845
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 545 | 2.186 | [20 empty] | -0.091942 | 1.4913 | 1.2754 | -0.71184 | 0.29644 | -0.0915 | 0.020921 | -0.0030076 | 0.00019172 | [12 empty] | 1.2716 | 0.93971 | 0.60979 | -0.045785 | 0.00010925
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 546 | 2.191 | [20 empty] | -0.091942 | 1.4941 | 1.2801 | -0.71577 | 0.29862 | -0.092342 | 0.021153 | -0.0030464 | 0.00019455 | [12 empty] | 1.2716 | 0.94143 | 0.61203 | -0.046037 | 0.00011005
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 547 | 2.196 | [20 empty] | -0.091942 | 1.4968 | 1.2847 | -0.71971 | 0.30082 | -0.093191 | 0.021386 | -0.0030857 | 0.00019742 | [12 empty] | 1.2716 | 0.94316 | 0.61427 | -0.046291 | 0.00011086
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 548 | 2.2011 | [20 empty] | -0.091942 | 1.4995 | 1.2894 | -0.72366 | 0.30302 | -0.094046 | 0.021622 | -0.0031254 | 0.00020033 | [12 empty] | 1.2716 | 0.94488 | 0.61652 | -0.046545 | 0.00011167
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 549 | 2.2061 | [20 empty] | -0.091942 | 1.5023 | 1.2942 | -0.72763 | 0.30524 | -0.094907 | 0.02186 | -0.0031656 | 0.00020327 | [12 empty] | 1.2716 | 0.94661 | 0.61877 | -0.0468 | 0.00011249
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 550 | 2.2111 | [20 empty] | -0.091942 | 1.505 | 1.2989 | -0.73161 | 0.30747 | -0.095775 | 0.0221 | -0.0032061 | 0.00020625 | [12 empty] | 1.2716 | 0.94833 | 0.62103 | -0.047056 | 0.00011331
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 551 | 2.2162 | [20 empty] | -0.091942 | 1.5077 | 1.3036 | -0.73561 | 0.30971 | -0.096648 | 0.022342 | -0.0032472 | 0.00020927 | [12 empty] | 1.2716 | 0.95005 | 0.62329 | -0.047314 | 0.00011414
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 552 | 2.2212 | [20 empty] | -0.091942 | 1.5105 | 1.3083 | -0.73962 | 0.31197 | -0.097529 | 0.022586 | -0.0032887 | 0.00021233 | [12 empty] | 1.2716 | 0.95178 | 0.62555 | -0.047572 | 0.00011497
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 553 | 2.2263 | [20 empty] | -0.091942 | 1.5132 | 1.3131 | -0.74365 | 0.31423 | -0.098415 | 0.022833 | -0.0033306 | 0.00021543 | [12 empty] | 1.2716 | 0.9535 | 0.62782 | -0.047831 | 0.0001158
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 554 | 2.2313 | [20 empty] | -0.091942 | 1.516 | 1.3178 | -0.74769 | 0.31651 | -0.099308 | 0.023082 | -0.003373 | 0.00021856 | [12 empty] | 1.2716 | 0.95523 | 0.63009 | -0.048091 | 0.00011664
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 555 | 2.2363 | [20 empty] | -0.091942 | 1.5187 | 1.3226 | -0.75175 | 0.3188 | -0.10021 | 0.023333 | -0.0034158 | 0.00022174 | [12 empty] | 1.2716 | 0.95695 | 0.63237 | -0.048352 | 0.00011749
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 556 | 2.2414 | [20 empty] | -0.091942 | 1.5214 | 1.3274 | -0.75582 | 0.32111 | -0.10111 | 0.023586 | -0.0034591 | 0.00022496 | [12 empty] | 1.2716 | 0.95867 | 0.63465 | -0.048613 | 0.00011834
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 557 | 2.2464 | [20 empty] | -0.091942 | 1.5242 | 1.3322 | -0.7599 | 0.32342 | -0.10203 | 0.023842 | -0.0035029 | 0.00022821 | [12 empty] | 1.2716 | 0.9604 | 0.63694 | -0.048876 | 0.00011919
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 558 | 2.2515 | [20 empty] | -0.091942 | 1.5269 | 1.3369 | -0.764 | 0.32575 | -0.10295 | 0.0241 | -0.0035472 | 0.00023151 | [12 empty] | 1.2716 | 0.96212 | 0.63923 | -0.04914 | 0.00012005
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 559 | 2.2565 | [20 empty] | -0.091942 | 1.5296 | 1.3417 | -0.76812 | 0.32809 | -0.10387 | 0.02436 | -0.0035919 | 0.00023485 | [12 empty] | 1.2716 | 0.96385 | 0.64152 | -0.049405 | 0.00012091
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 560 | 2.2616 | [20 empty] | -0.091942 | 1.5324 | 1.3465 | -0.77225 | 0.33045 | -0.1048 | 0.024623 | -0.0036371 | 0.00023823 | [12 empty] | 1.2716 | 0.96557 | 0.64382 | -0.04967 | 0.00012178
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 561 | 2.2667 | [20 empty] | -0.091942 | 1.5351 | 1.3514 | -0.77639 | 0.33282 | -0.10574 | 0.024888 | -0.0036829 | 0.00024166 | [12 empty] | 1.2716 | 0.9673 | 0.64612 | -0.049937 | 0.00012265
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 562 | 2.2717 | [20 empty] | -0.091942 | 1.5378 | 1.3562 | -0.78055 | 0.33519 | -0.10669 | 0.025155 | -0.0037291 | 0.00024513 | [12 empty] | 1.2716 | 0.96902 | 0.64842 | -0.050204 | 0.00012353
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 563 | 2.2768 | [20 empty] | -0.091942 | 1.5406 | 1.361 | -0.78473 | 0.33759 | -0.10764 | 0.025425 | -0.0037757 | 0.00024864 | [12 empty] | 1.2716 | 0.97074 | 0.65073 | -0.050473 | 0.00012441
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 564 | 2.2818 | [20 empty] | -0.091942 | 1.5433 | 1.3658 | -0.78892 | 0.33999 | -0.1086 | 0.025697 | -0.0038229 | 0.00025219 | [12 empty] | 1.2716 | 0.97247 | 0.65305 | -0.050742 | 0.0001253
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 565 | 2.2869 | [20 empty] | -0.091942 | 1.5461 | 1.3707 | -0.79312 | 0.34241 | -0.10957 | 0.025972 | -0.0038706 | 0.00025579 | [12 empty] | 1.2716 | 0.97419 | 0.65536 | -0.051013 | 0.00012619
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 566 | 2.292 | [20 empty] | -0.091942 | 1.5488 | 1.3755 | -0.79734 | 0.34484 | -0.11054 | 0.026249 | -0.0039189 | 0.00025944 | [12 empty] | 1.2716 | 0.97592 | 0.65769 | -0.051284 | 0.00012708
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 567 | 2.297 | [20 empty] | -0.091942 | 1.5515 | 1.3804 | -0.80157 | 0.34728 | -0.11152 | 0.026528 | -0.0039676 | 0.00026313 | [12 empty] | 1.2716 | 0.97764 | 0.66001 | -0.051556 | 0.00012798
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 568 | 2.3021 | [20 empty] | -0.091942 | 1.5543 | 1.3853 | -0.80582 | 0.34974 | -0.11251 | 0.02681 | -0.0040168 | 0.00026686 | [12 empty] | 1.2716 | 0.97937 | 0.66234 | -0.051829 | 0.00012889
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 569 | 2.3072 | [20 empty] | -0.091942 | 1.557 | 1.3902 | -0.81008 | 0.35221 | -0.1135 | 0.027095 | -0.0040666 | 0.00027064 | [12 empty] | 1.2716 | 0.98109 | 0.66468 | -0.052104 | 0.0001298
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 570 | 2.3122 | [20 empty] | -0.091942 | 1.5597 | 1.3951 | -0.81436 | 0.35469 | -0.1145 | 0.027382 | -0.0041169 | 0.00027447 | [12 empty] | 1.2716 | 0.98281 | 0.66701 | -0.052379 | 0.00013071
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 571 | 2.3173 | [20 empty] | -0.091942 | 1.5625 | 1.4 | -0.81866 | 0.35719 | -0.11551 | 0.027671 | -0.0041677 | 0.00027835 | [12 empty] | 1.2716 | 0.98454 | 0.66936 | -0.052655 | 0.00013163
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 572 | 2.3224 | [20 empty] | -0.091942 | 1.5652 | 1.4049 | -0.82296 | 0.3597 | -0.11652 | 0.027963 | -0.0042191 | 0.00028227 | [12 empty] | 1.2716 | 0.98626 | 0.6717 | -0.052932 | 0.00013256
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 573 | 2.3274 | [20 empty] | -0.091942 | 1.5679 | 1.4098 | -0.82729 | 0.36222 | -0.11755 | 0.028258 | -0.004271 | 0.00028624 | [12 empty] | 1.2716 | 0.98799 | 0.67405 | -0.05321 | 0.00013349
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 574 | 2.3325 | [20 empty] | -0.091942 | 1.5707 | 1.4147 | -0.83163 | 0.36475 | -0.11858 | 0.028555 | -0.0043234 | 0.00029026 | [12 empty] | 1.2716 | 0.98971 | 0.67641 | -0.053489 | 0.00013442
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 575 | 2.3376 | [20 empty] | -0.091942 | 1.5734 | 1.4196 | -0.83598 | 0.3673 | -0.11961 | 0.028855 | -0.0043764 | 0.00029433 | [12 empty] | 1.2716 | 0.99144 | 0.67877 | -0.053769 | 0.00013536
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 576 | 2.3427 | [20 empty] | -0.091942 | 1.5762 | 1.4246 | -0.84035 | 0.36986 | -0.12066 | 0.029157 | -0.00443 | 0.00029845 | [12 empty] | 1.2716 | 0.99316 | 0.68113 | -0.05405 | 0.0001363
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 577 | 2.3478 | [20 empty] | -0.091942 | 1.5789 | 1.4295 | -0.84473 | 0.37244 | -0.12171 | 0.029462 | -0.0044841 | 0.00030263 | [12 empty] | 1.2716 | 0.99488 | 0.6835 | -0.054332 | 0.00013725
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 578 | 2.3528 | [20 empty] | -0.091942 | 1.5816 | 1.4345 | -0.84913 | 0.37503 | -0.12277 | 0.02977 | -0.0045388 | 0.00030685 | [12 empty] | 1.2716 | 0.99661 | 0.68587 | -0.054615 | 0.00013821
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 579 | 2.3579 | [20 empty] | -0.091942 | 1.5844 | 1.4395 | -0.85355 | 0.37763 | -0.12383 | 0.03008 | -0.004594 | 0.00031112 | [12 empty] | 1.2716 | 0.99833 | 0.68824 | -0.054899 | 0.00013917
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 580 | 2.363 | [20 empty] | -0.091942 | 1.5871 | 1.4444 | -0.85798 | 0.38025 | -0.1249 | 0.030393 | -0.0046499 | 0.00031544 | [12 empty] | 1.2716 | 1.0001 | 0.69062 | -0.055184 | 0.00014013
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 581 | 2.3681 | [20 empty] | -0.091942 | 1.5898 | 1.4494 | -0.86242 | 0.38287 | -0.12598 | 0.030709 | -0.0047063 | 0.00031982 | [12 empty] | 1.2716 | 1.0018 | 0.69301 | -0.05547 | 0.0001411
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 582 | 2.3732 | [20 empty] | -0.091942 | 1.5926 | 1.4544 | -0.86689 | 0.38552 | -0.12707 | 0.031028 | -0.0047633 | 0.00032425 | [12 empty] | 1.2716 | 1.0035 | 0.69539 | -0.055757 | 0.00014207
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 583 | 2.3783 | [20 empty] | -0.091942 | 1.5953 | 1.4594 | -0.87136 | 0.38817 | -0.12817 | 0.031349 | -0.0048209 | 0.00032874 | [12 empty] | 1.2716 | 1.0052 | 0.69779 | -0.056045 | 0.00014305
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 584 | 2.3834 | [20 empty] | -0.091942 | 1.5981 | 1.4644 | -0.87585 | 0.39084 | -0.12927 | 0.031673 | -0.004879 | 0.00033327 | [12 empty] | 1.2716 | 1.007 | 0.70018 | -0.056334 | 0.00014404
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 585 | 2.3885 | [20 empty] | -0.091942 | 1.6008 | 1.4694 | -0.88036 | 0.39353 | -0.13038 | 0.032 | -0.0049378 | 0.00033787 | [12 empty] | 1.2716 | 1.0087 | 0.70258 | -0.056624 | 0.00014502
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 586 | 2.3936 | [20 empty] | -0.091942 | 1.6035 | 1.4745 | -0.88488 | 0.39623 | -0.1315 | 0.032329 | -0.0049972 | 0.00034252 | [12 empty] | 1.2716 | 1.0104 | 0.70499 | -0.056915 | 0.00014602
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 587 | 2.3987 | [20 empty] | -0.091942 | 1.6063 | 1.4795 | -0.88942 | 0.39894 | -0.13263 | 0.032662 | -0.0050572 | 0.00034722 | [12 empty] | 1.2716 | 1.0121 | 0.70739 | -0.057207 | 0.00014702
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 588 | 2.4038 | [20 empty] | -0.091942 | 1.609 | 1.4846 | -0.89397 | 0.40166 | -0.13376 | 0.032997 | -0.0051178 | 0.00035198 | [12 empty] | 1.2716 | 1.0139 | 0.70981 | -0.057499 | 0.00014802
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 589 | 2.4089 | [20 empty] | -0.091942 | 1.6117 | 1.4896 | -0.89854 | 0.4044 | -0.1349 | 0.033335 | -0.0051791 | 0.0003568 | [12 empty] | 1.2716 | 1.0156 | 0.71222 | -0.057793 | 0.00014903
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 590 | 2.414 | [20 empty] | -0.091942 | 1.6145 | 1.4947 | -0.90313 | 0.40716 | -0.13605 | 0.033676 | -0.0052409 | 0.00036167 | [12 empty] | 1.2716 | 1.0173 | 0.71464 | -0.058088 | 0.00015005
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 591 | 2.4191 | [20 empty] | -0.091942 | 1.6172 | 1.4997 | -0.90773 | 0.40992 | -0.13721 | 0.03402 | -0.0053034 | 0.00036661 | [12 empty] | 1.2716 | 1.019 | 0.71707 | -0.058384 | 0.00015107
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 592 | 2.4242 | [20 empty] | -0.091942 | 1.6199 | 1.5048 | -0.91234 | 0.4127 | -0.13837 | 0.034367 | -0.0053666 | 0.0003716 | [12 empty] | 1.2716 | 1.0207 | 0.7195 | -0.058681 | 0.00015209
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 593 | 2.4293 | [20 empty] | -0.091942 | 1.6227 | 1.5099 | -0.91697 | 0.4155 | -0.13954 | 0.034717 | -0.0054303 | 0.00037665 | [12 empty] | 1.2716 | 1.0225 | 0.72193 | -0.058979 | 0.00015312
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 594 | 2.4344 | [20 empty] | -0.091942 | 1.6254 | 1.515 | -0.92162 | 0.41831 | -0.14072 | 0.03507 | -0.0054948 | 0.00038176 | [12 empty] | 1.2716 | 1.0242 | 0.72437 | -0.059278 | 0.00015416
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 595 | 2.4395 | [20 empty] | -0.091942 | 1.6282 | 1.5201 | -0.92628 | 0.42113 | -0.14191 | 0.035425 | -0.0055599 | 0.00038693 | [12 empty] | 1.2716 | 1.0259 | 0.72681 | -0.059577 | 0.0001552
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 596 | 2.4446 | [20 empty] | -0.091942 | 1.6309 | 1.5252 | -0.93096 | 0.42397 | -0.14311 | 0.035784 | -0.0056256 | 0.00039217 | [12 empty] | 1.2716 | 1.0276 | 0.72925 | -0.059878 | 0.00015624
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 597 | 2.4497 | [20 empty] | -0.091942 | 1.6336 | 1.5304 | -0.93566 | 0.42682 | -0.14431 | 0.036146 | -0.005692 | 0.00039746 | [12 empty] | 1.2716 | 1.0294 | 0.7317 | -0.06018 | 0.0001573
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 598 | 2.4549 | [20 empty] | -0.091942 | 1.6364 | 1.5355 | -0.94037 | 0.42969 | -0.14553 | 0.036511 | -0.0057591 | 0.00040282 | [12 empty] | 1.2716 | 1.0311 | 0.73416 | -0.060483 | 0.00015835
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 599 | 2.46 | [20 empty] | -0.091942 | 1.6391 | 1.5406 | -0.94509 | 0.43257 | -0.14675 | 0.036879 | -0.0058268 | 0.00040824 | [12 empty] | 1.2716 | 1.0328 | 0.73661 | -0.060787 | 0.00015941
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 600 | 2.4651 | [20 empty] | -0.091942 | 1.6418 | 1.5458 | -0.94983 | 0.43547 | -0.14798 | 0.037249 | -0.0058953 | 0.00041372 | [12 empty] | 1.2716 | 1.0345 | 0.73907 | -0.061092 | 0.00016048
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 601 | 2.4702 | [20 empty] | -0.091942 | 1.6446 | 1.5509 | -0.95459 | 0.43838 | -0.14921 | 0.037624 | -0.0059644 | 0.00041927 | [12 empty] | 1.2716 | 1.0363 | 0.74154 | -0.061398 | 0.00016155
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 602 | 2.4753 | [20 empty] | -0.091942 | 1.6473 | 1.5561 | -0.95936 | 0.4413 | -0.15046 | 0.038001 | -0.0060342 | 0.00042488 | [12 empty] | 1.2716 | 1.038 | 0.74401 | -0.061705 | 0.00016263
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 603 | 2.4805 | [20 empty] | -0.091942 | 1.65 | 1.5613 | -0.96415 | 0.44424 | -0.15171 | 0.038381 | -0.0061047 | 0.00043056 | [12 empty] | 1.2716 | 1.0397 | 0.74648 | -0.062013 | 0.00016371
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 604 | 2.4856 | [20 empty] | -0.091942 | 1.6528 | 1.5665 | -0.96896 | 0.4472 | -0.15298 | 0.038764 | -0.0061759 | 0.00043631 | [12 empty] | 1.2716 | 1.0414 | 0.74896 | -0.062322 | 0.0001648
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 605 | 2.4907 | [20 empty] | -0.091942 | 1.6555 | 1.5716 | -0.97378 | 0.45017 | -0.15425 | 0.039151 | -0.0062479 | 0.00044212 | [12 empty] | 1.2716 | 1.0432 | 0.75144 | -0.062632 | 0.0001659
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 606 | 2.4958 | [20 empty] | -0.091942 | 1.6583 | 1.5768 | -0.97861 | 0.45315 | -0.15552 | 0.039541 | -0.0063205 | 0.000448 | [12 empty] | 1.2716 | 1.0449 | 0.75393 | -0.062943 | 0.000167
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 607 | 2.501 | [20 empty] | -0.091942 | 1.661 | 1.5821 | -0.98347 | 0.45615 | -0.15681 | 0.039934 | -0.0063939 | 0.00045395 | [12 empty] | 1.2716 | 1.0466 | 0.75642 | -0.063255 | 0.0001681
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 608 | 2.5061 | [20 empty] | -0.091942 | 1.6637 | 1.5873 | -0.98833 | 0.45916 | -0.15811 | 0.040331 | -0.006468 | 0.00045997 | [12 empty] | 1.2716 | 1.0483 | 0.75891 | -0.063569 | 0.00016921
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 609 | 2.5112 | [20 empty] | -0.091942 | 1.6665 | 1.5925 | -0.99322 | 0.46219 | -0.15941 | 0.04073 | -0.0065428 | 0.00046606 | [12 empty] | 1.2716 | 1.0501 | 0.76141 | -0.063883 | 0.00017033
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 610 | 2.5164 | [20 empty] | -0.091942 | 1.6692 | 1.5977 | -0.99812 | 0.46523 | -0.16073 | 0.041133 | -0.0066184 | 0.00047221 | [12 empty] | 1.2716 | 1.0518 | 0.76392 | -0.064198 | 0.00017145
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 611 | 2.5215 | [20 empty] | -0.091942 | 1.6719 | 1.603 | -1.003 | 0.46829 | -0.16205 | 0.041539 | -0.0066947 | 0.00047844 | [12 empty] | 1.2716 | 1.0535 | 0.76642 | -0.064514 | 0.00017258
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 612 | 2.5266 | [20 empty] | -0.091942 | 1.6747 | 1.6082 | -1.008 | 0.47137 | -0.16338 | 0.041949 | -0.0067718 | 0.00048474 | [12 empty] | 1.2716 | 1.0552 | 0.76893 | -0.064831 | 0.00017371
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 613 | 2.5318 | [20 empty] | -0.091942 | 1.6774 | 1.6135 | -1.0129 | 0.47445 | -0.16472 | 0.042362 | -0.0068496 | 0.00049111 | [12 empty] | 1.2716 | 1.057 | 0.77145 | -0.06515 | 0.00017485
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 614 | 2.5369 | [20 empty] | -0.091942 | 1.6801 | 1.6187 | -1.0179 | 0.47756 | -0.16607 | 0.042778 | -0.0069282 | 0.00049756 | [12 empty] | 1.2716 | 1.0587 | 0.77397 | -0.065469 | 0.00017599
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 615 | 2.542 | [20 empty] | -0.091942 | 1.6829 | 1.624 | -1.0229 | 0.48068 | -0.16742 | 0.043198 | -0.0070076 | 0.00050408 | [12 empty] | 1.2716 | 1.0604 | 0.77649 | -0.06579 | 0.00017714
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 616 | 2.5472 | [20 empty] | -0.091942 | 1.6856 | 1.6293 | -1.0279 | 0.48381 | -0.16879 | 0.043621 | -0.0070878 | 0.00051068 | [12 empty] | 1.2716 | 1.0621 | 0.77902 | -0.066111 | 0.0001783
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 617 | 2.5523 | [20 empty] | -0.091942 | 1.6884 | 1.6346 | -1.0329 | 0.48696 | -0.17016 | 0.044048 | -0.0071687 | 0.00051735 | [12 empty] | 1.2716 | 1.0639 | 0.78155 | -0.066433 | 0.00017946
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 618 | 2.5575 | [20 empty] | -0.091942 | 1.6911 | 1.6399 | -1.0379 | 0.49012 | -0.17155 | 0.044478 | -0.0072504 | 0.00052409 | [12 empty] | 1.2716 | 1.0656 | 0.78408 | -0.066757 | 0.00018062
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 619 | 2.5626 | [20 empty] | -0.091942 | 1.6938 | 1.6452 | -1.043 | 0.4933 | -0.17294 | 0.044911 | -0.007333 | 0.00053091 | [12 empty] | 1.2716 | 1.0673 | 0.78662 | -0.067082 | 0.00018179
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 620 | 2.5678 | [20 empty] | -0.091942 | 1.6966 | 1.6505 | -1.048 | 0.4965 | -0.17434 | 0.045348 | -0.0074163 | 0.00053781 | [12 empty] | 1.2716 | 1.069 | 0.78917 | -0.067407 | 0.00018297
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 621 | 2.5729 | [20 empty] | -0.091942 | 1.6993 | 1.6559 | -1.0531 | 0.49971 | -0.17575 | 0.045789 | -0.0075004 | 0.00054479 | [12 empty] | 1.2716 | 1.0707 | 0.79171 | -0.067734 | 0.00018416
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 622 | 2.5781 | [20 empty] | -0.091942 | 1.702 | 1.6612 | -1.0582 | 0.50294 | -0.17717 | 0.046233 | -0.0075854 | 0.00055185 | [12 empty] | 1.2716 | 1.0725 | 0.79427 | -0.068062 | 0.00018534
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 623 | 2.5832 | [20 empty] | -0.091942 | 1.7048 | 1.6666 | -1.0633 | 0.50618 | -0.1786 | 0.046681 | -0.0076712 | 0.00055899 | [12 empty] | 1.2716 | 1.0742 | 0.79682 | -0.06839 | 0.00018654
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 624 | 2.5884 | [20 empty] | -0.091942 | 1.7075 | 1.6719 | -1.0684 | 0.50944 | -0.18004 | 0.047132 | -0.0077578 | 0.00056621 | [12 empty] | 1.2716 | 1.0759 | 0.79938 | -0.06872 | 0.00018774
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 625 | 2.5935 | [20 empty] | -0.091942 | 1.7102 | 1.6773 | -1.0736 | 0.51271 | -0.18148 | 0.047587 | -0.0078452 | 0.00057351 | [12 empty] | 1.2716 | 1.0776 | 0.80195 | -0.069051 | 0.00018895
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 626 | 2.5987 | [20 empty] | -0.091942 | 1.713 | 1.6826 | -1.0787 | 0.516 | -0.18294 | 0.048046 | -0.0079335 | 0.00058089 | [12 empty] | 1.2716 | 1.0794 | 0.80452 | -0.069383 | 0.00019016
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 627 | 2.6039 | [20 empty] | -0.091942 | 1.7157 | 1.688 | -1.0839 | 0.5193 | -0.18441 | 0.048508 | -0.0080226 | 0.00058835 | [12 empty] | 1.2716 | 1.0811 | 0.80709 | -0.069716 | 0.00019138
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 628 | 2.609 | [20 empty] | -0.091942 | 1.7185 | 1.6934 | -1.0891 | 0.52263 | -0.18588 | 0.048975 | -0.0081126 | 0.0005959 | [12 empty] | 1.2716 | 1.0828 | 0.80966 | -0.07005 | 0.0001926
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 629 | 2.6142 | [20 empty] | -0.091942 | 1.7212 | 1.6988 | -1.0943 | 0.52596 | -0.18737 | 0.049444 | -0.0082035 | 0.00060354 | [12 empty] | 1.2716 | 1.0845 | 0.81224 | -0.070386 | 0.00019383
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 630 | 2.6193 | [20 empty] | -0.091942 | 1.7239 | 1.7042 | -1.0995 | 0.52931 | -0.18886 | 0.049918 | -0.0082952 | 0.00061126 | [12 empty] | 1.2716 | 1.0863 | 0.81483 | -0.070722 | 0.00019507
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 631 | 2.6245 | [20 empty] | -0.091942 | 1.7267 | 1.7096 | -1.1048 | 0.53268 | -0.19036 | 0.050395 | -0.0083878 | 0.00061906 | [12 empty] | 1.2716 | 1.088 | 0.81742 | -0.071059 | 0.00019631
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 632 | 2.6297 | [20 empty] | -0.091942 | 1.7294 | 1.7151 | -1.1101 | 0.53607 | -0.19188 | 0.050876 | -0.0084813 | 0.00062695 | [12 empty] | 1.2716 | 1.0897 | 0.82001 | -0.071397 | 0.00019755
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 633 | 2.6348 | [20 empty] | -0.091942 | 1.7321 | 1.7205 | -1.1153 | 0.53947 | -0.1934 | 0.051361 | -0.0085757 | 0.00063493 | [12 empty] | 1.2716 | 1.0914 | 0.82261 | -0.071737 | 0.00019881
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 634 | 2.64 | [20 empty] | -0.091942 | 1.7349 | 1.7259 | -1.1206 | 0.54289 | -0.19493 | 0.05185 | -0.008671 | 0.000643 | [12 empty] | 1.2716 | 1.0932 | 0.82521 | -0.072077 | 0.00020007
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 635 | 2.6452 | [20 empty] | -0.091942 | 1.7376 | 1.7314 | -1.1259 | 0.54632 | -0.19647 | 0.052343 | -0.0087672 | 0.00065116 | [12 empty] | 1.2716 | 1.0949 | 0.82781 | -0.072419 | 0.00020133
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 636 | 2.6503 | [20 empty] | -0.091942 | 1.7403 | 1.7368 | -1.1313 | 0.54977 | -0.19803 | 0.052839 | -0.0088643 | 0.00065941 | [12 empty] | 1.2716 | 1.0966 | 0.83042 | -0.072762 | 0.0002026
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 637 | 2.6555 | [20 empty] | -0.091942 | 1.7431 | 1.7423 | -1.1366 | 0.55324 | -0.19959 | 0.05334 | -0.0089623 | 0.00066775 | [12 empty] | 1.2716 | 1.0983 | 0.83304 | -0.073105 | 0.00020388
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 638 | 2.6607 | [20 empty] | -0.091942 | 1.7458 | 1.7478 | -1.142 | 0.55672 | -0.20116 | 0.053844 | -0.0090613 | 0.00067618 | [12 empty] | 1.2716 | 1.1001 | 0.83565 | -0.07345 | 0.00020516
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 639 | 2.6659 | [20 empty] | -0.091942 | 1.7486 | 1.7533 | -1.1473 | 0.56022 | -0.20274 | 0.054352 | -0.0091612 | 0.00068471 | [12 empty] | 1.2716 | 1.1018 | 0.83828 | -0.073796 | 0.00020645
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 640 | 2.671 | [20 empty] | -0.091942 | 1.7513 | 1.7587 | -1.1527 | 0.56373 | -0.20433 | 0.054865 | -0.009262 | 0.00069333 | [12 empty] | 1.2716 | 1.1035 | 0.8409 | -0.074143 | 0.00020775
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 641 | 2.6762 | [20 empty] | -0.091942 | 1.754 | 1.7642 | -1.1582 | 0.56726 | -0.20593 | 0.055381 | -0.0093638 | 0.00070204 | [12 empty] | 1.2716 | 1.1052 | 0.84353 | -0.074491 | 0.00020905
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 642 | 2.6814 | [20 empty] | -0.091942 | 1.7568 | 1.7698 | -1.1636 | 0.57081 | -0.20754 | 0.055901 | -0.0094665 | 0.00071085 | [12 empty] | 1.2716 | 1.107 | 0.84617 | -0.07484 | 0.00021036
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 643 | 2.6866 | [20 empty] | -0.091942 | 1.7595 | 1.7753 | -1.169 | 0.57438 | -0.20917 | 0.056426 | -0.0095702 | 0.00071976 | [12 empty] | 1.2716 | 1.1087 | 0.8488 | -0.075191 | 0.00021167
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 644 | 2.6918 | [20 empty] | -0.091942 | 1.7622 | 1.7808 | -1.1745 | 0.57796 | -0.2108 | 0.056954 | -0.0096749 | 0.00072876 | [12 empty] | 1.2716 | 1.1104 | 0.85145 | -0.075542 | 0.00021299
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 645 | 2.6969 | [20 empty] | -0.091942 | 1.765 | 1.7863 | -1.18 | 0.58156 | -0.21244 | 0.057487 | -0.0097805 | 0.00073786 | [12 empty] | 1.2716 | 1.1121 | 0.85409 | -0.075895 | 0.00021432
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 646 | 2.7021 | [20 empty] | -0.091942 | 1.7677 | 1.7919 | -1.1855 | 0.58517 | -0.21409 | 0.058024 | -0.0098872 | 0.00074707 | [12 empty] | 1.2716 | 1.1139 | 0.85674 | -0.076248 | 0.00021565
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 647 | 2.7073 | [20 empty] | -0.091942 | 1.7704 | 1.7974 | -1.191 | 0.5888 | -0.21575 | 0.058565 | -0.0099948 | 0.00075637 | [12 empty] | 1.2716 | 1.1156 | 0.8594 | -0.076603 | 0.00021699
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 648 | 2.7125 | [20 empty] | -0.091942 | 1.7732 | 1.803 | -1.1965 | 0.59245 | -0.21743 | 0.05911 | -0.010103 | 0.00076577 | [12 empty] | 1.2716 | 1.1173 | 0.86206 | -0.076958 | 0.00021833
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 649 | 2.7177 | [20 empty] | -0.091942 | 1.7759 | 1.8086 | -1.2021 | 0.59612 | -0.21911 | 0.05966 | -0.010213 | 0.00077528 | [12 empty] | 1.2716 | 1.119 | 0.86472 | -0.077315 | 0.00021968
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 650 | 2.7229 | [20 empty] | -0.091942 | 1.7787 | 1.8141 | -1.2076 | 0.5998 | -0.2208 | 0.060213 | -0.010324 | 0.00078488 | [12 empty] | 1.2716 | 1.1208 | 0.86739 | -0.077673 | 0.00022104
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 651 | 2.7281 | [20 empty] | -0.091942 | 1.7814 | 1.8197 | -1.2132 | 0.6035 | -0.22251 | 0.060771 | -0.010435 | 0.0007946 | [12 empty] | 1.2716 | 1.1225 | 0.87006 | -0.078032 | 0.0002224
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 652 | 2.7333 | [20 empty] | -0.091942 | 1.7841 | 1.8253 | -1.2188 | 0.60721 | -0.22422 | 0.061334 | -0.010548 | 0.00080441 | [12 empty] | 1.2716 | 1.1242 | 0.87273 | -0.078392 | 0.00022377
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 653 | 2.7385 | [20 empty] | -0.091942 | 1.7869 | 1.8309 | -1.2244 | 0.61095 | -0.22595 | 0.0619 | -0.010662 | 0.00081434 | [12 empty] | 1.2716 | 1.1259 | 0.87541 | -0.078754 | 0.00022515
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 654 | 2.7437 | [20 empty] | -0.091942 | 1.7896 | 1.8365 | -1.2301 | 0.6147 | -0.22768 | 0.062471 | -0.010777 | 0.00082437 | [12 empty] | 1.2716 | 1.1276 | 0.87809 | -0.079116 | 0.00022653
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 655 | 2.7489 | [20 empty] | -0.091942 | 1.7923 | 1.8422 | -1.2357 | 0.61847 | -0.22943 | 0.063046 | -0.010893 | 0.00083451 | [12 empty] | 1.2716 | 1.1294 | 0.88078 | -0.07948 | 0.00022792
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 656 | 2.7541 | [20 empty] | -0.091942 | 1.7951 | 1.8478 | -1.2414 | 0.62225 | -0.23118 | 0.063626 | -0.01101 | 0.00084475 | [12 empty] | 1.2716 | 1.1311 | 0.88347 | -0.079844 | 0.00022932
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 657 | 2.7593 | [20 empty] | -0.091942 | 1.7978 | 1.8534 | -1.2471 | 0.62606 | -0.23295 | 0.06421 | -0.011128 | 0.00085511 | [12 empty] | 1.2716 | 1.1328 | 0.88617 | -0.08021 | 0.00023072
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 658 | 2.7645 | [20 empty] | -0.091942 | 1.8005 | 1.8591 | -1.2528 | 0.62988 | -0.23473 | 0.064799 | -0.011247 | 0.00086558 | [12 empty] | 1.2716 | 1.1345 | 0.88887 | -0.080577 | 0.00023213
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 659 | 2.7697 | [20 empty] | -0.091942 | 1.8033 | 1.8647 | -1.2585 | 0.63371 | -0.23652 | 0.065392 | -0.011367 | 0.00087616 | [12 empty] | 1.2716 | 1.1363 | 0.89157 | -0.080945 | 0.00023354
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 660 | 2.7749 | [20 empty] | -0.091942 | 1.806 | 1.8704 | -1.2642 | 0.63757 | -0.23832 | 0.06599 | -0.011488 | 0.00088685 | [12 empty] | 1.2716 | 1.138 | 0.89428 | -0.081314 | 0.00023496
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 661 | 2.7801 | [20 empty] | -0.091942 | 1.8088 | 1.8761 | -1.27 | 0.64144 | -0.24013 | 0.066592 | -0.011611 | 0.00089766 | [12 empty] | 1.2716 | 1.1397 | 0.89699 | -0.081684 | 0.00023639
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 662 | 2.7853 | [20 empty] | -0.091942 | 1.8115 | 1.8817 | -1.2758 | 0.64533 | -0.24195 | 0.067199 | -0.011734 | 0.00090858 | [12 empty] | 1.2716 | 1.1414 | 0.89971 | -0.082055 | 0.00023782
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 663 | 2.7905 | [20 empty] | -0.091942 | 1.8142 | 1.8874 | -1.2815 | 0.64924 | -0.24378 | 0.06781 | -0.011859 | 0.00091962 | [12 empty] | 1.2716 | 1.1432 | 0.90243 | -0.082427 | 0.00023926
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 664 | 2.7957 | [20 empty] | -0.091942 | 1.817 | 1.8931 | -1.2874 | 0.65317 | -0.24563 | 0.068426 | -0.011985 | 0.00093077 | [12 empty] | 1.2716 | 1.1449 | 0.90515 | -0.082801 | 0.00024071
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 665 | 2.8009 | [20 empty] | -0.091942 | 1.8197 | 1.8988 | -1.2932 | 0.65711 | -0.24748 | 0.069047 | -0.012111 | 0.00094205 | [12 empty] | 1.2716 | 1.1466 | 0.90788 | -0.083176 | 0.00024216
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 666 | 2.8061 | [20 empty] | -0.091942 | 1.8224 | 1.9045 | -1.299 | 0.66107 | -0.24935 | 0.069672 | -0.01224 | 0.00095344 | [12 empty] | 1.2716 | 1.1483 | 0.91061 | -0.083551 | 0.00024362
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 667 | 2.8113 | [20 empty] | -0.091942 | 1.8252 | 1.9103 | -1.3049 | 0.66505 | -0.25123 | 0.070302 | -0.012369 | 0.00096495 | [12 empty] | 1.2716 | 1.1501 | 0.91335 | -0.083928 | 0.00024509
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 668 | 2.8166 | [20 empty] | -0.091942 | 1.8279 | 1.916 | -1.3108 | 0.66905 | -0.25312 | 0.070937 | -0.012499 | 0.00097659 | [12 empty] | 1.2716 | 1.1518 | 0.91609 | -0.084306 | 0.00024656
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 669 | 2.8218 | [20 empty] | -0.091942 | 1.8306 | 1.9217 | -1.3167 | 0.67306 | -0.25502 | 0.071576 | -0.012631 | 0.00098834 | [12 empty] | 1.2716 | 1.1535 | 0.91884 | -0.084686 | 0.00024804
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 670 | 2.827 | [20 empty] | -0.091942 | 1.8334 | 1.9275 | -1.3226 | 0.6771 | -0.25693 | 0.072221 | -0.012763 | 0.0010002 | [12 empty] | 1.2716 | 1.1552 | 0.92158 | -0.085066 | 0.00024953
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 671 | 2.8322 | [20 empty] | -0.091942 | 1.8361 | 1.9332 | -1.3285 | 0.68115 | -0.25885 | 0.07287 | -0.012897 | 0.0010122 | [12 empty] | 1.2716 | 1.157 | 0.92434 | -0.085447 | 0.00025102
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 672 | 2.8374 | [20 empty] | -0.091942 | 1.8389 | 1.939 | -1.3344 | 0.68522 | -0.26079 | 0.073524 | -0.013033 | 0.0010244 | [12 empty] | 1.2716 | 1.1587 | 0.92709 | -0.08583 | 0.00025252
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 673 | 2.8427 | [20 empty] | -0.091942 | 1.8416 | 1.9448 | -1.3404 | 0.68931 | -0.26273 | 0.074183 | -0.013169 | 0.0010366 | [12 empty] | 1.2716 | 1.1604 | 0.92986 | -0.086214 | 0.00025403
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 674 | 2.8479 | [20 empty] | -0.091942 | 1.8443 | 1.9506 | -1.3464 | 0.69341 | -0.26469 | 0.074847 | -0.013306 | 0.001049 | [12 empty] | 1.2716 | 1.1621 | 0.93262 | -0.086599 | 0.00025554
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 675 | 2.8531 | [20 empty] | -0.091942 | 1.8471 | 1.9564 | -1.3524 | 0.69754 | -0.26666 | 0.075515 | -0.013445 | 0.0010615 | [12 empty] | 1.2716 | 1.1639 | 0.93539 | -0.086985 | 0.00025706
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 676 | 2.8583 | [20 empty] | -0.091942 | 1.8498 | 1.9622 | -1.3584 | 0.70168 | -0.26864 | 0.076189 | -0.013585 | 0.0010742 | [12 empty] | 1.2716 | 1.1656 | 0.93816 | -0.087372 | 0.00025859
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 677 | 2.8636 | [20 empty] | -0.091942 | 1.8525 | 1.968 | -1.3645 | 0.70584 | -0.27063 | 0.076868 | -0.013727 | 0.0010869 | [12 empty] | 1.2716 | 1.1673 | 0.94094 | -0.08776 | 0.00026012
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 678 | 2.8688 | [20 empty] | -0.091942 | 1.8553 | 1.9738 | -1.3705 | 0.71002 | -0.27264 | 0.077552 | -0.013869 | 0.0010999 | [12 empty] | 1.2716 | 1.169 | 0.94372 | -0.08815 | 0.00026166
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 679 | 2.874 | [20 empty] | -0.091942 | 1.858 | 1.9796 | -1.3766 | 0.71422 | -0.27465 | 0.07824 | -0.014013 | 0.0011129 | [12 empty] | 1.2716 | 1.1708 | 0.94651 | -0.08854 | 0.00026321
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 680 | 2.8792 | [20 empty] | -0.091942 | 1.8607 | 1.9855 | -1.3827 | 0.71844 | -0.27668 | 0.078934 | -0.014158 | 0.0011261 | [12 empty] | 1.2716 | 1.1725 | 0.9493 | -0.088932 | 0.00026476
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 681 | 2.8845 | [20 empty] | -0.091942 | 1.8635 | 1.9913 | -1.3888 | 0.72267 | -0.27872 | 0.079633 | -0.014305 | 0.0011394 | [12 empty] | 1.2716 | 1.1742 | 0.95209 | -0.089325 | 0.00026632
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 682 | 2.8897 | [20 empty] | -0.091942 | 1.8662 | 1.9972 | -1.3949 | 0.72692 | -0.28078 | 0.080338 | -0.014452 | 0.0011529 | [12 empty] | 1.2716 | 1.1759 | 0.95489 | -0.089719 | 0.00026789
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 683 | 2.8949 | [20 empty] | -0.091942 | 1.869 | 2.003 | -1.4011 | 0.7312 | -0.28284 | 0.081047 | -0.014601 | 0.0011664 | [12 empty] | 1.2716 | 1.1777 | 0.95769 | -0.090114 | 0.00026946
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 684 | 2.9002 | [20 empty] | -0.091942 | 1.8717 | 2.0089 | -1.4072 | 0.73549 | -0.28492 | 0.081762 | -0.014752 | 0.0011802 | [12 empty] | 1.2716 | 1.1794 | 0.9605 | -0.090511 | 0.00027105
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 685 | 2.9054 | [20 empty] | -0.091942 | 1.8744 | 2.0148 | -1.4134 | 0.7398 | -0.28701 | 0.082481 | -0.014903 | 0.0011941 | [12 empty] | 1.2716 | 1.1811 | 0.96331 | -0.090908 | 0.00027263
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 686 | 2.9107 | [20 empty] | -0.091942 | 1.8772 | 2.0207 | -1.4196 | 0.74413 | -0.28911 | 0.083206 | -0.015056 | 0.0012081 | [12 empty] | 1.2716 | 1.1828 | 0.96613 | -0.091307 | 0.00027423
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 687 | 2.9159 | [20 empty] | -0.091942 | 1.8799 | 2.0265 | -1.4258 | 0.74848 | -0.29122 | 0.083937 | -0.01521 | 0.0012222 | [12 empty] | 1.2716 | 1.1845 | 0.96894 | -0.091707 | 0.00027583
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 688 | 2.9211 | [20 empty] | -0.091942 | 1.8826 | 2.0324 | -1.4321 | 0.75285 | -0.29335 | 0.084673 | -0.015366 | 0.0012365 | [12 empty] | 1.2716 | 1.1863 | 0.97177 | -0.092108 | 0.00027744
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 689 | 2.9264 | [20 empty] | -0.091942 | 1.8854 | 2.0384 | -1.4383 | 0.75723 | -0.29548 | 0.085414 | -0.015523 | 0.001251 | [12 empty] | 1.2716 | 1.188 | 0.97459 | -0.09251 | 0.00027906
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 690 | 2.9316 | [20 empty] | -0.091942 | 1.8881 | 2.0443 | -1.4446 | 0.76164 | -0.29763 | 0.08616 | -0.015682 | 0.0012656 | [12 empty] | 1.2716 | 1.1897 | 0.97743 | -0.092913 | 0.00028068
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 691 | 2.9369 | [20 empty] | -0.091942 | 1.8908 | 2.0502 | -1.4509 | 0.76606 | -0.2998 | 0.086912 | -0.015841 | 0.0012803 | [12 empty] | 1.2716 | 1.1914 | 0.98026 | -0.093318 | 0.00028231
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 692 | 2.9421 | [20 empty] | -0.091942 | 1.8936 | 2.0562 | -1.4572 | 0.77051 | -0.30197 | 0.08767 | -0.016002 | 0.0012952 | [12 empty] | 1.2716 | 1.1932 | 0.9831 | -0.093724 | 0.00028395
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 693 | 2.9474 | [20 empty] | -0.091942 | 1.8963 | 2.0621 | -1.4635 | 0.77497 | -0.30416 | 0.088433 | -0.016165 | 0.0013103 | [12 empty] | 1.2716 | 1.1949 | 0.98594 | -0.094131 | 0.0002856
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 694 | 2.9526 | [20 empty] | -0.091942 | 1.8991 | 2.0681 | -1.4698 | 0.77945 | -0.30636 | 0.089201 | -0.016329 | 0.0013255 | [12 empty] | 1.2716 | 1.1966 | 0.98879 | -0.094539 | 0.00028725
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 695 | 2.9579 | [20 empty] | -0.091942 | 1.9018 | 2.074 | -1.4762 | 0.78396 | -0.30858 | 0.089975 | -0.016494 | 0.0013408 | [12 empty] | 1.2716 | 1.1983 | 0.99164 | -0.094948 | 0.00028891
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 696 | 2.9631 | [20 empty] | -0.091942 | 1.9045 | 2.08 | -1.4826 | 0.78848 | -0.3108 | 0.090754 | -0.016661 | 0.0013564 | [12 empty] | 1.2716 | 1.2001 | 0.9945 | -0.095358 | 0.00029057
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 697 | 2.9684 | [20 empty] | -0.091942 | 1.9073 | 2.086 | -1.489 | 0.79302 | -0.31304 | 0.09154 | -0.01683 | 0.001372 | [12 empty] | 1.2716 | 1.2018 | 0.99736 | -0.09577 | 0.00029225
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 698 | 2.9736 | [20 empty] | -0.091942 | 1.91 | 2.092 | -1.4954 | 0.79758 | -0.31529 | 0.09233 | -0.016999 | 0.0013878 | [12 empty] | 1.2716 | 1.2035 | 1.0002 | -0.096183 | 0.00029393
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 699 | 2.9789 | [20 empty] | -0.091942 | 1.9127 | 2.098 | -1.5018 | 0.80216 | -0.31756 | 0.093127 | -0.017171 | 0.0014038 | [12 empty] | 1.2716 | 1.2052 | 1.0031 | -0.096597 | 0.00029562
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 700 | 2.9841 | [20 empty] | -0.091942 | 1.9155 | 2.104 | -1.5083 | 0.80676 | -0.31984 | 0.093929 | -0.017343 | 0.00142 | [12 empty] | 1.2716 | 1.207 | 1.006 | -0.097012 | 0.00029731
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 701 | 2.9894 | [20 empty] | -0.091942 | 1.9182 | 2.11 | -1.5148 | 0.81138 | -0.32213 | 0.094737 | -0.017517 | 0.0014363 | [12 empty] | 1.2716 | 1.2087 | 1.0088 | -0.097428 | 0.00029901
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 702 | 2.9946 | [20 empty] | -0.091942 | 1.9209 | 2.116 | -1.5213 | 0.81602 | -0.32443 | 0.095551 | -0.017693 | 0.0014528 | [12 empty] | 1.2716 | 1.2104 | 1.0117 | -0.097846 | 0.00030072
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 703 | 2.9999 | [20 empty] | -0.091942 | 1.9237 | 2.122 | -1.5278 | 0.82068 | -0.32675 | 0.096371 | -0.01787 | 0.0014694 | [12 empty] | 1.2716 | 1.2121 | 1.0146 | -0.098265 | 0.00030244
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 704 | 3.0052 | [20 empty] | -0.091942 | 1.9264 | 2.1281 | -1.5343 | 0.82536 | -0.32908 | 0.097196 | -0.018049 | 0.0014862 | [12 empty] | 1.2716 | 1.2139 | 1.0175 | -0.098685 | 0.00030416
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 705 | 3.0104 | [20 empty] | -0.091942 | 1.9292 | 2.1341 | -1.5408 | 0.83006 | -0.33142 | 0.098027 | -0.018229 | 0.0015032 | [12 empty] | 1.2716 | 1.2156 | 1.0204 | -0.099106 | 0.0003059
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 706 | 3.0157 | [20 empty] | -0.091942 | 1.9319 | 2.1402 | -1.5474 | 0.83478 | -0.33378 | 0.098865 | -0.018411 | 0.0015203 | [12 empty] | 1.2716 | 1.2173 | 1.0233 | -0.099528 | 0.00030764
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 707 | 3.021 | [20 empty] | -0.091942 | 1.9346 | 2.1463 | -1.554 | 0.83952 | -0.33615 | 0.099708 | -0.018594 | 0.0015376 | [12 empty] | 1.2716 | 1.219 | 1.0262 | -0.099952 | 0.00030938
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 708 | 3.0262 | [20 empty] | -0.091942 | 1.9374 | 2.1523 | -1.5606 | 0.84428 | -0.33853 | 0.10056 | -0.018779 | 0.0015551 | [12 empty] | 1.2716 | 1.2208 | 1.0291 | -0.10038 | 0.00031114
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 709 | 3.0315 | [20 empty] | -0.091942 | 1.9401 | 2.1584 | -1.5672 | 0.84906 | -0.34093 | 0.10141 | -0.018966 | 0.0015728 | [12 empty] | 1.2716 | 1.2225 | 1.032 | -0.1008 | 0.0003129
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 710 | 3.0368 | [20 empty] | -0.091942 | 1.9428 | 2.1645 | -1.5739 | 0.85386 | -0.34334 | 0.10227 | -0.019154 | 0.0015906 | [12 empty] | 1.2716 | 1.2242 | 1.0349 | -0.10123 | 0.00031467
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 711 | 3.042 | [20 empty] | -0.091942 | 1.9456 | 2.1706 | -1.5805 | 0.85868 | -0.34577 | 0.10314 | -0.019343 | 0.0016086 | [12 empty] | 1.2716 | 1.2259 | 1.0378 | -0.10166 | 0.00031644
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 712 | 3.0473 | [20 empty] | -0.091942 | 1.9483 | 2.1767 | -1.5872 | 0.86352 | -0.34821 | 0.10401 | -0.019535 | 0.0016268 | [12 empty] | 1.2716 | 1.2277 | 1.0407 | -0.10209 | 0.00031823
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 713 | 3.0526 | [20 empty] | -0.091942 | 1.951 | 2.1828 | -1.5939 | 0.86838 | -0.35066 | 0.10489 | -0.019727 | 0.0016452 | [12 empty] | 1.2716 | 1.2294 | 1.0437 | -0.10252 | 0.00032002
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 714 | 3.0578 | [20 empty] | -0.091942 | 1.9538 | 2.189 | -1.6006 | 0.87326 | -0.35312 | 0.10578 | -0.019922 | 0.0016637 | [12 empty] | 1.2716 | 1.2311 | 1.0466 | -0.10295 | 0.00032182
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 715 | 3.0631 | [20 empty] | -0.091942 | 1.9565 | 2.1951 | -1.6074 | 0.87816 | -0.3556 | 0.10667 | -0.020118 | 0.0016825 | [12 empty] | 1.2716 | 1.2328 | 1.0495 | -0.10338 | 0.00032363
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 716 | 3.0684 | [20 empty] | -0.091942 | 1.9593 | 2.2012 | -1.6141 | 0.88309 | -0.3581 | 0.10757 | -0.020316 | 0.0017014 | [12 empty] | 1.2716 | 1.2346 | 1.0525 | -0.10382 | 0.00032544
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 717 | 3.0737 | [20 empty] | -0.091942 | 1.962 | 2.2074 | -1.6209 | 0.88803 | -0.36061 | 0.10847 | -0.020515 | 0.0017205 | [12 empty] | 1.2716 | 1.2363 | 1.0554 | -0.10425 | 0.00032726
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 718 | 3.079 | [20 empty] | -0.091942 | 1.9647 | 2.2136 | -1.6277 | 0.893 | -0.36313 | 0.10939 | -0.020716 | 0.0017398 | [12 empty] | 1.2716 | 1.238 | 1.0584 | -0.10469 | 0.00032909
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 719 | 3.0842 | [20 empty] | -0.091942 | 1.9675 | 2.2197 | -1.6345 | 0.89798 | -0.36566 | 0.1103 | -0.020919 | 0.0017593 | [12 empty] | 1.2716 | 1.2397 | 1.0613 | -0.10513 | 0.00033093
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 720 | 3.0895 | [20 empty] | -0.091942 | 1.9702 | 2.2259 | -1.6413 | 0.90299 | -0.36821 | 0.11123 | -0.021124 | 0.0017789 | [12 empty] | 1.2716 | 1.2414 | 1.0643 | -0.10557 | 0.00033277
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 721 | 3.0948 | [20 empty] | -0.091942 | 1.9729 | 2.2321 | -1.6482 | 0.90801 | -0.37078 | 0.11216 | -0.02133 | 0.0017988 | [12 empty] | 1.2716 | 1.2432 | 1.0672 | -0.10601 | 0.00033463
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 722 | 3.1001 | [20 empty] | -0.091942 | 1.9757 | 2.2383 | -1.655 | 0.91306 | -0.37336 | 0.11309 | -0.021538 | 0.0018188 | [12 empty] | 1.2716 | 1.2449 | 1.0702 | -0.10645 | 0.00033649
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 723 | 3.1054 | [20 empty] | -0.091942 | 1.9784 | 2.2445 | -1.6619 | 0.91813 | -0.37595 | 0.11404 | -0.021748 | 0.0018391 | [12 empty] | 1.2716 | 1.2466 | 1.0732 | -0.10689 | 0.00033835
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 724 | 3.1107 | [20 empty] | -0.091942 | 1.9811 | 2.2507 | -1.6688 | 0.92322 | -0.37856 | 0.11499 | -0.021959 | 0.0018595 | [12 empty] | 1.2716 | 1.2483 | 1.0761 | -0.10734 | 0.00034023
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 725 | 3.1159 | [20 empty] | -0.091942 | 1.9839 | 2.2569 | -1.6757 | 0.92833 | -0.38118 | 0.11594 | -0.022172 | 0.0018802 | [12 empty] | 1.2716 | 1.2501 | 1.0791 | -0.10778 | 0.00034211
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 726 | 3.1212 | [20 empty] | -0.091942 | 1.9866 | 2.2632 | -1.6827 | 0.93347 | -0.38381 | 0.1169 | -0.022387 | 0.001901 | [12 empty] | 1.2716 | 1.2518 | 1.0821 | -0.10823 | 0.00034401
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 727 | 3.1265 | [20 empty] | -0.091942 | 1.9894 | 2.2694 | -1.6896 | 0.93862 | -0.38646 | 0.11787 | -0.022604 | 0.0019221 | [12 empty] | 1.2716 | 1.2535 | 1.0851 | -0.10868 | 0.0003459
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 728 | 3.1318 | [20 empty] | -0.091942 | 1.9921 | 2.2757 | -1.6966 | 0.94379 | -0.38913 | 0.11885 | -0.022823 | 0.0019433 | [12 empty] | 1.2716 | 1.2552 | 1.088 | -0.10913 | 0.00034781
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 729 | 3.1371 | [20 empty] | -0.091942 | 1.9948 | 2.2819 | -1.7036 | 0.94899 | -0.39181 | 0.11983 | -0.023043 | 0.0019648 | [12 empty] | 1.2716 | 1.257 | 1.091 | -0.10958 | 0.00034973
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 730 | 3.1424 | [20 empty] | -0.091942 | 1.9976 | 2.2882 | -1.7107 | 0.95421 | -0.3945 | 0.12082 | -0.023265 | 0.0019865 | [12 empty] | 1.2716 | 1.2587 | 1.094 | -0.11003 | 0.00035165
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 731 | 3.1477 | [20 empty] | -0.091942 | 2.0003 | 2.2944 | -1.7177 | 0.95945 | -0.39721 | 0.12182 | -0.023489 | 0.0020083 | [12 empty] | 1.2716 | 1.2604 | 1.097 | -0.11048 | 0.00035358
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 732 | 3.153 | [20 empty] | -0.091942 | 2.003 | 2.3007 | -1.7248 | 0.96471 | -0.39994 | 0.12282 | -0.023715 | 0.0020304 | [12 empty] | 1.2716 | 1.2621 | 1.1 | -0.11093 | 0.00035552
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 733 | 3.1583 | [20 empty] | -0.091942 | 2.0058 | 2.307 | -1.7318 | 0.96999 | -0.40268 | 0.12383 | -0.023943 | 0.0020527 | [12 empty] | 1.2716 | 1.2639 | 1.103 | -0.11139 | 0.00035747
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 734 | 3.1636 | [20 empty] | -0.091942 | 2.0085 | 2.3133 | -1.7389 | 0.9753 | -0.40543 | 0.12485 | -0.024172 | 0.0020752 | [12 empty] | 1.2716 | 1.2656 | 1.1061 | -0.11185 | 0.00035942
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 735 | 3.1689 | [20 empty] | -0.091942 | 2.0112 | 2.3196 | -1.746 | 0.98062 | -0.4082 | 0.12587 | -0.024404 | 0.002098 | [12 empty] | 1.2716 | 1.2673 | 1.1091 | -0.1123 | 0.00036138
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 736 | 3.1742 | [20 empty] | -0.091942 | 2.014 | 2.3259 | -1.7532 | 0.98597 | -0.41099 | 0.12691 | -0.024637 | 0.0021209 | [12 empty] | 1.2716 | 1.269 | 1.1121 | -0.11276 | 0.00036335
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 737 | 3.1795 | [20 empty] | -0.091942 | 2.0167 | 2.3323 | -1.7603 | 0.99134 | -0.41378 | 0.12794 | -0.024872 | 0.0021441 | [12 empty] | 1.2716 | 1.2708 | 1.1151 | -0.11322 | 0.00036533
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 738 | 3.1848 | [20 empty] | -0.091942 | 2.0195 | 2.3386 | -1.7675 | 0.99673 | -0.4166 | 0.12899 | -0.02511 | 0.0021675 | [12 empty] | 1.2716 | 1.2725 | 1.1181 | -0.11368 | 0.00036732
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 739 | 3.1901 | [20 empty] | -0.091942 | 2.0222 | 2.3449 | -1.7747 | 1.0021 | -0.41943 | 0.13004 | -0.025349 | 0.0021911 | [12 empty] | 1.2716 | 1.2742 | 1.1212 | -0.11415 | 0.00036931
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 740 | 3.1954 | [20 empty] | -0.091942 | 2.0249 | 2.3513 | -1.7819 | 1.0076 | -0.42228 | 0.1311 | -0.02559 | 0.0022149 | [12 empty] | 1.2716 | 1.2759 | 1.1242 | -0.11461 | 0.00037132
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 741 | 3.2007 | [20 empty] | -0.091942 | 2.0277 | 2.3577 | -1.7892 | 1.013 | -0.42514 | 0.13217 | -0.025833 | 0.002239 | [12 empty] | 1.2716 | 1.2777 | 1.1273 | -0.11508 | 0.00037333
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 742 | 3.206 | [20 empty] | -0.091942 | 2.0304 | 2.364 | -1.7964 | 1.0185 | -0.42801 | 0.13324 | -0.026078 | 0.0022632 | [12 empty] | 1.2716 | 1.2794 | 1.1303 | -0.11554 | 0.00037535
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 743 | 3.2113 | [20 empty] | -0.091942 | 2.0331 | 2.3704 | -1.8037 | 1.024 | -0.4309 | 0.13432 | -0.026325 | 0.0022878 | [12 empty] | 1.2716 | 1.2811 | 1.1333 | -0.11601 | 0.00037738
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 744 | 3.2166 | [20 empty] | -0.091942 | 2.0359 | 2.3768 | -1.811 | 1.0295 | -0.43381 | 0.13541 | -0.026574 | 0.0023125 | [12 empty] | 1.2716 | 1.2828 | 1.1364 | -0.11648 | 0.00037941
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 745 | 3.2219 | [20 empty] | -0.091942 | 2.0386 | 2.3832 | -1.8183 | 1.0351 | -0.43674 | 0.13651 | -0.026825 | 0.0023375 | [12 empty] | 1.2716 | 1.2846 | 1.1395 | -0.11695 | 0.00038146
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 746 | 3.2272 | [20 empty] | -0.091942 | 2.0413 | 2.3896 | -1.8256 | 1.0407 | -0.43967 | 0.13761 | -0.027078 | 0.0023627 | [12 empty] | 1.2716 | 1.2863 | 1.1425 | -0.11742 | 0.00038351
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 747 | 3.2325 | [20 empty] | -0.091942 | 2.0441 | 2.396 | -1.833 | 1.0462 | -0.44263 | 0.13872 | -0.027333 | 0.0023882 | [12 empty] | 1.2716 | 1.288 | 1.1456 | -0.11789 | 0.00038557
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 748 | 3.2379 | [20 empty] | -0.091942 | 2.0468 | 2.4024 | -1.8403 | 1.0519 | -0.4456 | 0.13984 | -0.02759 | 0.0024139 | [12 empty] | 1.2716 | 1.2897 | 1.1487 | -0.11837 | 0.00038764
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 749 | 3.2432 | [20 empty] | -0.091942 | 2.0496 | 2.4088 | -1.8477 | 1.0575 | -0.44859 | 0.14096 | -0.027849 | 0.0024398 | [12 empty] | 1.2716 | 1.2915 | 1.1517 | -0.11884 | 0.00038971
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 750 | 3.2485 | [20 empty] | -0.091942 | 2.0523 | 2.4153 | -1.8551 | 1.0632 | -0.45159 | 0.1421 | -0.028111 | 0.002466 | [12 empty] | 1.2716 | 1.2932 | 1.1548 | -0.11932 | 0.0003918
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 751 | 3.2538 | [20 empty] | -0.091942 | 2.055 | 2.4217 | -1.8626 | 1.0688 | -0.45461 | 0.14324 | -0.028374 | 0.0024924 | [12 empty] | 1.2716 | 1.2949 | 1.1579 | -0.1198 | 0.00039389
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 752 | 3.2591 | [20 empty] | -0.091942 | 2.0578 | 2.4282 | -1.87 | 1.0745 | -0.45764 | 0.14438 | -0.02864 | 0.0025191 | [12 empty] | 1.2716 | 1.2966 | 1.161 | -0.12028 | 0.000396
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 753 | 3.2645 | [20 empty] | -0.091942 | 2.0605 | 2.4346 | -1.8775 | 1.0803 | -0.46069 | 0.14554 | -0.028907 | 0.002546 | [12 empty] | 1.2716 | 1.2983 | 1.1641 | -0.12076 | 0.00039811
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 754 | 3.2698 | [20 empty] | -0.091942 | 2.0632 | 2.4411 | -1.885 | 1.086 | -0.46376 | 0.1467 | -0.029177 | 0.0025732 | [12 empty] | 1.2716 | 1.3001 | 1.1672 | -0.12124 | 0.00040022
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 755 | 3.2751 | [20 empty] | -0.091942 | 2.066 | 2.4476 | -1.8925 | 1.0918 | -0.46684 | 0.14787 | -0.029449 | 0.0026006 | [12 empty] | 1.2716 | 1.3018 | 1.1703 | -0.12172 | 0.00040235
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 756 | 3.2804 | [20 empty] | -0.091942 | 2.0687 | 2.4541 | -1.9 | 1.0976 | -0.46994 | 0.14905 | -0.029723 | 0.0026283 | [12 empty] | 1.2716 | 1.3035 | 1.1734 | -0.12221 | 0.00040449
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 757 | 3.2857 | [20 empty] | -0.091942 | 2.0714 | 2.4606 | -1.9076 | 1.1034 | -0.47306 | 0.15024 | -0.03 | 0.0026562 | [12 empty] | 1.2716 | 1.3052 | 1.1765 | -0.12269 | 0.00040663
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 758 | 3.2911 | [20 empty] | -0.091942 | 2.0742 | 2.4671 | -1.9151 | 1.1092 | -0.47619 | 0.15144 | -0.030278 | 0.0026844 | [12 empty] | 1.2716 | 1.307 | 1.1796 | -0.12318 | 0.00040879
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 759 | 3.2964 | [20 empty] | -0.091942 | 2.0769 | 2.4736 | -1.9227 | 1.1151 | -0.47934 | 0.15264 | -0.030559 | 0.0027129 | [12 empty] | 1.2716 | 1.3087 | 1.1827 | -0.12367 | 0.00041095
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 760 | 3.3017 | [20 empty] | -0.091942 | 2.0797 | 2.4801 | -1.9303 | 1.121 | -0.48251 | 0.15385 | -0.030842 | 0.0027416 | [12 empty] | 1.2716 | 1.3104 | 1.1858 | -0.12416 | 0.00041312
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 761 | 3.307 | [20 empty] | -0.091942 | 2.0824 | 2.4866 | -1.938 | 1.1269 | -0.48569 | 0.15507 | -0.031127 | 0.0027706 | [12 empty] | 1.2716 | 1.3121 | 1.1889 | -0.12465 | 0.0004153
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 762 | 3.3124 | [20 empty] | -0.091942 | 2.0851 | 2.4932 | -1.9456 | 1.1328 | -0.48889 | 0.15629 | -0.031414 | 0.0027999 | [12 empty] | 1.2716 | 1.3139 | 1.1921 | -0.12514 | 0.00041748
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 763 | 3.3177 | [20 empty] | -0.091942 | 2.0879 | 2.4997 | -1.9533 | 1.1388 | -0.49211 | 0.15753 | -0.031704 | 0.0028294 | [12 empty] | 1.2716 | 1.3156 | 1.1952 | -0.12563 | 0.00041968
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 764 | 3.323 | [20 empty] | -0.091942 | 2.0906 | 2.5063 | -1.961 | 1.1448 | -0.49534 | 0.15877 | -0.031996 | 0.0028592 | [12 empty] | 1.2716 | 1.3173 | 1.1983 | -0.12613 | 0.00042188
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 765 | 3.3284 | [20 empty] | -0.091942 | 2.0933 | 2.5128 | -1.9687 | 1.1508 | -0.49859 | 0.16002 | -0.03229 | 0.0028893 | [12 empty] | 1.2716 | 1.319 | 1.2015 | -0.12662 | 0.0004241
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 766 | 3.3337 | [20 empty] | -0.091942 | 2.0961 | 2.5194 | -1.9764 | 1.1568 | -0.50186 | 0.16128 | -0.032587 | 0.0029196 | [12 empty] | 1.2716 | 1.3208 | 1.2046 | -0.12712 | 0.00042632
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 767 | 3.339 | [20 empty] | -0.091942 | 2.0988 | 2.526 | -1.9842 | 1.1629 | -0.50514 | 0.16255 | -0.032886 | 0.0029503 | [12 empty] | 1.2716 | 1.3225 | 1.2077 | -0.12762 | 0.00042855
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 768 | 3.3444 | [20 empty] | -0.091942 | 2.1015 | 2.5326 | -1.9919 | 1.169 | -0.50844 | 0.16382 | -0.033187 | 0.0029812 | [12 empty] | 1.2716 | 1.3242 | 1.2109 | -0.12812 | 0.00043079
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 769 | 3.3497 | [20 empty] | -0.091942 | 2.1043 | 2.5392 | -1.9997 | 1.1751 | -0.51176 | 0.16511 | -0.033491 | 0.0030124 | [12 empty] | 1.2716 | 1.3259 | 1.2141 | -0.12862 | 0.00043304
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 770 | 3.3551 | [20 empty] | -0.091942 | 2.107 | 2.5458 | -2.0075 | 1.1812 | -0.5151 | 0.1664 | -0.033797 | 0.0030439 | [12 empty] | 1.2716 | 1.3277 | 1.2172 | -0.12912 | 0.00043529
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 771 | 3.3604 | [20 empty] | -0.091942 | 2.1098 | 2.5524 | -2.0154 | 1.1873 | -0.51845 | 0.1677 | -0.034106 | 0.0030756 | [12 empty] | 1.2716 | 1.3294 | 1.2204 | -0.12963 | 0.00043756
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 772 | 3.3657 | [20 empty] | -0.091942 | 2.1125 | 2.559 | -2.0232 | 1.1935 | -0.52182 | 0.16901 | -0.034416 | 0.0031077 | [12 empty] | 1.2716 | 1.3311 | 1.2235 | -0.13013 | 0.00043983
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 773 | 3.3711 | [20 empty] | -0.091942 | 2.1152 | 2.5657 | -2.0311 | 1.1997 | -0.52521 | 0.17033 | -0.03473 | 0.00314 | [12 empty] | 1.2716 | 1.3328 | 1.2267 | -0.13064 | 0.00044212
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 774 | 3.3764 | [20 empty] | -0.091942 | 2.118 | 2.5723 | -2.039 | 1.2059 | -0.52862 | 0.17166 | -0.035045 | 0.0031727 | [12 empty] | 1.2716 | 1.3346 | 1.2299 | -0.13115 | 0.00044441
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 775 | 3.3818 | [20 empty] | -0.091942 | 2.1207 | 2.579 | -2.0469 | 1.2122 | -0.53204 | 0.17299 | -0.035364 | 0.0032056 | [12 empty] | 1.2716 | 1.3363 | 1.2331 | -0.13165 | 0.00044671
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 776 | 3.3871 | [20 empty] | -0.091942 | 2.1234 | 2.5856 | -2.0548 | 1.2184 | -0.53548 | 0.17433 | -0.035684 | 0.0032389 | [12 empty] | 1.2716 | 1.338 | 1.2363 | -0.13217 | 0.00044902
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 777 | 3.3925 | [20 empty] | -0.091942 | 2.1262 | 2.5923 | -2.0628 | 1.2247 | -0.53894 | 0.17569 | -0.036007 | 0.0032724 | [12 empty] | 1.2716 | 1.3397 | 1.2394 | -0.13268 | 0.00045134
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 778 | 3.3978 | [20 empty] | -0.091942 | 2.1289 | 2.599 | -2.0708 | 1.231 | -0.54242 | 0.17705 | -0.036333 | 0.0033063 | [12 empty] | 1.2716 | 1.3415 | 1.2426 | -0.13319 | 0.00045367
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 779 | 3.4032 | [20 empty] | -0.091942 | 2.1316 | 2.6057 | -2.0788 | 1.2374 | -0.54591 | 0.17842 | -0.036661 | 0.0033404 | [12 empty] | 1.2716 | 1.3432 | 1.2458 | -0.1337 | 0.000456
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 780 | 3.4085 | [20 empty] | -0.091942 | 2.1344 | 2.6124 | -2.0868 | 1.2437 | -0.54943 | 0.1798 | -0.036992 | 0.0033749 | [12 empty] | 1.2716 | 1.3449 | 1.249 | -0.13422 | 0.00045835
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 781 | 3.4139 | [20 empty] | -0.091942 | 2.1371 | 2.6191 | -2.0948 | 1.2501 | -0.55296 | 0.18118 | -0.037325 | 0.0034096 | [12 empty] | 1.2716 | 1.3466 | 1.2522 | -0.13474 | 0.0004607
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 782 | 3.4192 | [20 empty] | -0.091942 | 2.1399 | 2.6258 | -2.1029 | 1.2565 | -0.55651 | 0.18258 | -0.037661 | 0.0034447 | [12 empty] | 1.2716 | 1.3484 | 1.2554 | -0.13525 | 0.00046307
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 783 | 3.4246 | [20 empty] | -0.091942 | 2.1426 | 2.6325 | -2.111 | 1.263 | -0.56007 | 0.18399 | -0.037999 | 0.0034801 | [12 empty] | 1.2716 | 1.3501 | 1.2587 | -0.13577 | 0.00046544
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 784 | 3.4299 | [20 empty] | -0.091942 | 2.1453 | 2.6392 | -2.119 | 1.2695 | -0.56366 | 0.1854 | -0.03834 | 0.0035158 | [12 empty] | 1.2716 | 1.3518 | 1.2619 | -0.13629 | 0.00046782
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 785 | 3.4353 | [20 empty] | -0.091942 | 2.1481 | 2.646 | -2.1272 | 1.2759 | -0.56726 | 0.18682 | -0.038684 | 0.0035519 | [12 empty] | 1.2716 | 1.3535 | 1.2651 | -0.13682 | 0.00047022
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 786 | 3.4406 | [20 empty] | -0.091942 | 2.1508 | 2.6527 | -2.1353 | 1.2825 | -0.57089 | 0.18826 | -0.03903 | 0.0035882 | [12 empty] | 1.2716 | 1.3552 | 1.2683 | -0.13734 | 0.00047262
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 787 | 3.446 | [20 empty] | -0.091942 | 2.1535 | 2.6595 | -2.1435 | 1.289 | -0.57453 | 0.1897 | -0.039379 | 0.0036249 | [12 empty] | 1.2716 | 1.357 | 1.2716 | -0.13787 | 0.00047503
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 788 | 3.4513 | [20 empty] | -0.091942 | 2.1563 | 2.6662 | -2.1516 | 1.2956 | -0.57819 | 0.19115 | -0.039731 | 0.0036619 | [12 empty] | 1.2716 | 1.3587 | 1.2748 | -0.13839 | 0.00047745
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 789 | 3.4567 | [20 empty] | -0.091942 | 2.159 | 2.673 | -2.1599 | 1.3021 | -0.58186 | 0.19261 | -0.040085 | 0.0036993 | [12 empty] | 1.2716 | 1.3604 | 1.278 | -0.13892 | 0.00047987
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 790 | 3.4621 | [20 empty] | -0.091942 | 2.1617 | 2.6798 | -2.1681 | 1.3088 | -0.58556 | 0.19408 | -0.040442 | 0.0037369 | [12 empty] | 1.2716 | 1.3621 | 1.2813 | -0.13945 | 0.00048231
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 791 | 3.4674 | [20 empty] | -0.091942 | 2.1645 | 2.6866 | -2.1763 | 1.3154 | -0.58928 | 0.19556 | -0.040802 | 0.0037749 | [12 empty] | 1.2716 | 1.3639 | 1.2845 | -0.13998 | 0.00048476
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 792 | 3.4728 | [20 empty] | -0.091942 | 2.1672 | 2.6934 | -2.1846 | 1.3221 | -0.59301 | 0.19704 | -0.041164 | 0.0038133 | [12 empty] | 1.2716 | 1.3656 | 1.2878 | -0.14051 | 0.00048721
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 793 | 3.4782 | [20 empty] | -0.091942 | 2.17 | 2.7002 | -2.1929 | 1.3288 | -0.59676 | 0.19854 | -0.041529 | 0.003852 | [12 empty] | 1.2716 | 1.3673 | 1.291 | -0.14104 | 0.00048968
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 794 | 3.4835 | [20 empty] | -0.091942 | 2.1727 | 2.707 | -2.2012 | 1.3355 | -0.60054 | 0.20005 | -0.041897 | 0.003891 | [12 empty] | 1.2716 | 1.369 | 1.2943 | -0.14158 | 0.00049215
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 795 | 3.4889 | [20 empty] | -0.091942 | 2.1754 | 2.7138 | -2.2095 | 1.3422 | -0.60433 | 0.20157 | -0.042268 | 0.0039304 | [12 empty] | 1.2716 | 1.3708 | 1.2975 | -0.14211 | 0.00049464
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 796 | 3.4943 | [20 empty] | -0.091942 | 2.1782 | 2.7206 | -2.2178 | 1.349 | -0.60814 | 0.20309 | -0.042642 | 0.0039701 | [12 empty] | 1.2716 | 1.3725 | 1.3008 | -0.14265 | 0.00049713
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 797 | 3.4996 | [20 empty] | -0.091942 | 2.1809 | 2.7275 | -2.2262 | 1.3558 | -0.61197 | 0.20463 | -0.043018 | 0.0040102 | [12 empty] | 1.2716 | 1.3742 | 1.3041 | -0.14319 | 0.00049963
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 798 | 3.505 | [20 empty] | -0.091942 | 2.1836 | 2.7343 | -2.2346 | 1.3626 | -0.61582 | 0.20617 | -0.043397 | 0.0040506 | [12 empty] | 1.2716 | 1.3759 | 1.3073 | -0.14373 | 0.00050215
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 799 | 3.5104 | [20 empty] | -0.091942 | 2.1864 | 2.7412 | -2.243 | 1.3694 | -0.61968 | 0.20773 | -0.04378 | 0.0040914 | [12 empty] | 1.2716 | 1.3777 | 1.3106 | -0.14427 | 0.00050467
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 800 | 3.5157 | [20 empty] | -0.091942 | 2.1891 | 2.748 | -2.2515 | 1.3763 | -0.62357 | 0.20929 | -0.044165 | 0.0041326 | [12 empty] | 1.2716 | 1.3794 | 1.3139 | -0.14481 | 0.0005072
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 801 | 3.5211 | [20 empty] | -0.091942 | 2.1918 | 2.7549 | -2.2599 | 1.3832 | -0.62748 | 0.21087 | -0.044552 | 0.0041741 | [12 empty] | 1.2716 | 1.3811 | 1.3172 | -0.14535 | 0.00050974
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 802 | 3.5265 | [20 empty] | -0.091942 | 2.1946 | 2.7618 | -2.2684 | 1.3901 | -0.63141 | 0.21245 | -0.044943 | 0.0042159 | [12 empty] | 1.2716 | 1.3828 | 1.3205 | -0.1459 | 0.00051229
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 803 | 3.5319 | [20 empty] | -0.091942 | 2.1973 | 2.7687 | -2.2769 | 1.3971 | -0.63535 | 0.21405 | -0.045337 | 0.0042582 | [12 empty] | 1.2716 | 1.3846 | 1.3238 | -0.14645 | 0.00051485
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 804 | 3.5372 | [20 empty] | -0.091942 | 2.2001 | 2.7756 | -2.2854 | 1.404 | -0.63932 | 0.21565 | -0.045734 | 0.0043008 | [12 empty] | 1.2716 | 1.3863 | 1.3271 | -0.14699 | 0.00051742
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 805 | 3.5426 | [20 empty] | -0.091942 | 2.2028 | 2.7825 | -2.2939 | 1.411 | -0.6433 | 0.21726 | -0.046133 | 0.0043438 | [12 empty] | 1.2716 | 1.388 | 1.3304 | -0.14754 | 0.00052
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 806 | 3.548 | [20 empty] | -0.091942 | 2.2055 | 2.7894 | -2.3025 | 1.4181 | -0.64731 | 0.21889 | -0.046536 | 0.0043871 | [12 empty] | 1.2716 | 1.3897 | 1.3337 | -0.14809 | 0.00052259
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 807 | 3.5534 | [20 empty] | -0.091942 | 2.2083 | 2.7963 | -2.3111 | 1.4251 | -0.65134 | 0.22052 | -0.046942 | 0.0044309 | [12 empty] | 1.2716 | 1.3915 | 1.337 | -0.14865 | 0.00052519
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 808 | 3.5587 | [20 empty] | -0.091942 | 2.211 | 2.8033 | -2.3197 | 1.4322 | -0.65538 | 0.22217 | -0.04735 | 0.004475 | [12 empty] | 1.2716 | 1.3932 | 1.3403 | -0.1492 | 0.00052779
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 809 | 3.5641 | [20 empty] | -0.091942 | 2.2137 | 2.8102 | -2.3283 | 1.4393 | -0.65945 | 0.22382 | -0.047762 | 0.0045195 | [12 empty] | 1.2716 | 1.3949 | 1.3436 | -0.14975 | 0.00053041
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 810 | 3.5695 | [20 empty] | -0.091942 | 2.2165 | 2.8172 | -2.3369 | 1.4464 | -0.66353 | 0.22549 | -0.048177 | 0.0045644 | [12 empty] | 1.2716 | 1.3966 | 1.347 | -0.15031 | 0.00053304
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 811 | 3.5749 | [20 empty] | -0.091942 | 2.2192 | 2.8241 | -2.3456 | 1.4536 | -0.66764 | 0.22716 | -0.048595 | 0.0046096 | [12 empty] | 1.2716 | 1.3984 | 1.3503 | -0.15087 | 0.00053568
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 812 | 3.5803 | [20 empty] | -0.091942 | 2.2219 | 2.8311 | -2.3543 | 1.4608 | -0.67176 | 0.22885 | -0.049016 | 0.0046553 | [12 empty] | 1.2716 | 1.4001 | 1.3536 | -0.15143 | 0.00053832
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 813 | 3.5857 | [20 empty] | -0.091942 | 2.2247 | 2.8381 | -2.363 | 1.468 | -0.67591 | 0.23055 | -0.04944 | 0.0047014 | [12 empty] | 1.2716 | 1.4018 | 1.357 | -0.15199 | 0.00054098
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 814 | 3.5911 | [20 empty] | -0.091942 | 2.2274 | 2.8451 | -2.3717 | 1.4752 | -0.68008 | 0.23225 | -0.049867 | 0.0047478 | [12 empty] | 1.2716 | 1.4035 | 1.3603 | -0.15255 | 0.00054365
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 815 | 3.5964 | [20 empty] | -0.091942 | 2.2302 | 2.8521 | -2.3805 | 1.4825 | -0.68427 | 0.23397 | -0.050298 | 0.0047947 | [12 empty] | 1.2716 | 1.4053 | 1.3636 | -0.15311 | 0.00054632
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 816 | 3.6018 | [20 empty] | -0.091942 | 2.2329 | 2.8591 | -2.3893 | 1.4897 | -0.68847 | 0.2357 | -0.050731 | 0.0048419 | [12 empty] | 1.2716 | 1.407 | 1.367 | -0.15367 | 0.00054901
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 817 | 3.6072 | [20 empty] | -0.091942 | 2.2356 | 2.8661 | -2.3981 | 1.4971 | -0.6927 | 0.23744 | -0.051168 | 0.0048896 | [12 empty] | 1.2716 | 1.4087 | 1.3703 | -0.15424 | 0.0005517
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 818 | 3.6126 | [20 empty] | -0.091942 | 2.2384 | 2.8731 | -2.4069 | 1.5044 | -0.69695 | 0.23918 | -0.051608 | 0.0049377 | [12 empty] | 1.2716 | 1.4104 | 1.3737 | -0.15481 | 0.00055441
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 819 | 3.618 | [20 empty] | -0.091942 | 2.2411 | 2.8801 | -2.4157 | 1.5118 | -0.70122 | 0.24094 | -0.052051 | 0.0049862 | [12 empty] | 1.2716 | 1.4121 | 1.3771 | -0.15538 | 0.00055713
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 820 | 3.6234 | [20 empty] | -0.091942 | 2.2438 | 2.8872 | -2.4246 | 1.5192 | -0.70551 | 0.24271 | -0.052498 | 0.0050351 | [12 empty] | 1.2716 | 1.4139 | 1.3804 | -0.15595 | 0.00055985
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 821 | 3.6288 | [20 empty] | -0.091942 | 2.2466 | 2.8942 | -2.4334 | 1.5266 | -0.70983 | 0.2445 | -0.052948 | 0.0050845 | [12 empty] | 1.2716 | 1.4156 | 1.3838 | -0.15652 | 0.00056259
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 822 | 3.6342 | [20 empty] | -0.091942 | 2.2493 | 2.9013 | -2.4424 | 1.534 | -0.71416 | 0.24629 | -0.053401 | 0.0051342 | [12 empty] | 1.2716 | 1.4173 | 1.3872 | -0.15709 | 0.00056533
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 823 | 3.6396 | [20 empty] | -0.091942 | 2.252 | 2.9083 | -2.4513 | 1.5415 | -0.71852 | 0.24809 | -0.053857 | 0.0051844 | [12 empty] | 1.2716 | 1.419 | 1.3905 | -0.15766 | 0.00056809
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 824 | 3.645 | [20 empty] | -0.091942 | 2.2548 | 2.9154 | -2.4602 | 1.549 | -0.72289 | 0.24991 | -0.054317 | 0.005235 | [12 empty] | 1.2716 | 1.4208 | 1.3939 | -0.15824 | 0.00057086
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 825 | 3.6504 | [20 empty] | -0.091942 | 2.2575 | 2.9225 | -2.4692 | 1.5566 | -0.72729 | 0.25173 | -0.05478 | 0.005286 | [12 empty] | 1.2716 | 1.4225 | 1.3973 | -0.15882 | 0.00057363
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 826 | 3.6558 | [20 empty] | -0.091942 | 2.2603 | 2.9296 | -2.4782 | 1.5641 | -0.73171 | 0.25357 | -0.055246 | 0.0053375 | [12 empty] | 1.2716 | 1.4242 | 1.4007 | -0.15939 | 0.00057642
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 827 | 3.6612 | [20 empty] | -0.091942 | 2.263 | 2.9367 | -2.4872 | 1.5717 | -0.73615 | 0.25541 | -0.055716 | 0.0053894 | [12 empty] | 1.2716 | 1.4259 | 1.4041 | -0.15997 | 0.00057922
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 828 | 3.6666 | [20 empty] | -0.091942 | 2.2657 | 2.9438 | -2.4962 | 1.5793 | -0.74061 | 0.25727 | -0.05619 | 0.0054418 | [12 empty] | 1.2716 | 1.4277 | 1.4075 | -0.16055 | 0.00058202
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 829 | 3.672 | [20 empty] | -0.091942 | 2.2685 | 2.9509 | -2.5053 | 1.587 | -0.74509 | 0.25914 | -0.056666 | 0.0054946 | [12 empty] | 1.2716 | 1.4294 | 1.4109 | -0.16114 | 0.00058484
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 830 | 3.6774 | [20 empty] | -0.091942 | 2.2712 | 2.958 | -2.5144 | 1.5946 | -0.7496 | 0.26102 | -0.057147 | 0.0055478 | [12 empty] | 1.2716 | 1.4311 | 1.4143 | -0.16172 | 0.00058767
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 831 | 3.6828 | [20 empty] | -0.091942 | 2.2739 | 2.9651 | -2.5235 | 1.6023 | -0.75412 | 0.26292 | -0.05763 | 0.0056015 | [12 empty] | 1.2716 | 1.4328 | 1.4177 | -0.16231 | 0.0005905
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 832 | 3.6882 | [20 empty] | -0.091942 | 2.2767 | 2.9723 | -2.5326 | 1.6101 | -0.75867 | 0.26482 | -0.058118 | 0.0056557 | [12 empty] | 1.2716 | 1.4346 | 1.4211 | -0.16289 | 0.00059335
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 833 | 3.6936 | [20 empty] | -0.091942 | 2.2794 | 2.9794 | -2.5417 | 1.6178 | -0.76324 | 0.26674 | -0.058608 | 0.0057103 | [12 empty] | 1.2716 | 1.4363 | 1.4245 | -0.16348 | 0.00059621
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 834 | 3.699 | [20 empty] | -0.091942 | 2.2821 | 2.9866 | -2.5509 | 1.6256 | -0.76783 | 0.26866 | -0.059103 | 0.0057654 | [12 empty] | 1.2716 | 1.438 | 1.428 | -0.16407 | 0.00059908
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 835 | 3.7044 | [20 empty] | -0.091942 | 2.2849 | 2.9938 | -2.5601 | 1.6334 | -0.77245 | 0.2706 | -0.0596 | 0.0058209 | [12 empty] | 1.2716 | 1.4397 | 1.4314 | -0.16466 | 0.00060196
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 836 | 3.7098 | [20 empty] | -0.091942 | 2.2876 | 3.0009 | -2.5693 | 1.6413 | -0.77708 | 0.27255 | -0.060102 | 0.0058769 | [12 empty] | 1.2716 | 1.4415 | 1.4348 | -0.16525 | 0.00060484
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 837 | 3.7152 | [20 empty] | -0.091942 | 2.2904 | 3.0081 | -2.5785 | 1.6491 | -0.78174 | 0.27451 | -0.060607 | 0.0059334 | [12 empty] | 1.2716 | 1.4432 | 1.4383 | -0.16585 | 0.00060774
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 838 | 3.7206 | [20 empty] | -0.091942 | 2.2931 | 3.0153 | -2.5878 | 1.657 | -0.78642 | 0.27649 | -0.061116 | 0.0059903 | [12 empty] | 1.2716 | 1.4449 | 1.4417 | -0.16644 | 0.00061065
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 839 | 3.7261 | [20 empty] | -0.091942 | 2.2958 | 3.0225 | -2.597 | 1.6649 | -0.79113 | 0.27847 | -0.061628 | 0.0060478 | [12 empty] | 1.2716 | 1.4466 | 1.4451 | -0.16704 | 0.00061357
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 840 | 3.7315 | [20 empty] | -0.091942 | 2.2986 | 3.0297 | -2.6063 | 1.6729 | -0.79585 | 0.28047 | -0.062144 | 0.0061057 | [12 empty] | 1.2716 | 1.4484 | 1.4486 | -0.16764 | 0.0006165
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 841 | 3.7369 | [20 empty] | -0.091942 | 2.3013 | 3.0369 | -2.6157 | 1.6809 | -0.8006 | 0.28248 | -0.062664 | 0.0061641 | [12 empty] | 1.2716 | 1.4501 | 1.452 | -0.16824 | 0.00061944
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 842 | 3.7423 | [20 empty] | -0.091942 | 2.304 | 3.0442 | -2.625 | 1.6889 | -0.80537 | 0.2845 | -0.063187 | 0.0062229 | [12 empty] | 1.2716 | 1.4518 | 1.4555 | -0.16884 | 0.0006224
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 843 | 3.7477 | [20 empty] | -0.091942 | 2.3068 | 3.0514 | -2.6344 | 1.6969 | -0.81017 | 0.28654 | -0.063714 | 0.0062823 | [12 empty] | 1.2716 | 1.4535 | 1.459 | -0.16944 | 0.00062536
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 844 | 3.7531 | [20 empty] | -0.091942 | 2.3095 | 3.0586 | -2.6437 | 1.705 | -0.81498 | 0.28858 | -0.064245 | 0.0063422 | [12 empty] | 1.2716 | 1.4553 | 1.4624 | -0.17004 | 0.00062833
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 845 | 3.7586 | [20 empty] | -0.091942 | 2.3122 | 3.0659 | -2.6532 | 1.7131 | -0.81982 | 0.29064 | -0.06478 | 0.0064025 | [12 empty] | 1.2716 | 1.457 | 1.4659 | -0.17065 | 0.00063131
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 846 | 3.764 | [20 empty] | -0.091942 | 2.315 | 3.0731 | -2.6626 | 1.7212 | -0.82469 | 0.29271 | -0.065319 | 0.0064634 | [12 empty] | 1.2716 | 1.4587 | 1.4694 | -0.17125 | 0.00063431
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 847 | 3.7694 | [20 empty] | -0.091942 | 2.3177 | 3.0804 | -2.672 | 1.7294 | -0.82957 | 0.29479 | -0.065861 | 0.0065248 | [12 empty] | 1.2716 | 1.4604 | 1.4728 | -0.17186 | 0.00063731
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 848 | 3.7748 | [20 empty] | -0.091942 | 2.3205 | 3.0877 | -2.6815 | 1.7375 | -0.83448 | 0.29688 | -0.066407 | 0.0065867 | [12 empty] | 1.2716 | 1.4622 | 1.4763 | -0.17247 | 0.00064033
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 849 | 3.7802 | [20 empty] | -0.091942 | 2.3232 | 3.095 | -2.691 | 1.7458 | -0.83941 | 0.29899 | -0.066957 | 0.0066491 | [12 empty] | 1.2716 | 1.4639 | 1.4798 | -0.17308 | 0.00064335
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 850 | 3.7857 | [20 empty] | -0.091942 | 2.3259 | 3.1023 | -2.7005 | 1.754 | -0.84437 | 0.30111 | -0.067511 | 0.006712 | [12 empty] | 1.2716 | 1.4656 | 1.4833 | -0.1737 | 0.00064639
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 851 | 3.7911 | [20 empty] | -0.091942 | 2.3287 | 3.1096 | -2.7101 | 1.7623 | -0.84935 | 0.30324 | -0.068069 | 0.0067754 | [12 empty] | 1.2716 | 1.4673 | 1.4868 | -0.17431 | 0.00064944
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 852 | 3.7965 | [20 empty] | -0.091942 | 2.3314 | 3.1169 | -2.7196 | 1.7706 | -0.85435 | 0.30539 | -0.068631 | 0.0068394 | [12 empty] | 1.2716 | 1.469 | 1.4903 | -0.17492 | 0.00065249
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 853 | 3.8019 | [20 empty] | -0.091942 | 2.3341 | 3.1242 | -2.7292 | 1.7789 | -0.85937 | 0.30754 | -0.069197 | 0.0069038 | [12 empty] | 1.2716 | 1.4708 | 1.4938 | -0.17554 | 0.00065556
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 854 | 3.8074 | [20 empty] | -0.091942 | 2.3369 | 3.1315 | -2.7388 | 1.7872 | -0.86442 | 0.30971 | -0.069767 | 0.0069689 | [12 empty] | 1.2716 | 1.4725 | 1.4973 | -0.17616 | 0.00065864
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 855 | 3.8128 | [20 empty] | -0.091942 | 2.3396 | 3.1389 | -2.7485 | 1.7956 | -0.8695 | 0.3119 | -0.070341 | 0.0070344 | [12 empty] | 1.2716 | 1.4742 | 1.5008 | -0.17678 | 0.00066173
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 856 | 3.8182 | [20 empty] | -0.091942 | 2.3423 | 3.1462 | -2.7581 | 1.804 | -0.87459 | 0.31409 | -0.070919 | 0.0071005 | [12 empty] | 1.2716 | 1.4759 | 1.5043 | -0.1774 | 0.00066483
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 857 | 3.8237 | [20 empty] | -0.091942 | 2.3451 | 3.1536 | -2.7678 | 1.8125 | -0.87971 | 0.3163 | -0.071501 | 0.0071671 | [12 empty] | 1.2716 | 1.4777 | 1.5078 | -0.17802 | 0.00066795
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 858 | 3.8291 | [20 empty] | -0.091942 | 2.3478 | 3.1609 | -2.7775 | 1.821 | -0.88486 | 0.31852 | -0.072087 | 0.0072343 | [12 empty] | 1.2716 | 1.4794 | 1.5113 | -0.17865 | 0.00067107
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 859 | 3.8345 | [20 empty] | -0.091942 | 2.3506 | 3.1683 | -2.7872 | 1.8295 | -0.89003 | 0.32075 | -0.072677 | 0.007302 | [12 empty] | 1.2716 | 1.4811 | 1.5149 | -0.17927 | 0.0006742
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 860 | 3.84 | [20 empty] | -0.091942 | 2.3533 | 3.1757 | -2.797 | 1.838 | -0.89522 | 0.323 | -0.073271 | 0.0073703 | [12 empty] | 1.2716 | 1.4828 | 1.5184 | -0.1799 | 0.00067735
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 861 | 3.8454 | [20 empty] | -0.091942 | 2.356 | 3.1831 | -2.8067 | 1.8466 | -0.90044 | 0.32526 | -0.07387 | 0.0074392 | [12 empty] | 1.2716 | 1.4846 | 1.5219 | -0.18053 | 0.0006805
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 862 | 3.8508 | [20 empty] | -0.091942 | 2.3588 | 3.1905 | -2.8165 | 1.8552 | -0.90568 | 0.32753 | -0.074472 | 0.0075086 | [12 empty] | 1.2716 | 1.4863 | 1.5255 | -0.18116 | 0.00068367
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 863 | 3.8563 | [20 empty] | -0.091942 | 2.3615 | 3.1979 | -2.8263 | 1.8638 | -0.91094 | 0.32982 | -0.075079 | 0.0075785 | [12 empty] | 1.2716 | 1.488 | 1.529 | -0.18179 | 0.00068685
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 864 | 3.8617 | [20 empty] | -0.091942 | 2.3642 | 3.2053 | -2.8362 | 1.8724 | -0.91623 | 0.33212 | -0.07569 | 0.0076491 | [12 empty] | 1.2716 | 1.4897 | 1.5325 | -0.18242 | 0.00069004
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 865 | 3.8671 | [20 empty] | -0.091942 | 2.367 | 3.2127 | -2.846 | 1.8811 | -0.92155 | 0.33443 | -0.076306 | 0.0077202 | [12 empty] | 1.2716 | 1.4915 | 1.5361 | -0.18305 | 0.00069324
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 866 | 3.8726 | [20 empty] | -0.091942 | 2.3697 | 3.2202 | -2.8559 | 1.8898 | -0.92689 | 0.33676 | -0.076925 | 0.0077919 | [12 empty] | 1.2716 | 1.4932 | 1.5396 | -0.18369 | 0.00069645
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 867 | 3.878 | [20 empty] | -0.091942 | 2.3724 | 3.2276 | -2.8658 | 1.8986 | -0.93225 | 0.3391 | -0.077549 | 0.0078641 | [12 empty] | 1.2716 | 1.4949 | 1.5432 | -0.18433 | 0.00069967
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 868 | 3.8835 | [20 empty] | -0.091942 | 2.3752 | 3.2351 | -2.8758 | 1.9074 | -0.93764 | 0.34145 | -0.078178 | 0.007937 | [12 empty] | 1.2716 | 1.4966 | 1.5468 | -0.18497 | 0.00070291
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 869 | 3.8889 | [20 empty] | -0.091942 | 2.3779 | 3.2425 | -2.8857 | 1.9162 | -0.94305 | 0.34382 | -0.07881 | 0.0080105 | [12 empty] | 1.2716 | 1.4984 | 1.5503 | -0.18561 | 0.00070615
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 870 | 3.8944 | [20 empty] | -0.091942 | 2.3807 | 3.25 | -2.8957 | 1.925 | -0.94849 | 0.3462 | -0.079447 | 0.0080845 | [12 empty] | 1.2716 | 1.5001 | 1.5539 | -0.18625 | 0.00070941
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 871 | 3.8998 | [20 empty] | -0.091942 | 2.3834 | 3.2575 | -2.9057 | 1.9339 | -0.95395 | 0.3486 | -0.080089 | 0.0081591 | [12 empty] | 1.2716 | 1.5018 | 1.5575 | -0.18689 | 0.00071267
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 872 | 3.9053 | [20 empty] | -0.091942 | 2.3861 | 3.2649 | -2.9157 | 1.9428 | -0.95944 | 0.351 | -0.080735 | 0.0082344 | [12 empty] | 1.2716 | 1.5035 | 1.5611 | -0.18753 | 0.00071595
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 873 | 3.9107 | [20 empty] | -0.091942 | 2.3889 | 3.2724 | -2.9257 | 1.9517 | -0.96496 | 0.35343 | -0.081385 | 0.0083102 | [12 empty] | 1.2716 | 1.5053 | 1.5646 | -0.18818 | 0.00071924
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 874 | 3.9161 | [20 empty] | -0.091942 | 2.3916 | 3.2799 | -2.9358 | 1.9606 | -0.9705 | 0.35586 | -0.08204 | 0.0083867 | [12 empty] | 1.2716 | 1.507 | 1.5682 | -0.18883 | 0.00072254
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 875 | 3.9216 | [20 empty] | -0.091942 | 2.3943 | 3.2874 | -2.9459 | 1.9696 | -0.97606 | 0.35831 | -0.082699 | 0.0084638 | [12 empty] | 1.2716 | 1.5087 | 1.5718 | -0.18948 | 0.00072586
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 876 | 3.927 | [20 empty] | -0.091942 | 2.3971 | 3.295 | -2.956 | 1.9786 | -0.98165 | 0.36078 | -0.083363 | 0.0085415 | [12 empty] | 1.2716 | 1.5104 | 1.5754 | -0.19013 | 0.00072918
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 877 | 3.9325 | [20 empty] | -0.091942 | 2.3998 | 3.3025 | -2.9661 | 1.9877 | -0.98727 | 0.36325 | -0.084031 | 0.0086198 | [12 empty] | 1.2716 | 1.5122 | 1.579 | -0.19078 | 0.00073252
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 878 | 3.9379 | [20 empty] | -0.091942 | 2.4025 | 3.31 | -2.9763 | 1.9968 | -0.99291 | 0.36575 | -0.084704 | 0.0086987 | [12 empty] | 1.2716 | 1.5139 | 1.5826 | -0.19143 | 0.00073586
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 879 | 3.9434 | [20 empty] | -0.091942 | 2.4053 | 3.3176 | -2.9865 | 2.0059 | -0.99858 | 0.36825 | -0.085382 | 0.0087783 | [12 empty] | 1.2716 | 1.5156 | 1.5862 | -0.19209 | 0.00073922
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 880 | 3.9489 | [20 empty] | -0.091942 | 2.408 | 3.3251 | -2.9967 | 2.015 | -1.0043 | 0.37077 | -0.086064 | 0.0088585 | [12 empty] | 1.2716 | 1.5173 | 1.5898 | -0.19274 | 0.00074259
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 881 | 3.9543 | [20 empty] | -0.091942 | 2.4108 | 3.3327 | -3.0069 | 2.0242 | -1.01 | 0.37331 | -0.086751 | 0.0089394 | [12 empty] | 1.2716 | 1.5191 | 1.5934 | -0.1934 | 0.00074597
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 882 | 3.9598 | [20 empty] | -0.091942 | 2.4135 | 3.3403 | -3.0172 | 2.0334 | -1.0157 | 0.37586 | -0.087443 | 0.0090209 | [12 empty] | 1.2716 | 1.5208 | 1.5971 | -0.19406 | 0.00074936
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 883 | 3.9652 | [20 empty] | -0.091942 | 2.4162 | 3.3478 | -3.0274 | 2.0427 | -1.0215 | 0.37842 | -0.088139 | 0.009103 | [12 empty] | 1.2716 | 1.5225 | 1.6007 | -0.19472 | 0.00075277
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 884 | 3.9707 | [20 empty] | -0.091942 | 2.419 | 3.3554 | -3.0377 | 2.0519 | -1.0273 | 0.381 | -0.08884 | 0.0091858 | [12 empty] | 1.2716 | 1.5242 | 1.6043 | -0.19538 | 0.00075618
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 885 | 3.9761 | [20 empty] | -0.091942 | 2.4217 | 3.363 | -3.0481 | 2.0612 | -1.0331 | 0.38359 | -0.089546 | 0.0092693 | [12 empty] | 1.2716 | 1.5259 | 1.6079 | -0.19605 | 0.00075961
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 886 | 3.9816 | [20 empty] | -0.091942 | 2.4244 | 3.3706 | -3.0584 | 2.0706 | -1.039 | 0.3862 | -0.090257 | 0.0093534 | [12 empty] | 1.2716 | 1.5277 | 1.6116 | -0.19671 | 0.00076305
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 887 | 3.9871 | [20 empty] | -0.091942 | 2.4272 | 3.3782 | -3.0688 | 2.0799 | -1.0449 | 0.38883 | -0.090972 | 0.0094382 | [12 empty] | 1.2716 | 1.5294 | 1.6152 | -0.19738 | 0.0007665
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 888 | 3.9925 | [20 empty] | -0.091942 | 2.4299 | 3.3859 | -3.0792 | 2.0893 | -1.0508 | 0.39146 | -0.091693 | 0.0095236 | [12 empty] | 1.2716 | 1.5311 | 1.6189 | -0.19805 | 0.00076996
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 889 | 3.998 | [20 empty] | -0.091942 | 2.4326 | 3.3935 | -3.0896 | 2.0987 | -1.0567 | 0.39412 | -0.092418 | 0.0096098 | [12 empty] | 1.2716 | 1.5328 | 1.6225 | -0.19872 | 0.00077344
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 890 | 4.0035 | [20 empty] | -0.091942 | 2.4354 | 3.4011 | -3.1 | 2.1082 | -1.0626 | 0.39678 | -0.093148 | 0.0096966 | [12 empty] | 1.2716 | 1.5346 | 1.6262 | -0.19939 | 0.00077692
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 891 | 4.0089 | [20 empty] | -0.091942 | 2.4381 | 3.4088 | -3.1105 | 2.1177 | -1.0686 | 0.39947 | -0.093883 | 0.0097841 | [12 empty] | 1.2716 | 1.5363 | 1.6298 | -0.20006 | 0.00078042
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 892 | 4.0144 | [20 empty] | -0.091942 | 2.4409 | 3.4164 | -3.121 | 2.1272 | -1.0746 | 0.40216 | -0.094623 | 0.0098723 | [12 empty] | 1.2716 | 1.538 | 1.6335 | -0.20074 | 0.00078393
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 893 | 4.0198 | [20 empty] | -0.091942 | 2.4436 | 3.4241 | -3.1315 | 2.1368 | -1.0807 | 0.40488 | -0.095368 | 0.0099612 | [12 empty] | 1.2716 | 1.5397 | 1.6372 | -0.20141 | 0.00078745
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 894 | 4.0253 | [20 empty] | -0.091942 | 2.4463 | 3.4318 | -3.142 | 2.1464 | -1.0867 | 0.4076 | -0.096118 | 0.010051 | [12 empty] | 1.2716 | 1.5415 | 1.6408 | -0.20209 | 0.00079099
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 895 | 4.0308 | [20 empty] | -0.091942 | 2.4491 | 3.4394 | -3.1526 | 2.156 | -1.0928 | 0.41035 | -0.096874 | 0.010141 | [12 empty] | 1.2716 | 1.5432 | 1.6445 | -0.20277 | 0.00079453
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 896 | 4.0363 | [20 empty] | -0.091942 | 2.4518 | 3.4471 | -3.1631 | 2.1656 | -1.0989 | 0.41311 | -0.097634 | 0.010232 | [12 empty] | 1.2716 | 1.5449 | 1.6482 | -0.20345 | 0.00079809
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 897 | 4.0417 | [20 empty] | -0.091942 | 2.4545 | 3.4548 | -3.1737 | 2.1753 | -1.1051 | 0.41588 | -0.098399 | 0.010324 | [12 empty] | 1.2716 | 1.5466 | 1.6518 | -0.20413 | 0.00080166
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 898 | 4.0472 | [20 empty] | -0.091942 | 2.4573 | 3.4625 | -3.1844 | 2.185 | -1.1113 | 0.41867 | -0.09917 | 0.010416 | [12 empty] | 1.2716 | 1.5484 | 1.6555 | -0.20481 | 0.00080524
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 899 | 4.0527 | [20 empty] | -0.091942 | 2.46 | 3.4703 | -3.195 | 2.1948 | -1.1175 | 0.42147 | -0.099945 | 0.010509 | [12 empty] | 1.2716 | 1.5501 | 1.6592 | -0.2055 | 0.00080883
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 900 | 4.0581 | [20 empty] | -0.091942 | 2.4627 | 3.478 | -3.2057 | 2.2046 | -1.1237 | 0.42429 | -0.10073 | 0.010603 | [12 empty] | 1.2716 | 1.5518 | 1.6629 | -0.20619 | 0.00081243
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 901 | 4.0636 | [20 empty] | -0.091942 | 2.4655 | 3.4857 | -3.2164 | 2.2144 | -1.13 | 0.42713 | -0.10151 | 0.010698 | [12 empty] | 1.2716 | 1.5535 | 1.6666 | -0.20687 | 0.00081605
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 902 | 4.0691 | [20 empty] | -0.091942 | 2.4682 | 3.4935 | -3.2271 | 2.2242 | -1.1362 | 0.42998 | -0.1023 | 0.010793 | [12 empty] | 1.2716 | 1.5553 | 1.6703 | -0.20756 | 0.00081968
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 903 | 4.0746 | [20 empty] | -0.091942 | 2.471 | 3.5012 | -3.2378 | 2.2341 | -1.1426 | 0.43285 | -0.1031 | 0.010889 | [12 empty] | 1.2716 | 1.557 | 1.674 | -0.20825 | 0.00082332
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 904 | 4.0801 | [20 empty] | -0.091942 | 2.4737 | 3.509 | -3.2486 | 2.244 | -1.1489 | 0.43574 | -0.1039 | 0.010986 | [12 empty] | 1.2716 | 1.5587 | 1.6777 | -0.20895 | 0.00082697
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 905 | 4.0855 | [20 empty] | -0.091942 | 2.4764 | 3.5167 | -3.2594 | 2.254 | -1.1553 | 0.43864 | -0.10471 | 0.011084 | [12 empty] | 1.2716 | 1.5604 | 1.6814 | -0.20964 | 0.00083064
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 906 | 4.091 | [20 empty] | -0.091942 | 2.4792 | 3.5245 | -3.2702 | 2.2639 | -1.1617 | 0.44155 | -0.10552 | 0.011182 | [12 empty] | 1.2716 | 1.5622 | 1.6852 | -0.21034 | 0.00083432
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 907 | 4.0965 | [20 empty] | -0.091942 | 2.4819 | 3.5323 | -3.2811 | 2.274 | -1.1681 | 0.44448 | -0.10634 | 0.011281 | [12 empty] | 1.2716 | 1.5639 | 1.6889 | -0.21103 | 0.00083801
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 908 | 4.102 | [20 empty] | -0.091942 | 2.4846 | 3.5401 | -3.2919 | 2.284 | -1.1745 | 0.44743 | -0.10716 | 0.011381 | [12 empty] | 1.2716 | 1.5656 | 1.6926 | -0.21173 | 0.00084171
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 909 | 4.1075 | [20 empty] | -0.091942 | 2.4874 | 3.5479 | -3.3028 | 2.2941 | -1.181 | 0.4504 | -0.10799 | 0.011482 | [12 empty] | 1.2716 | 1.5673 | 1.6963 | -0.21243 | 0.00084542
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 910 | 4.1129 | [20 empty] | -0.091942 | 2.4901 | 3.5557 | -3.3137 | 2.3042 | -1.1875 | 0.45338 | -0.10883 | 0.011583 | [12 empty] | 1.2716 | 1.5691 | 1.7001 | -0.21314 | 0.00084915
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 911 | 4.1184 | [20 empty] | -0.091942 | 2.4928 | 3.5635 | -3.3247 | 2.3143 | -1.1941 | 0.45638 | -0.10967 | 0.011685 | [12 empty] | 1.2716 | 1.5708 | 1.7038 | -0.21384 | 0.00085289
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 912 | 4.1239 | [20 empty] | -0.091942 | 2.4956 | 3.5714 | -3.3356 | 2.3245 | -1.2006 | 0.45939 | -0.11051 | 0.011788 | [12 empty] | 1.2716 | 1.5725 | 1.7076 | -0.21454 | 0.00085664
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 913 | 4.1294 | [20 empty] | -0.091942 | 2.4983 | 3.5792 | -3.3466 | 2.3347 | -1.2072 | 0.46242 | -0.11136 | 0.011892 | [12 empty] | 1.2716 | 1.5742 | 1.7113 | -0.21525 | 0.0008604
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 914 | 4.1349 | [20 empty] | -0.091942 | 2.5011 | 3.587 | -3.3576 | 2.345 | -1.2139 | 0.46547 | -0.11222 | 0.011997 | [12 empty] | 1.2716 | 1.576 | 1.7151 | -0.21596 | 0.00086418
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 915 | 4.1404 | [20 empty] | -0.091942 | 2.5038 | 3.5949 | -3.3687 | 2.3552 | -1.2205 | 0.46853 | -0.11308 | 0.012102 | [12 empty] | 1.2716 | 1.5777 | 1.7188 | -0.21667 | 0.00086797
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 916 | 4.1459 | [20 empty] | -0.091942 | 2.5065 | 3.6027 | -3.3797 | 2.3656 | -1.2272 | 0.47161 | -0.11395 | 0.012209 | [12 empty] | 1.2716 | 1.5794 | 1.7226 | -0.21738 | 0.00087177
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 917 | 4.1514 | [20 empty] | -0.091942 | 2.5093 | 3.6106 | -3.3908 | 2.3759 | -1.2339 | 0.47471 | -0.11482 | 0.012316 | [12 empty] | 1.2716 | 1.5811 | 1.7263 | -0.21809 | 0.00087558
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 918 | 4.1568 | [20 empty] | -0.091942 | 2.512 | 3.6185 | -3.4019 | 2.3863 | -1.2407 | 0.47782 | -0.1157 | 0.012423 | [12 empty] | 1.2716 | 1.5828 | 1.7301 | -0.21881 | 0.00087941
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 919 | 4.1623 | [20 empty] | -0.091942 | 2.5147 | 3.6264 | -3.413 | 2.3967 | -1.2474 | 0.48096 | -0.11659 | 0.012532 | [12 empty] | 1.2716 | 1.5846 | 1.7339 | -0.21952 | 0.00088324
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 920 | 4.1678 | [20 empty] | -0.091942 | 2.5175 | 3.6343 | -3.4242 | 2.4072 | -1.2542 | 0.4841 | -0.11748 | 0.012642 | [12 empty] | 1.2716 | 1.5863 | 1.7376 | -0.22024 | 0.00088709
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 921 | 4.1733 | [20 empty] | -0.091942 | 2.5202 | 3.6422 | -3.4354 | 2.4176 | -1.2611 | 0.48727 | -0.11838 | 0.012752 | [12 empty] | 1.2716 | 1.588 | 1.7414 | -0.22096 | 0.00089096
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 922 | 4.1788 | [20 empty] | -0.091942 | 2.5229 | 3.6501 | -3.4466 | 2.4282 | -1.2679 | 0.49045 | -0.11928 | 0.012863 | [12 empty] | 1.2716 | 1.5897 | 1.7452 | -0.22168 | 0.00089483
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 923 | 4.1843 | [20 empty] | -0.091942 | 2.5257 | 3.658 | -3.4578 | 2.4387 | -1.2748 | 0.49365 | -0.12019 | 0.012975 | [12 empty] | 1.2716 | 1.5915 | 1.749 | -0.2224 | 0.00089872
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 924 | 4.1898 | [20 empty] | -0.091942 | 2.5284 | 3.666 | -3.469 | 2.4493 | -1.2817 | 0.49687 | -0.1211 | 0.013088 | [12 empty] | 1.2716 | 1.5932 | 1.7528 | -0.22312 | 0.00090262
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 925 | 4.1953 | [20 empty] | -0.091942 | 2.5312 | 3.6739 | -3.4803 | 2.4599 | -1.2887 | 0.50011 | -0.12202 | 0.013202 | [12 empty] | 1.2716 | 1.5949 | 1.7566 | -0.22385 | 0.00090654
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 926 | 4.2008 | [20 empty] | -0.091942 | 2.5339 | 3.6818 | -3.4916 | 2.4706 | -1.2957 | 0.50336 | -0.12295 | 0.013316 | [12 empty] | 1.2716 | 1.5966 | 1.7604 | -0.22458 | 0.00091046
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 927 | 4.2063 | [20 empty] | -0.091942 | 2.5366 | 3.6898 | -3.5029 | 2.4813 | -1.3027 | 0.50663 | -0.12388 | 0.013432 | [12 empty] | 1.2716 | 1.5984 | 1.7642 | -0.2253 | 0.0009144
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 928 | 4.2118 | [20 empty] | -0.091942 | 2.5394 | 3.6978 | -3.5143 | 2.492 | -1.3097 | 0.50992 | -0.12482 | 0.013548 | [12 empty] | 1.2716 | 1.6001 | 1.768 | -0.22603 | 0.00091835
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 929 | 4.2173 | [20 empty] | -0.091942 | 2.5421 | 3.7057 | -3.5257 | 2.5027 | -1.3168 | 0.51322 | -0.12576 | 0.013665 | [12 empty] | 1.2716 | 1.6018 | 1.7718 | -0.22677 | 0.00092232
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 930 | 4.2228 | [20 empty] | -0.091942 | 2.5448 | 3.7137 | -3.5371 | 2.5135 | -1.3239 | 0.51655 | -0.12671 | 0.013784 | [12 empty] | 1.2716 | 1.6035 | 1.7756 | -0.2275 | 0.0009263
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 931 | 4.2283 | [20 empty] | -0.091942 | 2.5476 | 3.7217 | -3.5485 | 2.5244 | -1.331 | 0.51989 | -0.12767 | 0.013903 | [12 empty] | 1.2716 | 1.6053 | 1.7794 | -0.22823 | 0.00093029
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 932 | 4.2338 | [20 empty] | -0.091942 | 2.5503 | 3.7297 | -3.5599 | 2.5352 | -1.3382 | 0.52325 | -0.12863 | 0.014023 | [12 empty] | 1.2716 | 1.607 | 1.7833 | -0.22897 | 0.00093429
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 933 | 4.2393 | [20 empty] | -0.091942 | 2.553 | 3.7377 | -3.5714 | 2.5461 | -1.3454 | 0.52663 | -0.1296 | 0.014143 | [12 empty] | 1.2716 | 1.6087 | 1.7871 | -0.22971 | 0.00093831
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 934 | 4.2448 | [20 empty] | -0.091942 | 2.5558 | 3.7457 | -3.5829 | 2.5571 | -1.3526 | 0.53002 | -0.13058 | 0.014265 | [12 empty] | 1.2716 | 1.6104 | 1.7909 | -0.23045 | 0.00094234
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 935 | 4.2503 | [20 empty] | -0.091942 | 2.5585 | 3.7538 | -3.5944 | 2.568 | -1.3599 | 0.53344 | -0.13156 | 0.014388 | [12 empty] | 1.2716 | 1.6122 | 1.7948 | -0.23119 | 0.00094638
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 936 | 4.2559 | [20 empty] | -0.091942 | 2.5613 | 3.7618 | -3.606 | 2.579 | -1.3672 | 0.53687 | -0.13255 | 0.014511 | [12 empty] | 1.2716 | 1.6139 | 1.7986 | -0.23193 | 0.00095043
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 937 | 4.2614 | [20 empty] | -0.091942 | 2.564 | 3.7698 | -3.6175 | 2.5901 | -1.3745 | 0.54032 | -0.13354 | 0.014636 | [12 empty] | 1.2716 | 1.6156 | 1.8025 | -0.23268 | 0.0009545
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 938 | 4.2669 | [20 empty] | -0.091942 | 2.5667 | 3.7779 | -3.6291 | 2.6011 | -1.3818 | 0.54379 | -0.13454 | 0.014761 | [12 empty] | 1.2716 | 1.6173 | 1.8063 | -0.23342 | 0.00095858
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 939 | 4.2724 | [20 empty] | -0.091942 | 2.5695 | 3.7859 | -3.6407 | 2.6122 | -1.3892 | 0.54728 | -0.13555 | 0.014888 | [12 empty] | 1.2716 | 1.6191 | 1.8102 | -0.23417 | 0.00096268
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 940 | 4.2779 | [20 empty] | -0.091942 | 2.5722 | 3.794 | -3.6524 | 2.6234 | -1.3966 | 0.55078 | -0.13656 | 0.015015 | [12 empty] | 1.2716 | 1.6208 | 1.814 | -0.23492 | 0.00096679
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 941 | 4.2834 | [20 empty] | -0.091942 | 2.5749 | 3.8021 | -3.6641 | 2.6346 | -1.4041 | 0.55431 | -0.13759 | 0.015143 | [12 empty] | 1.2716 | 1.6225 | 1.8179 | -0.23567 | 0.00097091
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 942 | 4.2889 | [20 empty] | -0.091942 | 2.5777 | 3.8102 | -3.6757 | 2.6458 | -1.4115 | 0.55785 | -0.13861 | 0.015272 | [12 empty] | 1.2716 | 1.6242 | 1.8217 | -0.23642 | 0.00097504
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 943 | 4.2944 | [20 empty] | -0.091942 | 2.5804 | 3.8183 | -3.6875 | 2.657 | -1.419 | 0.56141 | -0.13965 | 0.015403 | [12 empty] | 1.2716 | 1.626 | 1.8256 | -0.23717 | 0.00097919
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 944 | 4.3 | [20 empty] | -0.091942 | 2.5831 | 3.8264 | -3.6992 | 2.6683 | -1.4266 | 0.56499 | -0.14069 | 0.015534 | [12 empty] | 1.2716 | 1.6277 | 1.8295 | -0.23793 | 0.00098335
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 945 | 4.3055 | [20 empty] | -0.091942 | 2.5859 | 3.8345 | -3.711 | 2.6797 | -1.4342 | 0.5686 | -0.14173 | 0.015666 | [12 empty] | 1.2716 | 1.6294 | 1.8334 | -0.23869 | 0.00098752
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 946 | 4.311 | [20 empty] | -0.091942 | 2.5886 | 3.8426 | -3.7228 | 2.691 | -1.4418 | 0.57222 | -0.14278 | 0.015799 | [12 empty] | 1.2716 | 1.6311 | 1.8372 | -0.23944 | 0.00099171
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 947 | 4.3165 | [20 empty] | -0.091942 | 2.5914 | 3.8507 | -3.7346 | 2.7024 | -1.4494 | 0.57585 | -0.14384 | 0.015933 | [12 empty] | 1.2716 | 1.6329 | 1.8411 | -0.2402 | 0.00099591
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 948 | 4.322 | [20 empty] | -0.091942 | 2.5941 | 3.8589 | -3.7464 | 2.7138 | -1.4571 | 0.57951 | -0.14491 | 0.016068 | [12 empty] | 1.2716 | 1.6346 | 1.845 | -0.24097 | 0.0010001
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 949 | 4.3276 | [20 empty] | -0.091942 | 2.5968 | 3.867 | -3.7583 | 2.7253 | -1.4648 | 0.58319 | -0.14598 | 0.016204 | [12 empty] | 1.2716 | 1.6363 | 1.8489 | -0.24173 | 0.0010043
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 950 | 4.3331 | [20 empty] | -0.091942 | 2.5996 | 3.8752 | -3.7702 | 2.7368 | -1.4725 | 0.58689 | -0.14707 | 0.016341 | [12 empty] | 1.2716 | 1.638 | 1.8528 | -0.24249 | 0.0010086
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 951 | 4.3386 | [20 empty] | -0.091942 | 2.6023 | 3.8833 | -3.7821 | 2.7484 | -1.4803 | 0.5906 | -0.14815 | 0.016479 | [12 empty] | 1.2716 | 1.6397 | 1.8567 | -0.24326 | 0.0010128
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 952 | 4.3441 | [20 empty] | -0.091942 | 2.605 | 3.8915 | -3.7941 | 2.7599 | -1.4881 | 0.59434 | -0.14925 | 0.016619 | [12 empty] | 1.2716 | 1.6415 | 1.8606 | -0.24403 | 0.0010171
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 953 | 4.3497 | [20 empty] | -0.091942 | 2.6078 | 3.8997 | -3.806 | 2.7716 | -1.4959 | 0.59809 | -0.15035 | 0.016759 | [12 empty] | 1.2716 | 1.6432 | 1.8645 | -0.2448 | 0.0010214
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 954 | 4.3552 | [20 empty] | -0.091942 | 2.6105 | 3.9079 | -3.818 | 2.7832 | -1.5038 | 0.60187 | -0.15145 | 0.0169 | [12 empty] | 1.2716 | 1.6449 | 1.8685 | -0.24557 | 0.0010257
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 955 | 4.3607 | [20 empty] | -0.091942 | 2.6132 | 3.9161 | -3.83 | 2.7949 | -1.5117 | 0.60567 | -0.15257 | 0.017042 | [12 empty] | 1.2716 | 1.6466 | 1.8724 | -0.24634 | 0.00103
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 956 | 4.3662 | [20 empty] | -0.091942 | 2.616 | 3.9243 | -3.8421 | 2.8066 | -1.5196 | 0.60948 | -0.15369 | 0.017185 | [12 empty] | 1.2716 | 1.6484 | 1.8763 | -0.24712 | 0.0010343
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 957 | 4.3718 | [20 empty] | -0.091942 | 2.6187 | 3.9325 | -3.8542 | 2.8184 | -1.5276 | 0.61332 | -0.15482 | 0.01733 | [12 empty] | 1.2716 | 1.6501 | 1.8802 | -0.24789 | 0.0010386
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 958 | 4.3773 | [20 empty] | -0.091942 | 2.6215 | 3.9407 | -3.8662 | 2.8302 | -1.5356 | 0.61717 | -0.15596 | 0.017475 | [12 empty] | 1.2716 | 1.6518 | 1.8842 | -0.24867 | 0.001043
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 959 | 4.3828 | [20 empty] | -0.091942 | 2.6242 | 3.9489 | -3.8784 | 2.842 | -1.5436 | 0.62105 | -0.1571 | 0.017622 | [12 empty] | 1.2716 | 1.6535 | 1.8881 | -0.24945 | 0.0010474
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 960 | 4.3884 | [20 empty] | -0.091942 | 2.6269 | 3.9572 | -3.8905 | 2.8539 | -1.5516 | 0.62494 | -0.15825 | 0.017769 | [12 empty] | 1.2716 | 1.6553 | 1.892 | -0.25023 | 0.0010517
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 961 | 4.3939 | [20 empty] | -0.091942 | 2.6297 | 3.9654 | -3.9027 | 2.8658 | -1.5597 | 0.62886 | -0.15941 | 0.017918 | [12 empty] | 1.2716 | 1.657 | 1.896 | -0.25102 | 0.0010561
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 962 | 4.3994 | [20 empty] | -0.091942 | 2.6324 | 3.9737 | -3.9149 | 2.8777 | -1.5679 | 0.63279 | -0.16057 | 0.018068 | [12 empty] | 1.2716 | 1.6587 | 1.8999 | -0.2518 | 0.0010605
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 963 | 4.405 | [20 empty] | -0.091942 | 2.6351 | 3.9819 | -3.9271 | 2.8897 | -1.576 | 0.63675 | -0.16174 | 0.018218 | [12 empty] | 1.2716 | 1.6604 | 1.9039 | -0.25259 | 0.0010649
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 964 | 4.4105 | [20 empty] | -0.091942 | 2.6379 | 3.9902 | -3.9393 | 2.9018 | -1.5842 | 0.64073 | -0.16292 | 0.01837 | [12 empty] | 1.2716 | 1.6622 | 1.9078 | -0.25337 | 0.0010694
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 965 | 4.416 | [20 empty] | -0.091942 | 2.6406 | 3.9985 | -3.9516 | 2.9138 | -1.5925 | 0.64473 | -0.16411 | 0.018523 | [12 empty] | 1.2716 | 1.6639 | 1.9118 | -0.25416 | 0.0010738
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 966 | 4.4216 | [20 empty] | -0.091942 | 2.6433 | 4.0068 | -3.9639 | 2.9259 | -1.6007 | 0.64875 | -0.1653 | 0.018677 | [12 empty] | 1.2716 | 1.6656 | 1.9158 | -0.25495 | 0.0010783
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 967 | 4.4271 | [20 empty] | -0.091942 | 2.6461 | 4.0151 | -3.9762 | 2.938 | -1.609 | 0.65279 | -0.16651 | 0.018833 | [12 empty] | 1.2716 | 1.6673 | 1.9197 | -0.25575 | 0.0010827
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 968 | 4.4326 | [20 empty] | -0.091942 | 2.6488 | 4.0234 | -3.9886 | 2.9502 | -1.6174 | 0.65685 | -0.16771 | 0.018989 | [12 empty] | 1.2716 | 1.6691 | 1.9237 | -0.25654 | 0.0010872
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 969 | 4.4382 | [20 empty] | -0.091942 | 2.6516 | 4.0317 | -4.001 | 2.9624 | -1.6258 | 0.66093 | -0.16893 | 0.019146 | [12 empty] | 1.2716 | 1.6708 | 1.9277 | -0.25734 | 0.0010917
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 970 | 4.4437 | [20 empty] | -0.091942 | 2.6543 | 4.04 | -4.0134 | 2.9747 | -1.6342 | 0.66503 | -0.17016 | 0.019305 | [12 empty] | 1.2716 | 1.6725 | 1.9317 | -0.25813 | 0.0010962
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 971 | 4.4493 | [20 empty] | -0.091942 | 2.657 | 4.0484 | -4.0258 | 2.987 | -1.6426 | 0.66916 | -0.17139 | 0.019465 | [12 empty] | 1.2716 | 1.6742 | 1.9356 | -0.25893 | 0.0011008
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 972 | 4.4548 | [20 empty] | -0.091942 | 2.6598 | 4.0567 | -4.0382 | 2.9993 | -1.6511 | 0.6733 | -0.17263 | 0.019626 | [12 empty] | 1.2716 | 1.676 | 1.9396 | -0.25973 | 0.0011053
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 973 | 4.4604 | [20 empty] | -0.091942 | 2.6625 | 4.0651 | -4.0507 | 3.0116 | -1.6596 | 0.67747 | -0.17387 | 0.019788 | [12 empty] | 1.2716 | 1.6777 | 1.9436 | -0.26054 | 0.0011099
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 974 | 4.4659 | [20 empty] | -0.091942 | 2.6652 | 4.0734 | -4.0632 | 3.024 | -1.6681 | 0.68166 | -0.17513 | 0.019951 | [12 empty] | 1.2716 | 1.6794 | 1.9476 | -0.26134 | 0.0011144
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 975 | 4.4715 | [20 empty] | -0.091942 | 2.668 | 4.0818 | -4.0757 | 3.0365 | -1.6767 | 0.68587 | -0.17639 | 0.020116 | [12 empty] | 1.2716 | 1.6811 | 1.9516 | -0.26215 | 0.001119
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 976 | 4.477 | [20 empty] | -0.091942 | 2.6707 | 4.0902 | -4.0883 | 3.049 | -1.6853 | 0.6901 | -0.17766 | 0.020281 | [12 empty] | 1.2716 | 1.6829 | 1.9556 | -0.26295 | 0.0011236
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 977 | 4.4825 | [20 empty] | -0.091942 | 2.6735 | 4.0986 | -4.1009 | 3.0615 | -1.694 | 0.69435 | -0.17894 | 0.020448 | [12 empty] | 1.2716 | 1.6846 | 1.9596 | -0.26376 | 0.0011282
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 978 | 4.4881 | [20 empty] | -0.091942 | 2.6762 | 4.107 | -4.1135 | 3.074 | -1.7027 | 0.69863 | -0.18023 | 0.020616 | [12 empty] | 1.2716 | 1.6863 | 1.9636 | -0.26457 | 0.0011329
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 979 | 4.4936 | [20 empty] | -0.091942 | 2.6789 | 4.1154 | -4.1261 | 3.0866 | -1.7114 | 0.70292 | -0.18152 | 0.020786 | [12 empty] | 1.2716 | 1.688 | 1.9677 | -0.26539 | 0.0011375
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 980 | 4.4992 | [20 empty] | -0.091942 | 2.6817 | 4.1238 | -4.1388 | 3.0992 | -1.7202 | 0.70724 | -0.18282 | 0.020956 | [12 empty] | 1.2716 | 1.6898 | 1.9717 | -0.2662 | 0.0011421
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 981 | 4.5047 | [20 empty] | -0.091942 | 2.6844 | 4.1322 | -4.1515 | 3.1119 | -1.7289 | 0.71158 | -0.18413 | 0.021128 | [12 empty] | 1.2716 | 1.6915 | 1.9757 | -0.26702 | 0.0011468
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 982 | 4.5103 | [20 empty] | -0.091942 | 2.6871 | 4.1406 | -4.1642 | 3.1246 | -1.7378 | 0.71595 | -0.18545 | 0.021301 | [12 empty] | 1.2716 | 1.6932 | 1.9797 | -0.26783 | 0.0011515
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 983 | 4.5159 | [20 empty] | -0.091942 | 2.6899 | 4.1491 | -4.1769 | 3.1374 | -1.7466 | 0.72033 | -0.18678 | 0.021475 | [12 empty] | 1.2716 | 1.6949 | 1.9838 | -0.26865 | 0.0011562
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 984 | 4.5214 | [20 empty] | -0.091942 | 2.6926 | 4.1575 | -4.1897 | 3.1502 | -1.7555 | 0.72474 | -0.18811 | 0.02165 | [12 empty] | 1.2716 | 1.6966 | 1.9878 | -0.26947 | 0.0011609
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 985 | 4.527 | [20 empty] | -0.091942 | 2.6953 | 4.166 | -4.2024 | 3.163 | -1.7645 | 0.72917 | -0.18945 | 0.021827 | [12 empty] | 1.2716 | 1.6984 | 1.9919 | -0.2703 | 0.0011656
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 986 | 4.5325 | [20 empty] | -0.091942 | 2.6981 | 4.1744 | -4.2153 | 3.1758 | -1.7735 | 0.73362 | -0.1908 | 0.022005 | [12 empty] | 1.2716 | 1.7001 | 1.9959 | -0.27112 | 0.0011704
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 987 | 4.5381 | [20 empty] | -0.091942 | 2.7008 | 4.1829 | -4.2281 | 3.1888 | -1.7825 | 0.7381 | -0.19216 | 0.022184 | [12 empty] | 1.2716 | 1.7018 | 2 | -0.27195 | 0.0011751
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 988 | 4.5436 | [20 empty] | -0.091942 | 2.7036 | 4.1914 | -4.241 | 3.2017 | -1.7915 | 0.7426 | -0.19353 | 0.022364 | [12 empty] | 1.2716 | 1.7035 | 2.004 | -0.27277 | 0.0011799
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 989 | 4.5492 | [20 empty] | -0.091942 | 2.7063 | 4.1999 | -4.2538 | 3.2147 | -1.8006 | 0.74712 | -0.1949 | 0.022546 | [12 empty] | 1.2716 | 1.7053 | 2.0081 | -0.2736 | 0.0011847
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 990 | 4.5548 | [20 empty] | -0.091942 | 2.709 | 4.2084 | -4.2668 | 3.2277 | -1.8097 | 0.75166 | -0.19629 | 0.022729 | [12 empty] | 1.2716 | 1.707 | 2.0121 | -0.27443 | 0.0011895
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 991 | 4.5603 | [20 empty] | -0.091942 | 2.7118 | 4.2169 | -4.2797 | 3.2408 | -1.8189 | 0.75623 | -0.19768 | 0.022913 | [12 empty] | 1.2716 | 1.7087 | 2.0162 | -0.27527 | 0.0011943
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 992 | 4.5659 | [20 empty] | -0.091942 | 2.7145 | 4.2254 | -4.2927 | 3.2539 | -1.8281 | 0.76082 | -0.19908 | 0.023099 | [12 empty] | 1.2716 | 1.7104 | 2.0203 | -0.2761 | 0.0011991
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 993 | 4.5714 | [20 empty] | -0.091942 | 2.7172 | 4.2339 | -4.3057 | 3.267 | -1.8373 | 0.76543 | -0.20049 | 0.023286 | [12 empty] | 1.2716 | 1.7122 | 2.0243 | -0.27694 | 0.001204
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 994 | 4.577 | [20 empty] | -0.091942 | 2.72 | 4.2424 | -4.3187 | 3.2802 | -1.8466 | 0.77007 | -0.20191 | 0.023474 | [12 empty] | 1.2716 | 1.7139 | 2.0284 | -0.27777 | 0.0012088
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 995 | 4.5826 | [20 empty] | -0.091942 | 2.7227 | 4.251 | -4.3317 | 3.2934 | -1.8559 | 0.77473 | -0.20333 | 0.023664 | [12 empty] | 1.2716 | 1.7156 | 2.0325 | -0.27861 | 0.0012137
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 996 | 4.5881 | [20 empty] | -0.091942 | 2.7254 | 4.2595 | -4.3448 | 3.3067 | -1.8652 | 0.77941 | -0.20477 | 0.023855 | [12 empty] | 1.2716 | 1.7173 | 2.0366 | -0.27945 | 0.0012186
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 997 | 4.5937 | [20 empty] | -0.091942 | 2.7282 | 4.2681 | -4.3579 | 3.32 | -1.8746 | 0.78412 | -0.20621 | 0.024047 | [12 empty] | 1.2716 | 1.7191 | 2.0407 | -0.2803 | 0.0012235
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 998 | 4.5993 | [20 empty] | -0.091942 | 2.7309 | 4.2767 | -4.371 | 3.3333 | -1.884 | 0.78885 | -0.20766 | 0.024241 | [12 empty] | 1.2716 | 1.7208 | 2.0448 | -0.28114 | 0.0012284
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 999 | 4.6048 | [20 empty] | -0.091942 | 2.7337 | 4.2852 | -4.3842 | 3.3467 | -1.8935 | 0.79361 | -0.20912 | 0.024436 | [12 empty] | 1.2716 | 1.7225 | 2.0489 | -0.28199 | 0.0012333
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1000 | 4.6104 | [20 empty] | -0.091942 | 2.7364 | 4.2938 | -4.3974 | 3.3601 | -1.903 | 0.79839 | -0.21059 | 0.024632 | [12 empty] | 1.2716 | 1.7242 | 2.053 | -0.28283 | 0.0012383
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1001 | 4.616 | [20 empty] | -0.091942 | 2.7391 | 4.3024 | -4.4106 | 3.3736 | -1.9125 | 0.80319 | -0.21207 | 0.02483 | [12 empty] | 1.2716 | 1.726 | 2.0571 | -0.28368 | 0.0012432
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1002 | 4.6216 | [20 empty] | -0.091942 | 2.7419 | 4.311 | -4.4238 | 3.3871 | -1.9221 | 0.80801 | -0.21356 | 0.025029 | [12 empty] | 1.2716 | 1.7277 | 2.0612 | -0.28453 | 0.0012482
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1003 | 4.6271 | [20 empty] | -0.091942 | 2.7446 | 4.3196 | -4.4371 | 3.4006 | -1.9317 | 0.81286 | -0.21506 | 0.025229 | [12 empty] | 1.2716 | 1.7294 | 2.0653 | -0.28539 | 0.0012532
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1004 | 4.6327 | [20 empty] | -0.091942 | 2.7473 | 4.3282 | -4.4503 | 3.4142 | -1.9414 | 0.81774 | -0.21656 | 0.025431 | [12 empty] | 1.2716 | 1.7311 | 2.0694 | -0.28624 | 0.0012582
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1005 | 4.6383 | [20 empty] | -0.091942 | 2.7501 | 4.3369 | -4.4637 | 3.4278 | -1.951 | 0.82264 | -0.21808 | 0.025635 | [12 empty] | 1.2716 | 1.7329 | 2.0736 | -0.2871 | 0.0012632
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1006 | 4.6439 | [20 empty] | -0.091942 | 2.7528 | 4.3455 | -4.477 | 3.4415 | -1.9608 | 0.82756 | -0.2196 | 0.025839 | [12 empty] | 1.2716 | 1.7346 | 2.0777 | -0.28796 | 0.0012683
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1007 | 4.6494 | [20 empty] | -0.091942 | 2.7555 | 4.3541 | -4.4904 | 3.4552 | -1.9705 | 0.83251 | -0.22113 | 0.026046 | [12 empty] | 1.2716 | 1.7363 | 2.0818 | -0.28881 | 0.0012733
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1008 | 4.655 | [20 empty] | -0.091942 | 2.7583 | 4.3628 | -4.5038 | 3.4689 | -1.9803 | 0.83748 | -0.22267 | 0.026253 | [12 empty] | 1.2716 | 1.738 | 2.086 | -0.28968 | 0.0012784
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1009 | 4.6606 | [20 empty] | -0.091942 | 2.761 | 4.3714 | -4.5172 | 3.4827 | -1.9902 | 0.84248 | -0.22422 | 0.026462 | [12 empty] | 1.2716 | 1.7398 | 2.0901 | -0.29054 | 0.0012835
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1010 | 4.6662 | [20 empty] | -0.091942 | 2.7638 | 4.3801 | -4.5306 | 3.4965 | -2.0001 | 0.8475 | -0.22578 | 0.026673 | [12 empty] | 1.2716 | 1.7415 | 2.0942 | -0.2914 | 0.0012886
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1011 | 4.6718 | [20 empty] | -0.091942 | 2.7665 | 4.3888 | -4.5441 | 3.5104 | -2.01 | 0.85255 | -0.22735 | 0.026885 | [12 empty] | 1.2716 | 1.7432 | 2.0984 | -0.29227 | 0.0012937
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1012 | 4.6773 | [20 empty] | -0.091942 | 2.7692 | 4.3975 | -4.5576 | 3.5243 | -2.0199 | 0.85762 | -0.22893 | 0.027098 | [12 empty] | 1.2716 | 1.7449 | 2.1026 | -0.29314 | 0.0012988
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1013 | 4.6829 | [20 empty] | -0.091942 | 2.772 | 4.4062 | -4.5711 | 3.5383 | -2.0299 | 0.86272 | -0.23052 | 0.027313 | [12 empty] | 1.2716 | 1.7467 | 2.1067 | -0.29401 | 0.0013039
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1014 | 4.6885 | [20 empty] | -0.091942 | 2.7747 | 4.4149 | -4.5847 | 3.5523 | -2.04 | 0.86784 | -0.23212 | 0.02753 | [12 empty] | 1.2716 | 1.7484 | 2.1109 | -0.29488 | 0.0013091
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1015 | 4.6941 | [20 empty] | -0.091942 | 2.7774 | 4.4236 | -4.5982 | 3.5663 | -2.0501 | 0.87299 | -0.23373 | 0.027748 | [12 empty] | 1.2716 | 1.7501 | 2.115 | -0.29575 | 0.0013143
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1016 | 4.6997 | [20 empty] | -0.091942 | 2.7802 | 4.4323 | -4.6118 | 3.5804 | -2.0602 | 0.87816 | -0.23534 | 0.027967 | [12 empty] | 1.2716 | 1.7518 | 2.1192 | -0.29663 | 0.0013195
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1017 | 4.7053 | [20 empty] | -0.091942 | 2.7829 | 4.441 | -4.6255 | 3.5945 | -2.0703 | 0.88336 | -0.23697 | 0.028188 | [12 empty] | 1.2716 | 1.7535 | 2.1234 | -0.2975 | 0.0013247
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1018 | 4.7108 | [20 empty] | -0.091942 | 2.7856 | 4.4498 | -4.6391 | 3.6086 | -2.0805 | 0.88859 | -0.2386 | 0.028411 | [12 empty] | 1.2716 | 1.7553 | 2.1276 | -0.29838 | 0.0013299
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1019 | 4.7164 | [20 empty] | -0.091942 | 2.7884 | 4.4585 | -4.6528 | 3.6228 | -2.0908 | 0.89384 | -0.24025 | 0.028635 | [12 empty] | 1.2716 | 1.757 | 2.1317 | -0.29926 | 0.0013351
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1020 | 4.722 | [20 empty] | -0.091942 | 2.7911 | 4.4673 | -4.6665 | 3.6371 | -2.1011 | 0.89911 | -0.24191 | 0.02886 | [12 empty] | 1.2716 | 1.7587 | 2.1359 | -0.30015 | 0.0013404
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1021 | 4.7276 | [20 empty] | -0.091942 | 2.7939 | 4.476 | -4.6803 | 3.6514 | -2.1114 | 0.90441 | -0.24357 | 0.029087 | [12 empty] | 1.2716 | 1.7604 | 2.1401 | -0.30103 | 0.0013456
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1022 | 4.7332 | [20 empty] | -0.091942 | 2.7966 | 4.4848 | -4.694 | 3.6657 | -2.1217 | 0.90974 | -0.24525 | 0.029316 | [12 empty] | 1.2716 | 1.7622 | 2.1443 | -0.30191 | 0.0013509
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1023 | 4.7388 | [20 empty] | -0.091942 | 2.7993 | 4.4936 | -4.7078 | 3.6801 | -2.1321 | 0.9151 | -0.24693 | 0.029546 | [12 empty] | 1.2716 | 1.7639 | 2.1485 | -0.3028 | 0.0013562
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1024 | 4.7444 | [20 empty] | -0.091942 | 2.8021 | 4.5024 | -4.7216 | 3.6945 | -2.1426 | 0.92048 | -0.24862 | 0.029778 | [12 empty] | 1.2716 | 1.7656 | 2.1527 | -0.30369 | 0.0013615
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1025 | 4.75 | [20 empty] | -0.091942 | 2.8048 | 4.5112 | -4.7355 | 3.7089 | -2.1531 | 0.92588 | -0.25033 | 0.030012 | [12 empty] | 1.2716 | 1.7673 | 2.1569 | -0.30458 | 0.0013668
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1026 | 4.7556 | [20 empty] | -0.091942 | 2.8075 | 4.52 | -4.7494 | 3.7234 | -2.1636 | 0.93132 | -0.25204 | 0.030247 | [12 empty] | 1.2716 | 1.7691 | 2.1611 | -0.30547 | 0.0013722
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1027 | 4.7612 | [20 empty] | -0.091942 | 2.8103 | 4.5288 | -4.7633 | 3.738 | -2.1741 | 0.93677 | -0.25377 | 0.030483 | [12 empty] | 1.2716 | 1.7708 | 2.1653 | -0.30637 | 0.0013775
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1028 | 4.7668 | [20 empty] | -0.091942 | 2.813 | 4.5376 | -4.7772 | 3.7525 | -2.1848 | 0.94226 | -0.2555 | 0.030722 | [12 empty] | 1.2716 | 1.7725 | 2.1696 | -0.30726 | 0.0013829
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1029 | 4.7724 | [20 empty] | -0.091942 | 2.8157 | 4.5465 | -4.7911 | 3.7672 | -2.1954 | 0.94777 | -0.25725 | 0.030961 | [12 empty] | 1.2716 | 1.7742 | 2.1738 | -0.30816 | 0.0013883
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1030 | 4.778 | [20 empty] | -0.091942 | 2.8185 | 4.5553 | -4.8051 | 3.7818 | -2.2061 | 0.95331 | -0.259 | 0.031203 | [12 empty] | 1.2716 | 1.776 | 2.178 | -0.30906 | 0.0013937
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1031 | 4.7836 | [20 empty] | -0.091942 | 2.8212 | 4.5642 | -4.8191 | 3.7965 | -2.2168 | 0.95888 | -0.26077 | 0.031446 | [12 empty] | 1.2716 | 1.7777 | 2.1822 | -0.30996 | 0.0013991
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1032 | 4.7892 | [20 empty] | -0.091942 | 2.824 | 4.573 | -4.8332 | 3.8113 | -2.2276 | 0.96447 | -0.26254 | 0.031691 | [12 empty] | 1.2716 | 1.7794 | 2.1865 | -0.31086 | 0.0014046
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1033 | 4.7948 | [20 empty] | -0.091942 | 2.8267 | 4.5819 | -4.8472 | 3.8261 | -2.2384 | 0.9701 | -0.26433 | 0.031938 | [12 empty] | 1.2716 | 1.7811 | 2.1907 | -0.31177 | 0.00141
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1034 | 4.8004 | [20 empty] | -0.091942 | 2.8294 | 4.5908 | -4.8613 | 3.8409 | -2.2493 | 0.97574 | -0.26613 | 0.032186 | [12 empty] | 1.2716 | 1.7829 | 2.195 | -0.31267 | 0.0014155
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1035 | 4.806 | [20 empty] | -0.091942 | 2.8322 | 4.5996 | -4.8754 | 3.8558 | -2.2602 | 0.98142 | -0.26793 | 0.032436 | [12 empty] | 1.2716 | 1.7846 | 2.1992 | -0.31358 | 0.001421
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1036 | 4.8116 | [20 empty] | -0.091942 | 2.8349 | 4.6085 | -4.8896 | 3.8707 | -2.2711 | 0.98712 | -0.26975 | 0.032687 | [12 empty] | 1.2716 | 1.7863 | 2.2035 | -0.31449 | 0.0014265
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1037 | 4.8172 | [20 empty] | -0.091942 | 2.8376 | 4.6174 | -4.9038 | 3.8857 | -2.2821 | 0.99285 | -0.27158 | 0.03294 | [12 empty] | 1.2716 | 1.788 | 2.2077 | -0.3154 | 0.001432
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1038 | 4.8228 | [20 empty] | -0.091942 | 2.8404 | 4.6263 | -4.918 | 3.9007 | -2.2931 | 0.99861 | -0.27342 | 0.033195 | [12 empty] | 1.2716 | 1.7898 | 2.212 | -0.31632 | 0.0014375
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1039 | 4.8284 | [20 empty] | -0.091942 | 2.8431 | 4.6353 | -4.9322 | 3.9157 | -2.3042 | 1.0044 | -0.27527 | 0.033452 | [12 empty] | 1.2716 | 1.7915 | 2.2162 | -0.31723 | 0.001443
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1040 | 4.834 | [20 empty] | -0.091942 | 2.8458 | 4.6442 | -4.9464 | 3.9308 | -2.3153 | 1.0102 | -0.27713 | 0.033711 | [12 empty] | 1.2716 | 1.7932 | 2.2205 | -0.31815 | 0.0014486
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1041 | 4.8396 | [20 empty] | -0.091942 | 2.8486 | 4.6531 | -4.9607 | 3.946 | -2.3264 | 1.0161 | -0.279 | 0.033971 | [12 empty] | 1.2716 | 1.7949 | 2.2248 | -0.31907 | 0.0014542
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1042 | 4.8452 | [20 empty] | -0.091942 | 2.8513 | 4.6621 | -4.975 | 3.9612 | -2.3376 | 1.0219 | -0.28088 | 0.034233 | [12 empty] | 1.2716 | 1.7967 | 2.2291 | -0.31999 | 0.0014598
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1043 | 4.8508 | [20 empty] | -0.091942 | 2.8541 | 4.671 | -4.9894 | 3.9764 | -2.3489 | 1.0278 | -0.28277 | 0.034496 | [12 empty] | 1.2716 | 1.7984 | 2.2333 | -0.32091 | 0.0014654
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1044 | 4.8564 | [20 empty] | -0.091942 | 2.8568 | 4.68 | -5.0037 | 3.9917 | -2.3601 | 1.0338 | -0.28467 | 0.034762 | [12 empty] | 1.2716 | 1.8001 | 2.2376 | -0.32183 | 0.001471
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1045 | 4.8621 | [20 empty] | -0.091942 | 2.8595 | 4.6889 | -5.0181 | 4.007 | -2.3715 | 1.0397 | -0.28659 | 0.035029 | [12 empty] | 1.2716 | 1.8018 | 2.2419 | -0.32276 | 0.0014767
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1046 | 4.8677 | [20 empty] | -0.091942 | 2.8623 | 4.6979 | -5.0325 | 4.0223 | -2.3828 | 1.0457 | -0.28851 | 0.035298 | [12 empty] | 1.2716 | 1.8035 | 2.2462 | -0.32369 | 0.0014823
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1047 | 4.8733 | [20 empty] | -0.091942 | 2.865 | 4.7069 | -5.047 | 4.0377 | -2.3943 | 1.0517 | -0.29045 | 0.035569 | [12 empty] | 1.2716 | 1.8053 | 2.2505 | -0.32462 | 0.001488
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1048 | 4.8789 | [20 empty] | -0.091942 | 2.8677 | 4.7159 | -5.0615 | 4.0532 | -2.4057 | 1.0577 | -0.2924 | 0.035842 | [12 empty] | 1.2716 | 1.807 | 2.2548 | -0.32555 | 0.0014937
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1049 | 4.8845 | [20 empty] | -0.091942 | 2.8705 | 4.7249 | -5.076 | 4.0687 | -2.4172 | 1.0638 | -0.29436 | 0.036116 | [12 empty] | 1.2716 | 1.8087 | 2.2591 | -0.32648 | 0.0014994
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1050 | 4.8901 | [20 empty] | -0.091942 | 2.8732 | 4.7339 | -5.0905 | 4.0842 | -2.4288 | 1.0699 | -0.29633 | 0.036393 | [12 empty] | 1.2716 | 1.8104 | 2.2634 | -0.32742 | 0.0015051
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1051 | 4.8958 | [20 empty] | -0.091942 | 2.8759 | 4.7429 | -5.1051 | 4.0998 | -2.4403 | 1.076 | -0.29831 | 0.036671 | [12 empty] | 1.2716 | 1.8122 | 2.2677 | -0.32835 | 0.0015109
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1052 | 4.9014 | [20 empty] | -0.091942 | 2.8787 | 4.752 | -5.1196 | 4.1154 | -2.452 | 1.0822 | -0.3003 | 0.036951 | [12 empty] | 1.2716 | 1.8139 | 2.272 | -0.32929 | 0.0015166
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1053 | 4.907 | [20 empty] | -0.091942 | 2.8814 | 4.761 | -5.1343 | 4.1311 | -2.4636 | 1.0884 | -0.3023 | 0.037233 | [12 empty] | 1.2716 | 1.8156 | 2.2764 | -0.33023 | 0.0015224
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1054 | 4.9126 | [20 empty] | -0.091942 | 2.8842 | 4.7701 | -5.1489 | 4.1468 | -2.4754 | 1.0946 | -0.30432 | 0.037517 | [12 empty] | 1.2716 | 1.8173 | 2.2807 | -0.33117 | 0.0015282
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1055 | 4.9182 | [20 empty] | -0.091942 | 2.8869 | 4.7791 | -5.1636 | 4.1626 | -2.4871 | 1.1008 | -0.30635 | 0.037802 | [12 empty] | 1.2716 | 1.8191 | 2.285 | -0.33212 | 0.001534
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1056 | 4.9239 | [20 empty] | -0.091942 | 2.8896 | 4.7882 | -5.1783 | 4.1784 | -2.4989 | 1.1071 | -0.30838 | 0.03809 | [12 empty] | 1.2716 | 1.8208 | 2.2894 | -0.33306 | 0.0015398
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1057 | 4.9295 | [20 empty] | -0.091942 | 2.8924 | 4.7973 | -5.193 | 4.1942 | -2.5108 | 1.1134 | -0.31043 | 0.038379 | [12 empty] | 1.2716 | 1.8225 | 2.2937 | -0.33401 | 0.0015457
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1058 | 4.9351 | [20 empty] | -0.091942 | 2.8951 | 4.8063 | -5.2077 | 4.2101 | -2.5227 | 1.1198 | -0.3125 | 0.038671 | [12 empty] | 1.2716 | 1.8242 | 2.298 | -0.33496 | 0.0015515
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1059 | 4.9407 | [20 empty] | -0.091942 | 2.8978 | 4.8154 | -5.2225 | 4.2261 | -2.5346 | 1.1261 | -0.31457 | 0.038964 | [12 empty] | 1.2716 | 1.826 | 2.3024 | -0.33591 | 0.0015574
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1060 | 4.9464 | [20 empty] | -0.091942 | 2.9006 | 4.8245 | -5.2373 | 4.242 | -2.5466 | 1.1325 | -0.31665 | 0.03926 | [12 empty] | 1.2716 | 1.8277 | 2.3067 | -0.33686 | 0.0015633
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1061 | 4.952 | [20 empty] | -0.091942 | 2.9033 | 4.8336 | -5.2522 | 4.2581 | -2.5587 | 1.139 | -0.31875 | 0.039557 | [12 empty] | 1.2716 | 1.8294 | 2.3111 | -0.33781 | 0.0015692
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1062 | 4.9576 | [20 empty] | -0.091942 | 2.906 | 4.8427 | -5.267 | 4.2742 | -2.5707 | 1.1454 | -0.32086 | 0.039856 | [12 empty] | 1.2716 | 1.8311 | 2.3154 | -0.33877 | 0.0015751
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1063 | 4.9632 | [20 empty] | -0.091942 | 2.9088 | 4.8519 | -5.2819 | 4.2903 | -2.5829 | 1.1519 | -0.32298 | 0.040157 | [12 empty] | 1.2716 | 1.8329 | 2.3198 | -0.33973 | 0.0015811
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1064 | 4.9689 | [20 empty] | -0.091942 | 2.9115 | 4.861 | -5.2969 | 4.3064 | -2.595 | 1.1584 | -0.32511 | 0.040461 | [12 empty] | 1.2716 | 1.8346 | 2.3242 | -0.34069 | 0.001587
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1065 | 4.9745 | [20 empty] | -0.091942 | 2.9143 | 4.8701 | -5.3118 | 4.3227 | -2.6073 | 1.165 | -0.32726 | 0.040766 | [12 empty] | 1.2716 | 1.8363 | 2.3285 | -0.34165 | 0.001593
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1066 | 4.9801 | [20 empty] | -0.091942 | 2.917 | 4.8793 | -5.3268 | 4.3389 | -2.6195 | 1.1715 | -0.32942 | 0.041073 | [12 empty] | 1.2716 | 1.838 | 2.3329 | -0.34261 | 0.001599
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1067 | 4.9858 | [20 empty] | -0.091942 | 2.9197 | 4.8885 | -5.3418 | 4.3552 | -2.6318 | 1.1781 | -0.33159 | 0.041382 | [12 empty] | 1.2716 | 1.8398 | 2.3373 | -0.34358 | 0.001605
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1068 | 4.9914 | [20 empty] | -0.091942 | 2.9225 | 4.8976 | -5.3568 | 4.3716 | -2.6442 | 1.1848 | -0.33377 | 0.041694 | [12 empty] | 1.2716 | 1.8415 | 2.3417 | -0.34454 | 0.001611
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1069 | 4.997 | [20 empty] | -0.091942 | 2.9252 | 4.9068 | -5.3719 | 4.388 | -2.6566 | 1.1915 | -0.33596 | 0.042007 | [12 empty] | 1.2716 | 1.8432 | 2.3461 | -0.34551 | 0.0016171
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1070 | 5.0027 | [20 empty] | -0.091942 | 2.9279 | 4.916 | -5.387 | 4.4044 | -2.669 | 1.1982 | -0.33817 | 0.042322 | [12 empty] | 1.2716 | 1.8449 | 2.3505 | -0.34648 | 0.0016231
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1071 | 5.0083 | [20 empty] | -0.091942 | 2.9307 | 4.9252 | -5.4021 | 4.4209 | -2.6815 | 1.2049 | -0.34039 | 0.04264 | [12 empty] | 1.2716 | 1.8467 | 2.3549 | -0.34746 | 0.0016292
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1072 | 5.014 | [20 empty] | -0.091942 | 2.9334 | 4.9344 | -5.4172 | 4.4374 | -2.6941 | 1.2117 | -0.34262 | 0.042959 | [12 empty] | 1.2716 | 1.8484 | 2.3593 | -0.34843 | 0.0016353
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1073 | 5.0196 | [20 empty] | -0.091942 | 2.9361 | 4.9436 | -5.4324 | 4.454 | -2.7067 | 1.2185 | -0.34486 | 0.043281 | [12 empty] | 1.2716 | 1.8501 | 2.3637 | -0.34941 | 0.0016414
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1074 | 5.0252 | [20 empty] | -0.091942 | 2.9389 | 4.9528 | -5.4476 | 4.4706 | -2.7193 | 1.2253 | -0.34712 | 0.043605 | [12 empty] | 1.2716 | 1.8518 | 2.3681 | -0.35038 | 0.0016475
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1075 | 5.0309 | [20 empty] | -0.091942 | 2.9416 | 4.962 | -5.4628 | 4.4873 | -2.732 | 1.2321 | -0.34938 | 0.043931 | [12 empty] | 1.2716 | 1.8536 | 2.3725 | -0.35136 | 0.0016537
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1076 | 5.0365 | [20 empty] | -0.091942 | 2.9444 | 4.9713 | -5.4781 | 4.504 | -2.7447 | 1.239 | -0.35167 | 0.044259 | [12 empty] | 1.2716 | 1.8553 | 2.3769 | -0.35235 | 0.0016598
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1077 | 5.0422 | [20 empty] | -0.091942 | 2.9471 | 4.9805 | -5.4934 | 4.5208 | -2.7575 | 1.246 | -0.35396 | 0.044589 | [12 empty] | 1.2716 | 1.857 | 2.3813 | -0.35333 | 0.001666
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1078 | 5.0478 | [20 empty] | -0.091942 | 2.9498 | 4.9898 | -5.5087 | 4.5376 | -2.7703 | 1.2529 | -0.35627 | 0.044921 | [12 empty] | 1.2716 | 1.8587 | 2.3857 | -0.35431 | 0.0016722
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1079 | 5.0535 | [20 empty] | -0.091942 | 2.9526 | 4.999 | -5.524 | 4.5545 | -2.7832 | 1.2599 | -0.35859 | 0.045255 | [12 empty] | 1.2716 | 1.8604 | 2.3902 | -0.3553 | 0.0016784
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1080 | 5.0591 | [20 empty] | -0.091942 | 2.9553 | 5.0083 | -5.5394 | 4.5714 | -2.7961 | 1.2669 | -0.36092 | 0.045592 | [12 empty] | 1.2716 | 1.8622 | 2.3946 | -0.35629 | 0.0016847
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1081 | 5.0648 | [20 empty] | -0.091942 | 2.958 | 5.0176 | -5.5548 | 4.5883 | -2.8091 | 1.274 | -0.36327 | 0.045931 | [12 empty] | 1.2716 | 1.8639 | 2.399 | -0.35728 | 0.0016909
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1082 | 5.0704 | [20 empty] | -0.091942 | 2.9608 | 5.0269 | -5.5703 | 4.6053 | -2.8221 | 1.2811 | -0.36562 | 0.046272 | [12 empty] | 1.2716 | 1.8656 | 2.4035 | -0.35827 | 0.0016972
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1083 | 5.076 | [20 empty] | -0.091942 | 2.9635 | 5.0362 | -5.5857 | 4.6224 | -2.8352 | 1.2882 | -0.368 | 0.046615 | [12 empty] | 1.2716 | 1.8673 | 2.4079 | -0.35927 | 0.0017035
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1084 | 5.0817 | [20 empty] | -0.091942 | 2.9662 | 5.0455 | -5.6012 | 4.6395 | -2.8483 | 1.2954 | -0.37038 | 0.046961 | [12 empty] | 1.2716 | 1.8691 | 2.4124 | -0.36026 | 0.0017098
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1085 | 5.0874 | [20 empty] | -0.091942 | 2.969 | 5.0548 | -5.6167 | 4.6566 | -2.8614 | 1.3025 | -0.37278 | 0.047308 | [12 empty] | 1.2716 | 1.8708 | 2.4168 | -0.36126 | 0.0017161
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1086 | 5.093 | [20 empty] | -0.091942 | 2.9717 | 5.0641 | -5.6323 | 4.6738 | -2.8747 | 1.3098 | -0.37519 | 0.047658 | [12 empty] | 1.2716 | 1.8725 | 2.4213 | -0.36226 | 0.0017224
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1087 | 5.0987 | [20 empty] | -0.091942 | 2.9745 | 5.0734 | -5.6478 | 4.6911 | -2.8879 | 1.317 | -0.37762 | 0.04801 | [12 empty] | 1.2716 | 1.8742 | 2.4257 | -0.36326 | 0.0017288
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1088 | 5.1043 | [20 empty] | -0.091942 | 2.9772 | 5.0828 | -5.6634 | 4.7083 | -2.9012 | 1.3243 | -0.38006 | 0.048365 | [12 empty] | 1.2716 | 1.876 | 2.4302 | -0.36427 | 0.0017351
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1089 | 5.11 | [20 empty] | -0.091942 | 2.9799 | 5.0921 | -5.6791 | 4.7257 | -2.9146 | 1.3316 | -0.38251 | 0.048722 | [12 empty] | 1.2716 | 1.8777 | 2.4347 | -0.36527 | 0.0017415
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1090 | 5.1156 | [20 empty] | -0.091942 | 2.9827 | 5.1015 | -5.6947 | 4.7431 | -2.928 | 1.339 | -0.38497 | 0.049081 | [12 empty] | 1.2716 | 1.8794 | 2.4392 | -0.36628 | 0.0017479
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1091 | 5.1213 | [20 empty] | -0.091942 | 2.9854 | 5.1108 | -5.7104 | 4.7605 | -2.9414 | 1.3464 | -0.38745 | 0.049442 | [12 empty] | 1.2716 | 1.8811 | 2.4436 | -0.36729 | 0.0017544
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1092 | 5.1269 | [20 empty] | -0.091942 | 2.9881 | 5.1202 | -5.7261 | 4.778 | -2.955 | 1.3538 | -0.38994 | 0.049806 | [12 empty] | 1.2716 | 1.8829 | 2.4481 | -0.3683 | 0.0017608
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1093 | 5.1326 | [20 empty] | -0.091942 | 2.9909 | 5.1296 | -5.7419 | 4.7955 | -2.9685 | 1.3612 | -0.39245 | 0.050172 | [12 empty] | 1.2716 | 1.8846 | 2.4526 | -0.36931 | 0.0017673
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1094 | 5.1383 | [20 empty] | -0.091942 | 2.9936 | 5.139 | -5.7576 | 4.8131 | -2.9821 | 1.3687 | -0.39497 | 0.05054 | [12 empty] | 1.2716 | 1.8863 | 2.4571 | -0.37033 | 0.0017737
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1095 | 5.1439 | [20 empty] | -0.091942 | 2.9963 | 5.1484 | -5.7734 | 4.8307 | -2.9958 | 1.3762 | -0.39751 | 0.050911 | [12 empty] | 1.2716 | 1.888 | 2.4616 | -0.37134 | 0.0017802
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1096 | 5.1496 | [20 empty] | -0.091942 | 2.9991 | 5.1578 | -5.7893 | 4.8484 | -3.0095 | 1.3838 | -0.40005 | 0.051284 | [12 empty] | 1.2716 | 1.8898 | 2.4661 | -0.37236 | 0.0017867
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1097 | 5.1552 | [20 empty] | -0.091942 | 3.0018 | 5.1672 | -5.8051 | 4.8661 | -3.0232 | 1.3914 | -0.40262 | 0.05166 | [12 empty] | 1.2716 | 1.8915 | 2.4706 | -0.37338 | 0.0017933
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1098 | 5.1609 | [20 empty] | -0.091942 | 3.0046 | 5.1766 | -5.821 | 4.8838 | -3.037 | 1.399 | -0.40519 | 0.052038 | [12 empty] | 1.2716 | 1.8932 | 2.4751 | -0.3744 | 0.0017998
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1099 | 5.1666 | [20 empty] | -0.091942 | 3.0073 | 5.1861 | -5.837 | 4.9017 | -3.0509 | 1.4067 | -0.40778 | 0.052418 | [12 empty] | 1.2716 | 1.8949 | 2.4796 | -0.37543 | 0.0018064
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1100 | 5.1722 | [20 empty] | -0.091942 | 3.01 | 5.1955 | -5.8529 | 4.9195 | -3.0648 | 1.4144 | -0.41039 | 0.052801 | [12 empty] | 1.2716 | 1.8967 | 2.4841 | -0.37645 | 0.001813
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1101 | 5.1779 | [20 empty] | -0.091942 | 3.0128 | 5.205 | -5.8689 | 4.9374 | -3.0787 | 1.4221 | -0.41301 | 0.053186 | [12 empty] | 1.2716 | 1.8984 | 2.4886 | -0.37748 | 0.0018196
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1102 | 5.1836 | [20 empty] | -0.091942 | 3.0155 | 5.2144 | -5.8849 | 4.9554 | -3.0928 | 1.4299 | -0.41564 | 0.053574 | [12 empty] | 1.2716 | 1.9001 | 2.4932 | -0.37851 | 0.0018262
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1103 | 5.1892 | [20 empty] | -0.091942 | 3.0182 | 5.2239 | -5.9009 | 4.9734 | -3.1068 | 1.4377 | -0.41829 | 0.053964 | [12 empty] | 1.2716 | 1.9018 | 2.4977 | -0.37954 | 0.0018328
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1104 | 5.1949 | [20 empty] | -0.091942 | 3.021 | 5.2334 | -5.917 | 4.9915 | -3.1209 | 1.4455 | -0.42095 | 0.054357 | [12 empty] | 1.2716 | 1.9036 | 2.5022 | -0.38057 | 0.0018395
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1105 | 5.2006 | [20 empty] | -0.091942 | 3.0237 | 5.2429 | -5.9331 | 5.0096 | -3.1351 | 1.4534 | -0.42362 | 0.054752 | [12 empty] | 1.2716 | 1.9053 | 2.5067 | -0.38161 | 0.0018462
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1106 | 5.2062 | [20 empty] | -0.091942 | 3.0264 | 5.2523 | -5.9492 | 5.0277 | -3.1493 | 1.4613 | -0.42631 | 0.055149 | [12 empty] | 1.2716 | 1.907 | 2.5113 | -0.38265 | 0.0018528
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1107 | 5.2119 | [20 empty] | -0.091942 | 3.0292 | 5.2618 | -5.9653 | 5.0459 | -3.1636 | 1.4693 | -0.42902 | 0.055549 | [12 empty] | 1.2716 | 1.9087 | 2.5158 | -0.38368 | 0.0018596
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1108 | 5.2176 | [20 empty] | -0.091942 | 3.0319 | 5.2714 | -5.9815 | 5.0642 | -3.1779 | 1.4772 | -0.43174 | 0.055952 | [12 empty] | 1.2716 | 1.9105 | 2.5204 | -0.38473 | 0.0018663
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1109 | 5.2233 | [20 empty] | -0.091942 | 3.0347 | 5.2809 | -5.9977 | 5.0825 | -3.1922 | 1.4853 | -0.43448 | 0.056357 | [12 empty] | 1.2716 | 1.9122 | 2.5249 | -0.38577 | 0.001873
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1110 | 5.2289 | [20 empty] | -0.091942 | 3.0374 | 5.2904 | -6.014 | 5.1009 | -3.2067 | 1.4933 | -0.43723 | 0.056765 | [12 empty] | 1.2716 | 1.9139 | 2.5295 | -0.38681 | 0.0018798
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1111 | 5.2346 | [20 empty] | -0.091942 | 3.0401 | 5.2999 | -6.0302 | 5.1193 | -3.2211 | 1.5014 | -0.43999 | 0.057176 | [12 empty] | 1.2716 | 1.9156 | 2.534 | -0.38786 | 0.0018866
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1112 | 5.2403 | [20 empty] | -0.091942 | 3.0429 | 5.3095 | -6.0465 | 5.1377 | -3.2356 | 1.5095 | -0.44277 | 0.057589 | [12 empty] | 1.2716 | 1.9173 | 2.5386 | -0.38891 | 0.0018934
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1113 | 5.246 | [20 empty] | -0.091942 | 3.0456 | 5.319 | -6.0629 | 5.1562 | -3.2502 | 1.5177 | -0.44556 | 0.058004 | [12 empty] | 1.2716 | 1.9191 | 2.5432 | -0.38996 | 0.0019002
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1114 | 5.2516 | [20 empty] | -0.091942 | 3.0483 | 5.3286 | -6.0792 | 5.1748 | -3.2648 | 1.5259 | -0.44837 | 0.058423 | [12 empty] | 1.2716 | 1.9208 | 2.5477 | -0.39101 | 0.001907
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1115 | 5.2573 | [20 empty] | -0.091942 | 3.0511 | 5.3382 | -6.0956 | 5.1934 | -3.2795 | 1.5341 | -0.4512 | 0.058843 | [12 empty] | 1.2716 | 1.9225 | 2.5523 | -0.39206 | 0.0019139
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1116 | 5.263 | [20 empty] | -0.091942 | 3.0538 | 5.3478 | -6.112 | 5.2121 | -3.2943 | 1.5424 | -0.45404 | 0.059267 | [12 empty] | 1.2716 | 1.9242 | 2.5569 | -0.39312 | 0.0019208
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1117 | 5.2687 | [20 empty] | -0.091942 | 3.0565 | 5.3573 | -6.1285 | 5.2308 | -3.309 | 1.5507 | -0.4569 | 0.059693 | [12 empty] | 1.2716 | 1.926 | 2.5615 | -0.39418 | 0.0019277
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1118 | 5.2744 | [20 empty] | -0.091942 | 3.0593 | 5.3669 | -6.1449 | 5.2495 | -3.3239 | 1.5591 | -0.45977 | 0.060122 | [12 empty] | 1.2716 | 1.9277 | 2.5661 | -0.39524 | 0.0019346
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1119 | 5.2801 | [20 empty] | -0.091942 | 3.062 | 5.3765 | -6.1615 | 5.2683 | -3.3388 | 1.5674 | -0.46265 | 0.060554 | [12 empty] | 1.2716 | 1.9294 | 2.5707 | -0.3963 | 0.0019415
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1120 | 5.2857 | [20 empty] | -0.091942 | 3.0648 | 5.3862 | -6.178 | 5.2872 | -3.3537 | 1.5759 | -0.46555 | 0.060988 | [12 empty] | 1.2716 | 1.9311 | 2.5753 | -0.39736 | 0.0019485
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1121 | 5.2914 | [20 empty] | -0.091942 | 3.0675 | 5.3958 | -6.1945 | 5.3061 | -3.3687 | 1.5843 | -0.46847 | 0.061425 | [12 empty] | 1.2716 | 1.9329 | 2.5799 | -0.39843 | 0.0019554
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1122 | 5.2971 | [20 empty] | -0.091942 | 3.0702 | 5.4054 | -6.2111 | 5.325 | -3.3838 | 1.5928 | -0.47141 | 0.061865 | [12 empty] | 1.2716 | 1.9346 | 2.5845 | -0.39949 | 0.0019624
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1123 | 5.3028 | [20 empty] | -0.091942 | 3.073 | 5.415 | -6.2278 | 5.3441 | -3.3989 | 1.6014 | -0.47435 | 0.062307 | [12 empty] | 1.2716 | 1.9363 | 2.5891 | -0.40056 | 0.0019694
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1124 | 5.3085 | [20 empty] | -0.091942 | 3.0757 | 5.4247 | -6.2444 | 5.3631 | -3.414 | 1.6099 | -0.47732 | 0.062752 | [12 empty] | 1.2716 | 1.938 | 2.5937 | -0.40163 | 0.0019764
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1125 | 5.3142 | [20 empty] | -0.091942 | 3.0784 | 5.4344 | -6.2611 | 5.3822 | -3.4293 | 1.6186 | -0.4803 | 0.0632 | [12 empty] | 1.2716 | 1.9398 | 2.5983 | -0.40271 | 0.0019835
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1126 | 5.3199 | [20 empty] | -0.091942 | 3.0812 | 5.444 | -6.2778 | 5.4014 | -3.4445 | 1.6272 | -0.4833 | 0.063651 | [12 empty] | 1.2716 | 1.9415 | 2.6029 | -0.40378 | 0.0019905
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1127 | 5.3256 | [20 empty] | -0.091942 | 3.0839 | 5.4537 | -6.2945 | 5.4206 | -3.4598 | 1.6359 | -0.48631 | 0.064105 | [12 empty] | 1.2716 | 1.9432 | 2.6076 | -0.40486 | 0.0019976
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1128 | 5.3313 | [20 empty] | -0.091942 | 3.0866 | 5.4634 | -6.3113 | 5.4399 | -3.4752 | 1.6446 | -0.48934 | 0.064561 | [12 empty] | 1.2716 | 1.9449 | 2.6122 | -0.40594 | 0.0020047
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1129 | 5.3369 | [20 empty] | -0.091942 | 3.0894 | 5.4731 | -6.3281 | 5.4592 | -3.4907 | 1.6534 | -0.49238 | 0.065021 | [12 empty] | 1.2716 | 1.9467 | 2.6168 | -0.40702 | 0.0020118
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1130 | 5.3426 | [20 empty] | -0.091942 | 3.0921 | 5.4828 | -6.345 | 5.4786 | -3.5061 | 1.6622 | -0.49544 | 0.065483 | [12 empty] | 1.2716 | 1.9484 | 2.6215 | -0.4081 | 0.002019
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1131 | 5.3483 | [20 empty] | -0.091942 | 3.0949 | 5.4925 | -6.3618 | 5.498 | -3.5217 | 1.671 | -0.49852 | 0.065948 | [12 empty] | 1.2716 | 1.9501 | 2.6261 | -0.40918 | 0.0020261
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1132 | 5.354 | [20 empty] | -0.091942 | 3.0976 | 5.5022 | -6.3787 | 5.5174 | -3.5373 | 1.6799 | -0.50161 | 0.066416 | [12 empty] | 1.2716 | 1.9518 | 2.6307 | -0.41027 | 0.0020333
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1133 | 5.3597 | [20 empty] | -0.091942 | 3.1003 | 5.5119 | -6.3956 | 5.537 | -3.5529 | 1.6889 | -0.50472 | 0.066886 | [12 empty] | 1.2716 | 1.9536 | 2.6354 | -0.41136 | 0.0020405
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1134 | 5.3654 | [20 empty] | -0.091942 | 3.1031 | 5.5217 | -6.4126 | 5.5565 | -3.5686 | 1.6978 | -0.50785 | 0.06736 | [12 empty] | 1.2716 | 1.9553 | 2.64 | -0.41245 | 0.0020477
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1135 | 5.3711 | [20 empty] | -0.091942 | 3.1058 | 5.5314 | -6.4295 | 5.5762 | -3.5844 | 1.7068 | -0.51099 | 0.067837 | [12 empty] | 1.2716 | 1.957 | 2.6447 | -0.41354 | 0.002055
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1136 | 5.3768 | [20 empty] | -0.091942 | 3.1085 | 5.5411 | -6.4466 | 5.5958 | -3.6002 | 1.7159 | -0.51415 | 0.068317 | [12 empty] | 1.2716 | 1.9587 | 2.6494 | -0.41464 | 0.0020622
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1137 | 5.3825 | [20 empty] | -0.091942 | 3.1113 | 5.5509 | -6.4636 | 5.6156 | -3.6161 | 1.7249 | -0.51733 | 0.068799 | [12 empty] | 1.2716 | 1.9605 | 2.654 | -0.41573 | 0.0020695
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1138 | 5.3882 | [20 empty] | -0.091942 | 3.114 | 5.5607 | -6.4807 | 5.6354 | -3.632 | 1.7341 | -0.52052 | 0.069285 | [12 empty] | 1.2716 | 1.9622 | 2.6587 | -0.41683 | 0.0020768
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1139 | 5.3939 | [20 empty] | -0.091942 | 3.1167 | 5.5705 | -6.4978 | 5.6552 | -3.648 | 1.7432 | -0.52373 | 0.069773 | [12 empty] | 1.2716 | 1.9639 | 2.6634 | -0.41793 | 0.0020841
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1140 | 5.3996 | [20 empty] | -0.091942 | 3.1195 | 5.5802 | -6.5149 | 5.6751 | -3.664 | 1.7524 | -0.52696 | 0.070265 | [12 empty] | 1.2716 | 1.9656 | 2.6681 | -0.41903 | 0.0020914
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1141 | 5.4053 | [20 empty] | -0.091942 | 3.1222 | 5.59 | -6.5321 | 5.695 | -3.6801 | 1.7617 | -0.53021 | 0.070759 | [12 empty] | 1.2716 | 1.9674 | 2.6727 | -0.42013 | 0.0020988
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1142 | 5.4111 | [20 empty] | -0.091942 | 3.125 | 5.5998 | -6.5492 | 5.715 | -3.6963 | 1.771 | -0.53347 | 0.071257 | [12 empty] | 1.2716 | 1.9691 | 2.6774 | -0.42124 | 0.0021061
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1143 | 5.4168 | [20 empty] | -0.091942 | 3.1277 | 5.6096 | -6.5665 | 5.735 | -3.7125 | 1.7803 | -0.53674 | 0.071758 | [12 empty] | 1.2716 | 1.9708 | 2.6821 | -0.42235 | 0.0021135
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1144 | 5.4225 | [20 empty] | -0.091942 | 3.1304 | 5.6195 | -6.5837 | 5.7551 | -3.7288 | 1.7897 | -0.54004 | 0.072262 | [12 empty] | 1.2716 | 1.9725 | 2.6868 | -0.42346 | 0.0021209
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1145 | 5.4282 | [20 empty] | -0.091942 | 3.1332 | 5.6293 | -6.601 | 5.7753 | -3.7451 | 1.7991 | -0.54335 | 0.072768 | [12 empty] | 1.2716 | 1.9742 | 2.6915 | -0.42457 | 0.0021283
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1146 | 5.4339 | [20 empty] | -0.091942 | 3.1359 | 5.6391 | -6.6183 | 5.7955 | -3.7615 | 1.8085 | -0.54668 | 0.073278 | [12 empty] | 1.2716 | 1.976 | 2.6962 | -0.42568 | 0.0021358
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1147 | 5.4396 | [20 empty] | -0.091942 | 3.1386 | 5.649 | -6.6356 | 5.8157 | -3.7779 | 1.818 | -0.55003 | 0.073791 | [12 empty] | 1.2716 | 1.9777 | 2.7009 | -0.4268 | 0.0021432
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1148 | 5.4453 | [20 empty] | -0.091942 | 3.1414 | 5.6588 | -6.653 | 5.8361 | -3.7944 | 1.8275 | -0.5534 | 0.074308 | [12 empty] | 1.2716 | 1.9794 | 2.7056 | -0.42791 | 0.0021507
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1149 | 5.451 | [20 empty] | -0.091942 | 3.1441 | 5.6687 | -6.6704 | 5.8564 | -3.811 | 1.8371 | -0.55678 | 0.074827 | [12 empty] | 1.2716 | 1.9811 | 2.7104 | -0.42903 | 0.0021582
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1150 | 5.4567 | [20 empty] | -0.091942 | 3.1468 | 5.6786 | -6.6878 | 5.8768 | -3.8276 | 1.8467 | -0.56018 | 0.07535 | [12 empty] | 1.2716 | 1.9829 | 2.7151 | -0.43015 | 0.0021658
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1151 | 5.4625 | [20 empty] | -0.091942 | 3.1496 | 5.6884 | -6.7053 | 5.8973 | -3.8443 | 1.8564 | -0.5636 | 0.075875 | [12 empty] | 1.2716 | 1.9846 | 2.7198 | -0.43128 | 0.0021733
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1152 | 5.4682 | [20 empty] | -0.091942 | 3.1523 | 5.6983 | -6.7228 | 5.9178 | -3.861 | 1.8661 | -0.56704 | 0.076404 | [12 empty] | 1.2716 | 1.9863 | 2.7245 | -0.4324 | 0.0021809
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1153 | 5.4739 | [20 empty] | -0.091942 | 3.1551 | 5.7082 | -6.7403 | 5.9384 | -3.8778 | 1.8758 | -0.57049 | 0.076937 | [12 empty] | 1.2716 | 1.988 | 2.7293 | -0.43353 | 0.0021884
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1154 | 5.4796 | [20 empty] | -0.091942 | 3.1578 | 5.7181 | -6.7579 | 5.959 | -3.8946 | 1.8856 | -0.57396 | 0.077472 | [12 empty] | 1.2716 | 1.9898 | 2.734 | -0.43466 | 0.002196
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1155 | 5.4853 | [20 empty] | -0.091942 | 3.1605 | 5.7281 | -6.7755 | 5.9797 | -3.9115 | 1.8954 | -0.57746 | 0.078011 | [12 empty] | 1.2716 | 1.9915 | 2.7387 | -0.43579 | 0.0022037
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1156 | 5.491 | [20 empty] | -0.091942 | 3.1633 | 5.738 | -6.7931 | 6.0004 | -3.9285 | 1.9053 | -0.58096 | 0.078553 | [12 empty] | 1.2716 | 1.9932 | 2.7435 | -0.43692 | 0.0022113
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1157 | 5.4968 | [20 empty] | -0.091942 | 3.166 | 5.7479 | -6.8107 | 6.0212 | -3.9455 | 1.9152 | -0.58449 | 0.079098 | [12 empty] | 1.2716 | 1.9949 | 2.7482 | -0.43806 | 0.002219
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1158 | 5.5025 | [20 empty] | -0.091942 | 3.1687 | 5.7578 | -6.8284 | 6.0421 | -3.9626 | 1.9251 | -0.58804 | 0.079647 | [12 empty] | 1.2716 | 1.9967 | 2.753 | -0.43919 | 0.0022267
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1159 | 5.5082 | [20 empty] | -0.091942 | 3.1715 | 5.7678 | -6.8461 | 6.063 | -3.9797 | 1.9351 | -0.5916 | 0.080199 | [12 empty] | 1.2716 | 1.9984 | 2.7577 | -0.44033 | 0.0022344
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1160 | 5.5139 | [20 empty] | -0.091942 | 3.1742 | 5.7778 | -6.8638 | 6.0839 | -3.9969 | 1.9452 | -0.59518 | 0.080754 | [12 empty] | 1.2716 | 2.0001 | 2.7625 | -0.44147 | 0.0022421
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1161 | 5.5197 | [20 empty] | -0.091942 | 3.1769 | 5.7877 | -6.8816 | 6.1049 | -4.0142 | 1.9553 | -0.59878 | 0.081312 | [12 empty] | 1.2716 | 2.0018 | 2.7673 | -0.44262 | 0.0022498
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1162 | 5.5254 | [20 empty] | -0.091942 | 3.1797 | 5.7977 | -6.8994 | 6.126 | -4.0315 | 1.9654 | -0.6024 | 0.081874 | [12 empty] | 1.2716 | 2.0036 | 2.772 | -0.44376 | 0.0022576
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1163 | 5.5311 | [20 empty] | -0.091942 | 3.1824 | 5.8077 | -6.9172 | 6.1471 | -4.0489 | 1.9756 | -0.60604 | 0.08244 | [12 empty] | 1.2716 | 2.0053 | 2.7768 | -0.44491 | 0.0022654
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1164 | 5.5368 | [20 empty] | -0.091942 | 3.1852 | 5.8177 | -6.9351 | 6.1683 | -4.0663 | 1.9858 | -0.6097 | 0.083009 | [12 empty] | 1.2716 | 2.007 | 2.7816 | -0.44606 | 0.0022732
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1165 | 5.5426 | [20 empty] | -0.091942 | 3.1879 | 5.8277 | -6.953 | 6.1895 | -4.0838 | 1.996 | -0.61337 | 0.083581 | [12 empty] | 1.2716 | 2.0087 | 2.7864 | -0.44721 | 0.002281
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1166 | 5.5483 | [20 empty] | -0.091942 | 3.1906 | 5.8377 | -6.9709 | 6.2108 | -4.1014 | 2.0063 | -0.61707 | 0.084156 | [12 empty] | 1.2716 | 2.0105 | 2.7911 | -0.44836 | 0.0022888
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1167 | 5.554 | [20 empty] | -0.091942 | 3.1934 | 5.8477 | -6.9888 | 6.2321 | -4.119 | 2.0167 | -0.62078 | 0.084736 | [12 empty] | 1.2716 | 2.0122 | 2.7959 | -0.44951 | 0.0022967
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1168 | 5.5598 | [20 empty] | -0.091942 | 3.1961 | 5.8577 | -7.0068 | 6.2535 | -4.1367 | 2.0271 | -0.62452 | 0.085318 | [12 empty] | 1.2716 | 2.0139 | 2.8007 | -0.45067 | 0.0023046
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1169 | 5.5655 | [20 empty] | -0.091942 | 3.1988 | 5.8678 | -7.0248 | 6.2749 | -4.1544 | 2.0375 | -0.62827 | 0.085904 | [12 empty] | 1.2716 | 2.0156 | 2.8055 | -0.45183 | 0.0023125
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1170 | 5.5712 | [20 empty] | -0.091942 | 3.2016 | 5.8778 | -7.0429 | 6.2964 | -4.1722 | 2.048 | -0.63204 | 0.086494 | [12 empty] | 1.2716 | 2.0174 | 2.8103 | -0.45299 | 0.0023204
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1171 | 5.577 | [20 empty] | -0.091942 | 3.2043 | 5.8878 | -7.061 | 6.318 | -4.1901 | 2.0585 | -0.63583 | 0.087087 | [12 empty] | 1.2716 | 2.0191 | 2.8151 | -0.45415 | 0.0023283
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1172 | 5.5827 | [20 empty] | -0.091942 | 3.207 | 5.8979 | -7.0791 | 6.3396 | -4.208 | 2.0691 | -0.63964 | 0.087684 | [12 empty] | 1.2716 | 2.0208 | 2.8199 | -0.45532 | 0.0023363
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1173 | 5.5884 | [20 empty] | -0.091942 | 3.2098 | 5.908 | -7.0972 | 6.3613 | -4.226 | 2.0797 | -0.64347 | 0.088284 | [12 empty] | 1.2716 | 2.0225 | 2.8248 | -0.45648 | 0.0023443
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1174 | 5.5942 | [20 empty] | -0.091942 | 3.2125 | 5.9181 | -7.1154 | 6.383 | -4.244 | 2.0904 | -0.64732 | 0.088888 | [12 empty] | 1.2716 | 2.0243 | 2.8296 | -0.45765 | 0.0023523
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1175 | 5.5999 | [20 empty] | -0.091942 | 3.2153 | 5.9281 | -7.1336 | 6.4048 | -4.2621 | 2.1011 | -0.65119 | 0.089496 | [12 empty] | 1.2716 | 2.026 | 2.8344 | -0.45882 | 0.0023603
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1176 | 5.6056 | [20 empty] | -0.091942 | 3.218 | 5.9382 | -7.1518 | 6.4266 | -4.2803 | 2.1118 | -0.65508 | 0.090107 | [12 empty] | 1.2716 | 2.0277 | 2.8392 | -0.46 | 0.0023684
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1177 | 5.6114 | [20 empty] | -0.091942 | 3.2207 | 5.9483 | -7.1701 | 6.4485 | -4.2985 | 2.1226 | -0.65899 | 0.090722 | [12 empty] | 1.2716 | 2.0294 | 2.8441 | -0.46117 | 0.0023764
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1178 | 5.6171 | [20 empty] | -0.091942 | 3.2235 | 5.9585 | -7.1883 | 6.4704 | -4.3168 | 2.1335 | -0.66292 | 0.09134 | [12 empty] | 1.2716 | 2.0311 | 2.8489 | -0.46235 | 0.0023845
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1179 | 5.6229 | [20 empty] | -0.091942 | 3.2262 | 5.9686 | -7.2067 | 6.4924 | -4.3352 | 2.1443 | -0.66687 | 0.091962 | [12 empty] | 1.2716 | 2.0329 | 2.8537 | -0.46352 | 0.0023926
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1180 | 5.6286 | [20 empty] | -0.091942 | 3.2289 | 5.9787 | -7.225 | 6.5145 | -4.3536 | 2.1553 | -0.67084 | 0.092588 | [12 empty] | 1.2716 | 2.0346 | 2.8586 | -0.4647 | 0.0024007
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1181 | 5.6344 | [20 empty] | -0.091942 | 3.2317 | 5.9888 | -7.2434 | 6.5366 | -4.3721 | 2.1663 | -0.67483 | 0.093218 | [12 empty] | 1.2716 | 2.0363 | 2.8634 | -0.46589 | 0.0024089
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1182 | 5.6401 | [20 empty] | -0.091942 | 3.2344 | 5.999 | -7.2618 | 6.5588 | -4.3906 | 2.1773 | -0.67884 | 0.093851 | [12 empty] | 1.2716 | 2.038 | 2.8683 | -0.46707 | 0.0024171
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1183 | 5.6458 | [20 empty] | -0.091942 | 3.2371 | 6.0091 | -7.2803 | 6.581 | -4.4092 | 2.1884 | -0.68287 | 0.094488 | [12 empty] | 1.2716 | 2.0398 | 2.8731 | -0.46826 | 0.0024253
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1184 | 5.6516 | [20 empty] | -0.091942 | 3.2399 | 6.0193 | -7.2987 | 6.6033 | -4.4279 | 2.1995 | -0.68692 | 0.095129 | [12 empty] | 1.2716 | 2.0415 | 2.878 | -0.46945 | 0.0024335
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1185 | 5.6573 | [20 empty] | -0.091942 | 3.2426 | 6.0295 | -7.3173 | 6.6256 | -4.4466 | 2.2107 | -0.69099 | 0.095774 | [12 empty] | 1.2716 | 2.0432 | 2.8829 | -0.47064 | 0.0024417
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1186 | 5.6631 | [20 empty] | -0.091942 | 3.2454 | 6.0397 | -7.3358 | 6.648 | -4.4654 | 2.2219 | -0.69508 | 0.096422 | [12 empty] | 1.2716 | 2.0449 | 2.8877 | -0.47183 | 0.00245
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1187 | 5.6688 | [20 empty] | -0.091942 | 3.2481 | 6.0498 | -7.3544 | 6.6704 | -4.4843 | 2.2331 | -0.6992 | 0.097074 | [12 empty] | 1.2716 | 2.0467 | 2.8926 | -0.47302 | 0.0024582
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1188 | 5.6746 | [20 empty] | -0.091942 | 3.2508 | 6.06 | -7.373 | 6.693 | -4.5032 | 2.2445 | -0.70333 | 0.097731 | [12 empty] | 1.2716 | 2.0484 | 2.8975 | -0.47422 | 0.0024665
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1189 | 5.6803 | [20 empty] | -0.091942 | 3.2536 | 6.0703 | -7.3916 | 6.7155 | -4.5222 | 2.2558 | -0.70749 | 0.098391 | [12 empty] | 1.2716 | 2.0501 | 2.9023 | -0.47542 | 0.0024748
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1190 | 5.6861 | [20 empty] | -0.091942 | 3.2563 | 6.0805 | -7.4103 | 6.7381 | -4.5412 | 2.2672 | -0.71166 | 0.099055 | [12 empty] | 1.2716 | 2.0518 | 2.9072 | -0.47662 | 0.0024832
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1191 | 5.6919 | [20 empty] | -0.091942 | 3.259 | 6.0907 | -7.429 | 6.7608 | -4.5603 | 2.2787 | -0.71586 | 0.099722 | [12 empty] | 1.2716 | 2.0536 | 2.9121 | -0.47782 | 0.0024915
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1192 | 5.6976 | [20 empty] | -0.091942 | 3.2618 | 6.1009 | -7.4477 | 6.7835 | -4.5795 | 2.2902 | -0.72008 | 0.10039 | [12 empty] | 1.2716 | 2.0553 | 2.917 | -0.47903 | 0.0024999
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1193 | 5.7034 | [20 empty] | -0.091942 | 3.2645 | 6.1112 | -7.4665 | 6.8063 | -4.5987 | 2.3017 | -0.72431 | 0.10107 | [12 empty] | 1.2716 | 2.057 | 2.9219 | -0.48023 | 0.0025083
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1194 | 5.7091 | [20 empty] | -0.091942 | 3.2672 | 6.1214 | -7.4852 | 6.8292 | -4.618 | 2.3133 | -0.72858 | 0.10175 | [12 empty] | 1.2716 | 2.0587 | 2.9268 | -0.48144 | 0.0025167
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1195 | 5.7149 | [20 empty] | -0.091942 | 3.27 | 6.1317 | -7.5041 | 6.8521 | -4.6374 | 2.325 | -0.73286 | 0.10243 | [12 empty] | 1.2716 | 2.0605 | 2.9317 | -0.48265 | 0.0025252
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1196 | 5.7206 | [20 empty] | -0.091942 | 3.2727 | 6.1419 | -7.5229 | 6.8751 | -4.6569 | 2.3367 | -0.73716 | 0.10312 | [12 empty] | 1.2716 | 2.0622 | 2.9366 | -0.48387 | 0.0025336
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1197 | 5.7264 | [20 empty] | -0.091942 | 3.2755 | 6.1522 | -7.5418 | 6.8981 | -4.6764 | 2.3484 | -0.74149 | 0.10381 | [12 empty] | 1.2716 | 2.0639 | 2.9415 | -0.48508 | 0.0025421
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1198 | 5.7322 | [20 empty] | -0.091942 | 3.2782 | 6.1625 | -7.5607 | 6.9212 | -4.6959 | 2.3602 | -0.74583 | 0.10451 | [12 empty] | 1.2716 | 2.0656 | 2.9465 | -0.4863 | 0.0025506
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1199 | 5.7379 | [20 empty] | -0.091942 | 3.2809 | 6.1728 | -7.5797 | 6.9443 | -4.7156 | 2.3721 | -0.7502 | 0.10521 | [12 empty] | 1.2716 | 2.0674 | 2.9514 | -0.48752 | 0.0025591
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1200 | 5.7437 | [20 empty] | -0.091942 | 3.2837 | 6.1831 | -7.5987 | 6.9675 | -4.7353 | 2.384 | -0.75459 | 0.10591 | [12 empty] | 1.2716 | 2.0691 | 2.9563 | -0.48874 | 0.0025677
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1201 | 5.7495 | [20 empty] | -0.091942 | 3.2864 | 6.1934 | -7.6177 | 6.9908 | -4.755 | 2.3959 | -0.75901 | 0.10662 | [12 empty] | 1.2716 | 2.0708 | 2.9612 | -0.48996 | 0.0025763
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1202 | 5.7552 | [20 empty] | -0.091942 | 3.2891 | 6.2037 | -7.6367 | 7.0141 | -4.7748 | 2.4079 | -0.76344 | 0.10733 | [12 empty] | 1.2716 | 2.0725 | 2.9662 | -0.49118 | 0.0025849
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1203 | 5.761 | [20 empty] | -0.091942 | 3.2919 | 6.214 | -7.6558 | 7.0374 | -4.7947 | 2.4199 | -0.7679 | 0.10805 | [12 empty] | 1.2716 | 2.0743 | 2.9711 | -0.49241 | 0.0025935
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1204 | 5.7668 | [20 empty] | -0.091942 | 3.2946 | 6.2244 | -7.6749 | 7.0609 | -4.8147 | 2.432 | -0.77238 | 0.10877 | [12 empty] | 1.2716 | 2.076 | 2.976 | -0.49364 | 0.0026021
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1205 | 5.7725 | [20 empty] | -0.091942 | 3.2973 | 6.2347 | -7.694 | 7.0844 | -4.8347 | 2.4442 | -0.77688 | 0.1095 | [12 empty] | 1.2716 | 2.0777 | 2.981 | -0.49487 | 0.0026108
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1206 | 5.7783 | [20 empty] | -0.091942 | 3.3001 | 6.2451 | -7.7132 | 7.1079 | -4.8548 | 2.4564 | -0.7814 | 0.11022 | [12 empty] | 1.2716 | 2.0794 | 2.9859 | -0.4961 | 0.0026194
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1207 | 5.7841 | [20 empty] | -0.091942 | 3.3028 | 6.2554 | -7.7324 | 7.1315 | -4.875 | 2.4686 | -0.78595 | 0.11096 | [12 empty] | 1.2716 | 2.0812 | 2.9909 | -0.49734 | 0.0026281
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1208 | 5.7898 | [20 empty] | -0.091942 | 3.3056 | 6.2658 | -7.7516 | 7.1552 | -4.8952 | 2.4809 | -0.79052 | 0.1117 | [12 empty] | 1.2716 | 2.0829 | 2.9958 | -0.49858 | 0.0026369
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1209 | 5.7956 | [20 empty] | -0.091942 | 3.3083 | 6.2762 | -7.7709 | 7.1789 | -4.9155 | 2.4933 | -0.79511 | 0.11244 | [12 empty] | 1.2716 | 2.0846 | 3.0008 | -0.49982 | 0.0026456
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1210 | 5.8014 | [20 empty] | -0.091942 | 3.311 | 6.2866 | -7.7902 | 7.2027 | -4.9359 | 2.5057 | -0.79973 | 0.11318 | [12 empty] | 1.2716 | 2.0863 | 3.0058 | -0.50106 | 0.0026544
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1211 | 5.8072 | [20 empty] | -0.091942 | 3.3138 | 6.297 | -7.8095 | 7.2265 | -4.9563 | 2.5181 | -0.80437 | 0.11393 | [12 empty] | 1.2716 | 2.088 | 3.0107 | -0.5023 | 0.0026631
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1212 | 5.8129 | [20 empty] | -0.091942 | 3.3165 | 6.3074 | -7.8289 | 7.2504 | -4.9768 | 2.5306 | -0.80903 | 0.11469 | [12 empty] | 1.2716 | 2.0898 | 3.0157 | -0.50355 | 0.002672
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1213 | 5.8187 | [20 empty] | -0.091942 | 3.3192 | 6.3178 | -7.8483 | 7.2744 | -4.9974 | 2.5432 | -0.81371 | 0.11545 | [12 empty] | 1.2716 | 2.0915 | 3.0207 | -0.50479 | 0.0026808
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1214 | 5.8245 | [20 empty] | -0.091942 | 3.322 | 6.3282 | -7.8677 | 7.2984 | -5.018 | 2.5558 | -0.81842 | 0.11621 | [12 empty] | 1.2716 | 2.0932 | 3.0257 | -0.50604 | 0.0026896
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1215 | 5.8303 | [20 empty] | -0.091942 | 3.3247 | 6.3386 | -7.8872 | 7.3225 | -5.0387 | 2.5684 | -0.82315 | 0.11698 | [12 empty] | 1.2716 | 2.0949 | 3.0307 | -0.50729 | 0.0026985
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1216 | 5.836 | [20 empty] | -0.091942 | 3.3274 | 6.3491 | -7.9067 | 7.3466 | -5.0595 | 2.5812 | -0.8279 | 0.11775 | [12 empty] | 1.2716 | 2.0967 | 3.0357 | -0.50855 | 0.0027074
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1217 | 5.8418 | [20 empty] | -0.091942 | 3.3302 | 6.3595 | -7.9262 | 7.3708 | -5.0803 | 2.5939 | -0.83268 | 0.11853 | [12 empty] | 1.2716 | 2.0984 | 3.0407 | -0.5098 | 0.0027163
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1218 | 5.8476 | [20 empty] | -0.091942 | 3.3329 | 6.37 | -7.9457 | 7.395 | -5.1012 | 2.6067 | -0.83748 | 0.11931 | [12 empty] | 1.2716 | 2.1001 | 3.0457 | -0.51106 | 0.0027253
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1219 | 5.8534 | [20 empty] | -0.091942 | 3.3357 | 6.3804 | -7.9653 | 7.4194 | -5.1222 | 2.6196 | -0.84231 | 0.1201 | [12 empty] | 1.2716 | 2.1018 | 3.0507 | -0.51232 | 0.0027342
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1220 | 5.8592 | [20 empty] | -0.091942 | 3.3384 | 6.3909 | -7.985 | 7.4437 | -5.1432 | 2.6325 | -0.84716 | 0.12089 | [12 empty] | 1.2716 | 2.1036 | 3.0557 | -0.51358 | 0.0027432
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1221 | 5.8649 | [20 empty] | -0.091942 | 3.3411 | 6.4014 | -8.0046 | 7.4682 | -5.1643 | 2.6455 | -0.85203 | 0.12168 | [12 empty] | 1.2716 | 2.1053 | 3.0607 | -0.51485 | 0.0027522
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1222 | 5.8707 | [20 empty] | -0.091942 | 3.3439 | 6.4119 | -8.0243 | 7.4927 | -5.1855 | 2.6585 | -0.85692 | 0.12248 | [12 empty] | 1.2716 | 2.107 | 3.0657 | -0.51611 | 0.0027612
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1223 | 5.8765 | [20 empty] | -0.091942 | 3.3466 | 6.4224 | -8.044 | 7.5172 | -5.2068 | 2.6716 | -0.86185 | 0.12329 | [12 empty] | 1.2716 | 2.1087 | 3.0707 | -0.51738 | 0.0027703
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1224 | 5.8823 | [20 empty] | -0.091942 | 3.3493 | 6.4329 | -8.0638 | 7.5418 | -5.2281 | 2.6847 | -0.86679 | 0.12409 | [12 empty] | 1.2716 | 2.1105 | 3.0757 | -0.51865 | 0.0027794
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1225 | 5.8881 | [20 empty] | -0.091942 | 3.3521 | 6.4434 | -8.0835 | 7.5665 | -5.2495 | 2.6979 | -0.87176 | 0.12491 | [12 empty] | 1.2716 | 2.1122 | 3.0808 | -0.51992 | 0.0027885
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1226 | 5.8939 | [20 empty] | -0.091942 | 3.3548 | 6.4539 | -8.1033 | 7.5913 | -5.271 | 2.7112 | -0.87675 | 0.12573 | [12 empty] | 1.2716 | 2.1139 | 3.0858 | -0.5212 | 0.0027976
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1227 | 5.8997 | [20 empty] | -0.091942 | 3.3575 | 6.4645 | -8.1232 | 7.6161 | -5.2925 | 2.7245 | -0.88177 | 0.12655 | [12 empty] | 1.2716 | 2.1156 | 3.0908 | -0.52247 | 0.0028067
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1228 | 5.9055 | [20 empty] | -0.091942 | 3.3603 | 6.475 | -8.1431 | 7.6409 | -5.3141 | 2.7378 | -0.88681 | 0.12738 | [12 empty] | 1.2716 | 2.1174 | 3.0959 | -0.52375 | 0.0028159
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1229 | 5.9113 | [20 empty] | -0.091942 | 3.363 | 6.4855 | -8.163 | 7.6658 | -5.3358 | 2.7512 | -0.89188 | 0.12821 | [12 empty] | 1.2716 | 2.1191 | 3.1009 | -0.52503 | 0.002825
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1230 | 5.9171 | [20 empty] | -0.091942 | 3.3658 | 6.4961 | -8.1829 | 7.6908 | -5.3575 | 2.7647 | -0.89697 | 0.12904 | [12 empty] | 1.2716 | 2.1208 | 3.106 | -0.52632 | 0.0028343
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1231 | 5.9228 | [20 empty] | -0.091942 | 3.3685 | 6.5067 | -8.2029 | 7.7159 | -5.3793 | 2.7782 | -0.90209 | 0.12989 | [12 empty] | 1.2716 | 2.1225 | 3.111 | -0.5276 | 0.0028435
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1232 | 5.9286 | [20 empty] | -0.091942 | 3.3712 | 6.5172 | -8.2229 | 7.741 | -5.4012 | 2.7918 | -0.90723 | 0.13073 | [12 empty] | 1.2716 | 2.1243 | 3.1161 | -0.52889 | 0.0028527
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1233 | 5.9344 | [20 empty] | -0.091942 | 3.374 | 6.5278 | -8.2429 | 7.7661 | -5.4232 | 2.8054 | -0.9124 | 0.13158 | [12 empty] | 1.2716 | 2.126 | 3.1211 | -0.53018 | 0.002862
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1234 | 5.9402 | [20 empty] | -0.091942 | 3.3767 | 6.5384 | -8.263 | 7.7913 | -5.4452 | 2.8191 | -0.91759 | 0.13244 | [12 empty] | 1.2716 | 2.1277 | 3.1262 | -0.53147 | 0.0028713
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1235 | 5.946 | [20 empty] | -0.091942 | 3.3794 | 6.549 | -8.2831 | 7.8166 | -5.4673 | 2.8328 | -0.92281 | 0.1333 | [12 empty] | 1.2716 | 2.1294 | 3.1313 | -0.53276 | 0.0028806
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1236 | 5.9518 | [20 empty] | -0.091942 | 3.3822 | 6.5596 | -8.3033 | 7.842 | -5.4895 | 2.8466 | -0.92805 | 0.13417 | [12 empty] | 1.2716 | 2.1312 | 3.1363 | -0.53406 | 0.00289
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1237 | 5.9576 | [20 empty] | -0.091942 | 3.3849 | 6.5703 | -8.3234 | 7.8674 | -5.5117 | 2.8604 | -0.93332 | 0.13504 | [12 empty] | 1.2716 | 2.1329 | 3.1414 | -0.53535 | 0.0028993
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1238 | 5.9634 | [20 empty] | -0.091942 | 3.3876 | 6.5809 | -8.3436 | 7.8929 | -5.534 | 2.8743 | -0.93862 | 0.13591 | [12 empty] | 1.2716 | 2.1346 | 3.1465 | -0.53665 | 0.0029087
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1239 | 5.9692 | [20 empty] | -0.091942 | 3.3904 | 6.5915 | -8.3639 | 7.9184 | -5.5564 | 2.8883 | -0.94394 | 0.13679 | [12 empty] | 1.2716 | 2.1363 | 3.1516 | -0.53795 | 0.0029181
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1240 | 5.975 | [20 empty] | -0.091942 | 3.3931 | 6.6022 | -8.3841 | 7.944 | -5.5789 | 2.9023 | -0.94928 | 0.13768 | [12 empty] | 1.2716 | 2.1381 | 3.1567 | -0.53926 | 0.0029276
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1241 | 5.9808 | [20 empty] | -0.091942 | 3.3959 | 6.6128 | -8.4044 | 7.9696 | -5.6014 | 2.9164 | -0.95466 | 0.13857 | [12 empty] | 1.2716 | 2.1398 | 3.1618 | -0.54056 | 0.002937
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1242 | 5.9867 | [20 empty] | -0.091942 | 3.3986 | 6.6235 | -8.4248 | 7.9954 | -5.624 | 2.9305 | -0.96005 | 0.13947 | [12 empty] | 1.2716 | 2.1415 | 3.1669 | -0.54187 | 0.0029465
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1243 | 5.9925 | [20 empty] | -0.091942 | 3.4013 | 6.6341 | -8.4451 | 8.0211 | -5.6467 | 2.9447 | -0.96548 | 0.14037 | [12 empty] | 1.2716 | 2.1432 | 3.172 | -0.54318 | 0.002956
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1244 | 5.9983 | [20 empty] | -0.091942 | 3.4041 | 6.6448 | -8.4655 | 8.047 | -5.6694 | 2.9589 | -0.97093 | 0.14127 | [12 empty] | 1.2716 | 2.1449 | 3.1771 | -0.54449 | 0.0029655
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1245 | 6.0041 | [20 empty] | -0.091942 | 3.4068 | 6.6555 | -8.486 | 8.0729 | -5.6922 | 2.9732 | -0.9764 | 0.14218 | [12 empty] | 1.2716 | 2.1467 | 3.1822 | -0.54581 | 0.0029751
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1246 | 6.0099 | [20 empty] | -0.091942 | 3.4095 | 6.6662 | -8.5064 | 8.0989 | -5.7151 | 2.9876 | -0.98191 | 0.1431 | [12 empty] | 1.2716 | 2.1484 | 3.1873 | -0.54712 | 0.0029846
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1247 | 6.0157 | [20 empty] | -0.091942 | 3.4123 | 6.6769 | -8.5269 | 8.1249 | -5.7381 | 3.002 | -0.98744 | 0.14402 | [12 empty] | 1.2716 | 2.1501 | 3.1924 | -0.54844 | 0.0029942
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1248 | 6.0215 | [20 empty] | -0.091942 | 3.415 | 6.6876 | -8.5475 | 8.151 | -5.7612 | 3.0165 | -0.99299 | 0.14495 | [12 empty] | 1.2716 | 2.1518 | 3.1975 | -0.54976 | 0.0030038
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1249 | 6.0273 | [20 empty] | -0.091942 | 3.4177 | 6.6983 | -8.568 | 8.1771 | -5.7843 | 3.031 | -0.99858 | 0.14588 | [12 empty] | 1.2716 | 2.1536 | 3.2027 | -0.55109 | 0.0030135
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1250 | 6.0331 | [20 empty] | -0.091942 | 3.4205 | 6.7091 | -8.5886 | 8.2034 | -5.8075 | 3.0456 | -1.0042 | 0.14682 | [12 empty] | 1.2716 | 2.1553 | 3.2078 | -0.55241 | 0.0030231
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1251 | 6.0389 | [20 empty] | -0.091942 | 3.4232 | 6.7198 | -8.6092 | 8.2296 | -5.8307 | 3.0602 | -1.0098 | 0.14776 | [12 empty] | 1.2716 | 2.157 | 3.2129 | -0.55374 | 0.0030328
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1252 | 6.0448 | [20 empty] | -0.091942 | 3.426 | 6.7306 | -8.6299 | 8.256 | -5.8541 | 3.075 | -1.0155 | 0.14871 | [12 empty] | 1.2716 | 2.1587 | 3.2181 | -0.55507 | 0.0030425
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1253 | 6.0506 | [20 empty] | -0.091942 | 3.4287 | 6.7413 | -8.6506 | 8.2824 | -5.8775 | 3.0897 | -1.0212 | 0.14966 | [12 empty] | 1.2716 | 2.1605 | 3.2232 | -0.5564 | 0.0030523
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1254 | 6.0564 | [20 empty] | -0.091942 | 3.4314 | 6.7521 | -8.6713 | 8.3089 | -5.901 | 3.1045 | -1.0269 | 0.15062 | [12 empty] | 1.2716 | 2.1622 | 3.2284 | -0.55773 | 0.003062
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1255 | 6.0622 | [20 empty] | -0.091942 | 3.4342 | 6.7629 | -8.6921 | 8.3354 | -5.9246 | 3.1194 | -1.0326 | 0.15158 | [12 empty] | 1.2716 | 2.1639 | 3.2335 | -0.55907 | 0.0030718
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1256 | 6.068 | [20 empty] | -0.091942 | 3.4369 | 6.7736 | -8.7129 | 8.362 | -5.9482 | 3.1344 | -1.0384 | 0.15255 | [12 empty] | 1.2716 | 2.1656 | 3.2387 | -0.5604 | 0.0030816
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1257 | 6.0739 | [20 empty] | -0.091942 | 3.4396 | 6.7844 | -8.7337 | 8.3887 | -5.9719 | 3.1494 | -1.0442 | 0.15353 | [12 empty] | 1.2716 | 2.1674 | 3.2438 | -0.56174 | 0.0030914
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1258 | 6.0797 | [20 empty] | -0.091942 | 3.4424 | 6.7952 | -8.7546 | 8.4154 | -5.9957 | 3.1644 | -1.05 | 0.15451 | [12 empty] | 1.2716 | 2.1691 | 3.249 | -0.56308 | 0.0031013
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1259 | 6.0855 | [20 empty] | -0.091942 | 3.4451 | 6.806 | -8.7755 | 8.4422 | -6.0196 | 3.1796 | -1.0559 | 0.15549 | [12 empty] | 1.2716 | 2.1708 | 3.2541 | -0.56443 | 0.0031112
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1260 | 6.0913 | [20 empty] | -0.091942 | 3.4478 | 6.8169 | -8.7964 | 8.469 | -6.0435 | 3.1947 | -1.0618 | 0.15648 | [12 empty] | 1.2716 | 2.1725 | 3.2593 | -0.56577 | 0.003121
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1261 | 6.0972 | [20 empty] | -0.091942 | 3.4506 | 6.8277 | -8.8174 | 8.496 | -6.0675 | 3.21 | -1.0677 | 0.15748 | [12 empty] | 1.2716 | 2.1743 | 3.2645 | -0.56712 | 0.003131
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1262 | 6.103 | [20 empty] | -0.091942 | 3.4533 | 6.8385 | -8.8383 | 8.5229 | -6.0916 | 3.2253 | -1.0736 | 0.15848 | [12 empty] | 1.2716 | 2.176 | 3.2697 | -0.56847 | 0.0031409
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1263 | 6.1088 | [20 empty] | -0.091942 | 3.4561 | 6.8494 | -8.8594 | 8.55 | -6.1158 | 3.2407 | -1.0796 | 0.15949 | [12 empty] | 1.2716 | 2.1777 | 3.2749 | -0.56982 | 0.0031509
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1264 | 6.1146 | [20 empty] | -0.091942 | 3.4588 | 6.8602 | -8.8804 | 8.5771 | -6.1401 | 3.2561 | -1.0856 | 0.1605 | [12 empty] | 1.2716 | 2.1794 | 3.28 | -0.57118 | 0.0031609
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1265 | 6.1205 | [20 empty] | -0.091942 | 3.4615 | 6.8711 | -8.9015 | 8.6043 | -6.1644 | 3.2716 | -1.0916 | 0.16152 | [12 empty] | 1.2716 | 2.1812 | 3.2852 | -0.57254 | 0.0031709
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1266 | 6.1263 | [20 empty] | -0.091942 | 3.4643 | 6.8819 | -8.9227 | 8.6315 | -6.1888 | 3.2871 | -1.0977 | 0.16254 | [12 empty] | 1.2716 | 2.1829 | 3.2904 | -0.57389 | 0.0031809
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1267 | 6.1321 | [20 empty] | -0.091942 | 3.467 | 6.8928 | -8.9438 | 8.6588 | -6.2133 | 3.3027 | -1.1038 | 0.16357 | [12 empty] | 1.2716 | 2.1846 | 3.2956 | -0.57526 | 0.003191
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1268 | 6.138 | [20 empty] | -0.091942 | 3.4697 | 6.9037 | -8.965 | 8.6862 | -6.2378 | 3.3184 | -1.1099 | 0.16461 | [12 empty] | 1.2716 | 2.1863 | 3.3008 | -0.57662 | 0.0032011
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1269 | 6.1438 | [20 empty] | -0.091942 | 3.4725 | 6.9146 | -8.9862 | 8.7136 | -6.2625 | 3.3341 | -1.116 | 0.16565 | [12 empty] | 1.2716 | 2.1881 | 3.306 | -0.57798 | 0.0032112
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1270 | 6.1496 | [20 empty] | -0.091942 | 3.4752 | 6.9255 | -9.0075 | 8.7411 | -6.2872 | 3.3499 | -1.1222 | 0.1667 | [12 empty] | 1.2716 | 2.1898 | 3.3113 | -0.57935 | 0.0032213
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1271 | 6.1555 | [20 empty] | -0.091942 | 3.4779 | 6.9364 | -9.0288 | 8.7687 | -6.312 | 3.3658 | -1.1284 | 0.16775 | [12 empty] | 1.2716 | 2.1915 | 3.3165 | -0.58072 | 0.0032315
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1272 | 6.1613 | [20 empty] | -0.091942 | 3.4807 | 6.9473 | -9.0501 | 8.7963 | -6.3368 | 3.3817 | -1.1346 | 0.16881 | [12 empty] | 1.2716 | 2.1932 | 3.3217 | -0.58209 | 0.0032417
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1273 | 6.1671 | [20 empty] | -0.091942 | 3.4834 | 6.9582 | -9.0715 | 8.824 | -6.3618 | 3.3977 | -1.1409 | 0.16987 | [12 empty] | 1.2716 | 2.195 | 3.3269 | -0.58347 | 0.0032519
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1274 | 6.173 | [20 empty] | -0.091942 | 3.4862 | 6.9692 | -9.0929 | 8.8518 | -6.3868 | 3.4137 | -1.1472 | 0.17094 | [12 empty] | 1.2716 | 2.1967 | 3.3321 | -0.58484 | 0.0032621
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1275 | 6.1788 | [20 empty] | -0.091942 | 3.4889 | 6.9801 | -9.1143 | 8.8796 | -6.4119 | 3.4298 | -1.1535 | 0.17202 | [12 empty] | 1.2716 | 2.1984 | 3.3374 | -0.58622 | 0.0032723
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1276 | 6.1846 | [20 empty] | -0.091942 | 3.4916 | 6.9911 | -9.1358 | 8.9075 | -6.4371 | 3.446 | -1.1598 | 0.1731 | [12 empty] | 1.2716 | 2.2001 | 3.3426 | -0.5876 | 0.0032826
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1277 | 6.1905 | [20 empty] | -0.091942 | 3.4944 | 7.002 | -9.1573 | 8.9354 | -6.4624 | 3.4622 | -1.1662 | 0.17419 | [12 empty] | 1.2716 | 2.2018 | 3.3479 | -0.58898 | 0.0032929
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1278 | 6.1963 | [20 empty] | -0.091942 | 3.4971 | 7.013 | -9.1788 | 8.9635 | -6.4877 | 3.4785 | -1.1726 | 0.17529 | [12 empty] | 1.2716 | 2.2036 | 3.3531 | -0.59037 | 0.0033033
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1279 | 6.2022 | [20 empty] | -0.091942 | 3.4998 | 7.024 | -9.2004 | 8.9915 | -6.5131 | 3.4949 | -1.1791 | 0.17639 | [12 empty] | 1.2716 | 2.2053 | 3.3584 | -0.59176 | 0.0033136
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1280 | 6.208 | [20 empty] | -0.091942 | 3.5026 | 7.035 | -9.222 | 9.0197 | -6.5386 | 3.5113 | -1.1855 | 0.17749 | [12 empty] | 1.2716 | 2.207 | 3.3636 | -0.59315 | 0.003324
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1281 | 6.2139 | [20 empty] | -0.091942 | 3.5053 | 7.046 | -9.2436 | 9.0479 | -6.5642 | 3.5278 | -1.192 | 0.1786 | [12 empty] | 1.2716 | 2.2087 | 3.3689 | -0.59454 | 0.0033344
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1282 | 6.2197 | [20 empty] | -0.091942 | 3.508 | 7.057 | -9.2652 | 9.0762 | -6.5899 | 3.5444 | -1.1986 | 0.17972 | [12 empty] | 1.2716 | 2.2105 | 3.3741 | -0.59593 | 0.0033448
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1283 | 6.2256 | [20 empty] | -0.091942 | 3.5108 | 7.068 | -9.2869 | 9.1046 | -6.6156 | 3.561 | -1.2051 | 0.18085 | [12 empty] | 1.2716 | 2.2122 | 3.3794 | -0.59733 | 0.0033553
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1284 | 6.2314 | [20 empty] | -0.091942 | 3.5135 | 7.079 | -9.3087 | 9.133 | -6.6414 | 3.5777 | -1.2117 | 0.18198 | [12 empty] | 1.2716 | 2.2139 | 3.3847 | -0.59872 | 0.0033657
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1285 | 6.2373 | [20 empty] | -0.091942 | 3.5163 | 7.09 | -9.3304 | 9.1615 | -6.6673 | 3.5944 | -1.2183 | 0.18312 | [12 empty] | 1.2716 | 2.2156 | 3.3899 | -0.60012 | 0.0033762
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1286 | 6.2431 | [20 empty] | -0.091942 | 3.519 | 7.1011 | -9.3522 | 9.19 | -6.6933 | 3.6113 | -1.225 | 0.18426 | [12 empty] | 1.2716 | 2.2174 | 3.3952 | -0.60153 | 0.0033867
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1287 | 6.249 | [20 empty] | -0.091942 | 3.5217 | 7.1121 | -9.3741 | 9.2186 | -6.7194 | 3.6281 | -1.2317 | 0.18541 | [12 empty] | 1.2716 | 2.2191 | 3.4005 | -0.60293 | 0.0033973
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1288 | 6.2548 | [20 empty] | -0.091942 | 3.5245 | 7.1232 | -9.3959 | 9.2473 | -6.7455 | 3.6451 | -1.2384 | 0.18656 | [12 empty] | 1.2716 | 2.2208 | 3.4058 | -0.60434 | 0.0034079
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1289 | 6.2607 | [20 empty] | -0.091942 | 3.5272 | 7.1343 | -9.4178 | 9.2761 | -6.7718 | 3.6621 | -1.2451 | 0.18773 | [12 empty] | 1.2716 | 2.2225 | 3.4111 | -0.60575 | 0.0034185
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1290 | 6.2665 | [20 empty] | -0.091942 | 3.5299 | 7.1453 | -9.4398 | 9.3049 | -6.7981 | 3.6792 | -1.2519 | 0.18889 | [12 empty] | 1.2716 | 2.2243 | 3.4164 | -0.60716 | 0.0034291
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1291 | 6.2724 | [20 empty] | -0.091942 | 3.5327 | 7.1564 | -9.4618 | 9.3338 | -6.8245 | 3.6963 | -1.2587 | 0.19007 | [12 empty] | 1.2716 | 2.226 | 3.4217 | -0.60857 | 0.0034397
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1292 | 6.2782 | [20 empty] | -0.091942 | 3.5354 | 7.1675 | -9.4838 | 9.3627 | -6.8509 | 3.7135 | -1.2656 | 0.19125 | [12 empty] | 1.2716 | 2.2277 | 3.427 | -0.60998 | 0.0034504
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1293 | 6.2841 | [20 empty] | -0.091942 | 3.5381 | 7.1786 | -9.5058 | 9.3917 | -6.8775 | 3.7308 | -1.2724 | 0.19244 | [12 empty] | 1.2716 | 2.2294 | 3.4323 | -0.6114 | 0.0034611
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1294 | 6.2899 | [20 empty] | -0.091942 | 3.5409 | 7.1897 | -9.5279 | 9.4208 | -6.9041 | 3.7482 | -1.2793 | 0.19363 | [12 empty] | 1.2716 | 2.2312 | 3.4376 | -0.61282 | 0.0034718
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1295 | 6.2958 | [20 empty] | -0.091942 | 3.5436 | 7.2008 | -9.55 | 9.45 | -6.9308 | 3.7656 | -1.2863 | 0.19483 | [12 empty] | 1.2716 | 2.2329 | 3.4429 | -0.61424 | 0.0034826
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1296 | 6.3017 | [20 empty] | -0.091942 | 3.5464 | 7.212 | -9.5721 | 9.4792 | -6.9576 | 3.7831 | -1.2932 | 0.19604 | [12 empty] | 1.2716 | 2.2346 | 3.4482 | -0.61567 | 0.0034933
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1297 | 6.3075 | [20 empty] | -0.091942 | 3.5491 | 7.2231 | -9.5943 | 9.5085 | -6.9845 | 3.8006 | -1.3002 | 0.19725 | [12 empty] | 1.2716 | 2.2363 | 3.4535 | -0.61709 | 0.0035041
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1298 | 6.3134 | [20 empty] | -0.091942 | 3.5518 | 7.2342 | -9.6165 | 9.5379 | -7.0115 | 3.8182 | -1.3073 | 0.19847 | [12 empty] | 1.2716 | 2.2381 | 3.4589 | -0.61852 | 0.0035149
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1299 | 6.3193 | [20 empty] | -0.091942 | 3.5546 | 7.2454 | -9.6387 | 9.5673 | -7.0385 | 3.8359 | -1.3143 | 0.1997 | [12 empty] | 1.2716 | 2.2398 | 3.4642 | -0.61995 | 0.0035258
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1300 | 6.3251 | [20 empty] | -0.091942 | 3.5573 | 7.2565 | -9.661 | 9.5968 | -7.0657 | 3.8537 | -1.3214 | 0.20093 | [12 empty] | 1.2716 | 2.2415 | 3.4695 | -0.62139 | 0.0035366
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1301 | 6.331 | [20 empty] | -0.091942 | 3.56 | 7.2677 | -9.6833 | 9.6263 | -7.0929 | 3.8715 | -1.3286 | 0.20217 | [12 empty] | 1.2716 | 2.2432 | 3.4749 | -0.62282 | 0.0035475
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1302 | 6.3369 | [20 empty] | -0.091942 | 3.5628 | 7.2789 | -9.7057 | 9.656 | -7.1202 | 3.8894 | -1.3357 | 0.20342 | [12 empty] | 1.2716 | 2.245 | 3.4802 | -0.62426 | 0.0035585
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1303 | 6.3427 | [20 empty] | -0.091942 | 3.5655 | 7.2901 | -9.7281 | 9.6857 | -7.1476 | 3.9073 | -1.3429 | 0.20467 | [12 empty] | 1.2716 | 2.2467 | 3.4856 | -0.6257 | 0.0035694
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1304 | 6.3486 | [20 empty] | -0.091942 | 3.5682 | 7.3013 | -9.7505 | 9.7154 | -7.175 | 3.9253 | -1.3502 | 0.20593 | [12 empty] | 1.2716 | 2.2484 | 3.4909 | -0.62714 | 0.0035804
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1305 | 6.3545 | [20 empty] | -0.091942 | 3.571 | 7.3125 | -9.7729 | 9.7453 | -7.2026 | 3.9434 | -1.3574 | 0.2072 | [12 empty] | 1.2716 | 2.2501 | 3.4963 | -0.62858 | 0.0035914
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1306 | 6.3603 | [20 empty] | -0.091942 | 3.5737 | 7.3237 | -9.7954 | 9.7752 | -7.2302 | 3.9616 | -1.3647 | 0.20847 | [12 empty] | 1.2716 | 2.2519 | 3.5016 | -0.63003 | 0.0036024
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1307 | 6.3662 | [20 empty] | -0.091942 | 3.5765 | 7.3349 | -9.8179 | 9.8052 | -7.258 | 3.9798 | -1.3721 | 0.20975 | [12 empty] | 1.2716 | 2.2536 | 3.507 | -0.63148 | 0.0036134
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1308 | 6.3721 | [20 empty] | -0.091942 | 3.5792 | 7.3461 | -9.8405 | 9.8352 | -7.2858 | 3.9981 | -1.3794 | 0.21104 | [12 empty] | 1.2716 | 2.2553 | 3.5124 | -0.63293 | 0.0036245
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1309 | 6.378 | [20 empty] | -0.091942 | 3.5819 | 7.3574 | -9.8631 | 9.8653 | -7.3137 | 4.0165 | -1.3868 | 0.21233 | [12 empty] | 1.2716 | 2.257 | 3.5177 | -0.63438 | 0.0036356
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1310 | 6.3838 | [20 empty] | -0.091942 | 3.5847 | 7.3686 | -9.8857 | 9.8955 | -7.3416 | 4.035 | -1.3943 | 0.21363 | [12 empty] | 1.2716 | 2.2587 | 3.5231 | -0.63584 | 0.0036467
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1311 | 6.3897 | [20 empty] | -0.091942 | 3.5874 | 7.3799 | -9.9083 | 9.9257 | -7.3697 | 4.0535 | -1.4017 | 0.21494 | [12 empty] | 1.2716 | 2.2605 | 3.5285 | -0.63729 | 0.0036579
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1312 | 6.3956 | [20 empty] | -0.091942 | 3.5901 | 7.3911 | -9.931 | 9.9561 | -7.3979 | 4.0721 | -1.4092 | 0.21626 | [12 empty] | 1.2716 | 2.2622 | 3.5339 | -0.63875 | 0.0036691
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1313 | 6.4015 | [20 empty] | -0.091942 | 3.5929 | 7.4024 | -9.9538 | 9.9864 | -7.4261 | 4.0907 | -1.4168 | 0.21758 | [12 empty] | 1.2716 | 2.2639 | 3.5393 | -0.64021 | 0.0036803
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1314 | 6.4073 | [20 empty] | -0.091942 | 3.5956 | 7.4137 | -9.9765 | 10.0169 | -7.4544 | 4.1095 | -1.4243 | 0.21891 | [12 empty] | 1.2716 | 2.2656 | 3.5447 | -0.64168 | 0.0036915
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1315 | 6.4132 | [20 empty] | -0.091942 | 3.5983 | 7.425 | -9.9993 | 10.0474 | -7.4828 | 4.1283 | -1.4319 | 0.22024 | [12 empty] | 1.2716 | 2.2674 | 3.5501 | -0.64314 | 0.0037027
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1316 | 6.4191 | [20 empty] | -0.091942 | 3.6011 | 7.4363 | -10.0221 | 10.078 | -7.5113 | 4.1471 | -1.4396 | 0.22159 | [12 empty] | 1.2716 | 2.2691 | 3.5555 | -0.64461 | 0.003714
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1317 | 6.425 | [20 empty] | -0.091942 | 3.6038 | 7.4476 | -10.045 | 10.1087 | -7.5399 | 4.1661 | -1.4473 | 0.22294 | [12 empty] | 1.2716 | 2.2708 | 3.5609 | -0.64608 | 0.0037253
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1318 | 6.4309 | [20 empty] | -0.091942 | 3.6066 | 7.4589 | -10.0679 | 10.1394 | -7.5686 | 4.1851 | -1.455 | 0.2243 | [12 empty] | 1.2716 | 2.2725 | 3.5663 | -0.64756 | 0.0037366
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1319 | 6.4368 | [20 empty] | -0.091942 | 3.6093 | 7.4702 | -10.0908 | 10.1702 | -7.5973 | 4.2042 | -1.4627 | 0.22566 | [12 empty] | 1.2716 | 2.2743 | 3.5717 | -0.64903 | 0.003748
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1320 | 6.4426 | [20 empty] | -0.091942 | 3.612 | 7.4815 | -10.1138 | 10.2011 | -7.6262 | 4.2233 | -1.4705 | 0.22703 | [12 empty] | 1.2716 | 2.276 | 3.5771 | -0.65051 | 0.0037594
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1321 | 6.4485 | [20 empty] | -0.091942 | 3.6148 | 7.4929 | -10.1368 | 10.2321 | -7.6551 | 4.2426 | -1.4783 | 0.22841 | [12 empty] | 1.2716 | 2.2777 | 3.5825 | -0.65199 | 0.0037708
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1322 | 6.4544 | [20 empty] | -0.091942 | 3.6175 | 7.5042 | -10.1598 | 10.2631 | -7.6841 | 4.2619 | -1.4862 | 0.2298 | [12 empty] | 1.2716 | 2.2794 | 3.588 | -0.65347 | 0.0037822
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1323 | 6.4603 | [20 empty] | -0.091942 | 3.6202 | 7.5156 | -10.1829 | 10.2942 | -7.7132 | 4.2813 | -1.494 | 0.23119 | [12 empty] | 1.2716 | 2.2812 | 3.5934 | -0.65495 | 0.0037937
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1324 | 6.4662 | [20 empty] | -0.091942 | 3.623 | 7.5269 | -10.206 | 10.3253 | -7.7424 | 4.3007 | -1.502 | 0.2326 | [12 empty] | 1.2716 | 2.2829 | 3.5988 | -0.65644 | 0.0038051
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1325 | 6.4721 | [20 empty] | -0.091942 | 3.6257 | 7.5383 | -10.2292 | 10.3566 | -7.7717 | 4.3202 | -1.5099 | 0.23401 | [12 empty] | 1.2716 | 2.2846 | 3.6043 | -0.65793 | 0.0038166
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1326 | 6.478 | [20 empty] | -0.091942 | 3.6284 | 7.5497 | -10.2524 | 10.3879 | -7.8011 | 4.3398 | -1.5179 | 0.23542 | [12 empty] | 1.2716 | 2.2863 | 3.6097 | -0.65942 | 0.0038282
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1327 | 6.4839 | [20 empty] | -0.091942 | 3.6312 | 7.5611 | -10.2756 | 10.4192 | -7.8305 | 4.3595 | -1.526 | 0.23685 | [12 empty] | 1.2716 | 2.2881 | 3.6152 | -0.66091 | 0.0038397
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1328 | 6.4898 | [20 empty] | -0.091942 | 3.6339 | 7.5725 | -10.2988 | 10.4507 | -7.8601 | 4.3793 | -1.534 | 0.23828 | [12 empty] | 1.2716 | 2.2898 | 3.6206 | -0.66241 | 0.0038513
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1329 | 6.4957 | [20 empty] | -0.091942 | 3.6367 | 7.5839 | -10.3221 | 10.4822 | -7.8897 | 4.3991 | -1.5421 | 0.23972 | [12 empty] | 1.2716 | 2.2915 | 3.6261 | -0.66391 | 0.0038629
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1330 | 6.5016 | [20 empty] | -0.091942 | 3.6394 | 7.5953 | -10.3454 | 10.5138 | -7.9195 | 4.419 | -1.5503 | 0.24116 | [12 empty] | 1.2716 | 2.2932 | 3.6315 | -0.66541 | 0.0038746
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1331 | 6.5075 | [20 empty] | -0.091942 | 3.6421 | 7.6067 | -10.3688 | 10.5454 | -7.9493 | 4.439 | -1.5584 | 0.24262 | [12 empty] | 1.2716 | 2.295 | 3.637 | -0.66691 | 0.0038862
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1332 | 6.5134 | [20 empty] | -0.091942 | 3.6449 | 7.6182 | -10.3922 | 10.5772 | -7.9792 | 4.459 | -1.5667 | 0.24408 | [12 empty] | 1.2716 | 2.2967 | 3.6425 | -0.66841 | 0.0038979
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1333 | 6.5193 | [20 empty] | -0.091942 | 3.6476 | 7.6296 | -10.4156 | 10.609 | -8.0092 | 4.4791 | -1.5749 | 0.24555 | [12 empty] | 1.2716 | 2.2984 | 3.6479 | -0.66992 | 0.0039097
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1334 | 6.5252 | [20 empty] | -0.091942 | 3.6503 | 7.6411 | -10.439 | 10.6408 | -8.0393 | 4.4993 | -1.5832 | 0.24703 | [12 empty] | 1.2716 | 2.3001 | 3.6534 | -0.67143 | 0.0039214
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1335 | 6.5311 | [20 empty] | -0.091942 | 3.6531 | 7.6525 | -10.4625 | 10.6728 | -8.0694 | 4.5196 | -1.5915 | 0.24851 | [12 empty] | 1.2716 | 2.3019 | 3.6589 | -0.67294 | 0.0039332
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1336 | 6.537 | [20 empty] | -0.091942 | 3.6558 | 7.664 | -10.4861 | 10.7048 | -8.0997 | 4.54 | -1.5999 | 0.25001 | [12 empty] | 1.2716 | 2.3036 | 3.6644 | -0.67445 | 0.003945
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1337 | 6.5429 | [20 empty] | -0.091942 | 3.6585 | 7.6755 | -10.5096 | 10.7369 | -8.1301 | 4.5604 | -1.6083 | 0.25151 | [12 empty] | 1.2716 | 2.3053 | 3.6699 | -0.67597 | 0.0039568
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1338 | 6.5488 | [20 empty] | -0.091942 | 3.6613 | 7.687 | -10.5332 | 10.769 | -8.1605 | 4.5809 | -1.6167 | 0.25302 | [12 empty] | 1.2716 | 2.307 | 3.6753 | -0.67748 | 0.0039687
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1339 | 6.5547 | [20 empty] | -0.091942 | 3.664 | 7.6985 | -10.5569 | 10.8013 | -8.1911 | 4.6015 | -1.6252 | 0.25453 | [12 empty] | 1.2716 | 2.3088 | 3.6808 | -0.67901 | 0.0039805
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1340 | 6.5606 | [20 empty] | -0.091942 | 3.6668 | 7.71 | -10.5805 | 10.8336 | -8.2217 | 4.6221 | -1.6337 | 0.25606 | [12 empty] | 1.2716 | 2.3105 | 3.6863 | -0.68053 | 0.0039924
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1341 | 6.5665 | [20 empty] | -0.091942 | 3.6695 | 7.7215 | -10.6042 | 10.8659 | -8.2524 | 4.6429 | -1.6423 | 0.25759 | [12 empty] | 1.2716 | 2.3122 | 3.6918 | -0.68205 | 0.0040044
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1342 | 6.5724 | [20 empty] | -0.091942 | 3.6722 | 7.733 | -10.628 | 10.8984 | -8.2832 | 4.6637 | -1.6509 | 0.25913 | [12 empty] | 1.2716 | 2.3139 | 3.6974 | -0.68358 | 0.0040163
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1343 | 6.5783 | [20 empty] | -0.091942 | 3.675 | 7.7445 | -10.6517 | 10.9309 | -8.3141 | 4.6846 | -1.6595 | 0.26068 | [12 empty] | 1.2716 | 2.3156 | 3.7029 | -0.68511 | 0.0040283
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1344 | 6.5842 | [20 empty] | -0.091942 | 3.6777 | 7.7561 | -10.6756 | 10.9635 | -8.3451 | 4.7055 | -1.6682 | 0.26224 | [12 empty] | 1.2716 | 2.3174 | 3.7084 | -0.68664 | 0.0040403
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1345 | 6.5901 | [20 empty] | -0.091942 | 3.6804 | 7.7676 | -10.6994 | 10.9962 | -8.3762 | 4.7266 | -1.6769 | 0.2638 | [12 empty] | 1.2716 | 2.3191 | 3.7139 | -0.68817 | 0.0040524
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1346 | 6.596 | [20 empty] | -0.091942 | 3.6832 | 7.7792 | -10.7233 | 11.0289 | -8.4074 | 4.7477 | -1.6856 | 0.26537 | [12 empty] | 1.2716 | 2.3208 | 3.7194 | -0.68971 | 0.0040644
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1347 | 6.602 | [20 empty] | -0.091942 | 3.6859 | 7.7907 | -10.7472 | 11.0617 | -8.4387 | 4.7689 | -1.6944 | 0.26696 | [12 empty] | 1.2716 | 2.3225 | 3.725 | -0.69125 | 0.0040765
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1348 | 6.6079 | [20 empty] | -0.091942 | 3.6886 | 7.8023 | -10.7712 | 11.0946 | -8.4701 | 4.7902 | -1.7032 | 0.26855 | [12 empty] | 1.2716 | 2.3243 | 3.7305 | -0.69279 | 0.0040886
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1349 | 6.6138 | [20 empty] | -0.091942 | 3.6914 | 7.8139 | -10.7952 | 11.1276 | -8.5015 | 4.8115 | -1.7121 | 0.27014 | [12 empty] | 1.2716 | 2.326 | 3.736 | -0.69433 | 0.0041008
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1350 | 6.6197 | [20 empty] | -0.091942 | 3.6941 | 7.8255 | -10.8192 | 11.1606 | -8.5331 | 4.833 | -1.721 | 0.27175 | [12 empty] | 1.2716 | 2.3277 | 3.7416 | -0.69588 | 0.004113
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1351 | 6.6256 | [20 empty] | -0.091942 | 3.6969 | 7.8371 | -10.8432 | 11.1937 | -8.5647 | 4.8545 | -1.7299 | 0.27336 | [12 empty] | 1.2716 | 2.3294 | 3.7471 | -0.69742 | 0.0041252
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1352 | 6.6316 | [20 empty] | -0.091942 | 3.6996 | 7.8487 | -10.8673 | 11.2269 | -8.5965 | 4.8761 | -1.7389 | 0.27499 | [12 empty] | 1.2716 | 2.3312 | 3.7527 | -0.69897 | 0.0041374
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1353 | 6.6375 | [20 empty] | -0.091942 | 3.7023 | 7.8603 | -10.8915 | 11.2601 | -8.6283 | 4.8978 | -1.7479 | 0.27662 | [12 empty] | 1.2716 | 2.3329 | 3.7582 | -0.70053 | 0.0041496
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1354 | 6.6434 | [20 empty] | -0.091942 | 3.7051 | 7.8719 | -10.9156 | 11.2934 | -8.6602 | 4.9195 | -1.757 | 0.27826 | [12 empty] | 1.2716 | 2.3346 | 3.7638 | -0.70208 | 0.0041619
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1355 | 6.6493 | [20 empty] | -0.091942 | 3.7078 | 7.8835 | -10.9398 | 11.3268 | -8.6923 | 4.9414 | -1.7661 | 0.27991 | [12 empty] | 1.2716 | 2.3363 | 3.7693 | -0.70364 | 0.0041742
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1356 | 6.6552 | [20 empty] | -0.091942 | 3.7105 | 7.8952 | -10.9641 | 11.3603 | -8.7244 | 4.9633 | -1.7753 | 0.28156 | [12 empty] | 1.2716 | 2.3381 | 3.7749 | -0.7052 | 0.0041866
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1357 | 6.6612 | [20 empty] | -0.091942 | 3.7133 | 7.9068 | -10.9883 | 11.3939 | -8.7566 | 4.9853 | -1.7844 | 0.28323 | [12 empty] | 1.2716 | 2.3398 | 3.7805 | -0.70676 | 0.0041989
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1358 | 6.6671 | [20 empty] | -0.091942 | 3.716 | 7.9185 | -11.0127 | 11.4275 | -8.7889 | 5.0074 | -1.7937 | 0.2849 | [12 empty] | 1.2716 | 2.3415 | 3.786 | -0.70832 | 0.0042113
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1359 | 6.673 | [20 empty] | -0.091942 | 3.7187 | 7.9302 | -11.037 | 11.4612 | -8.8213 | 5.0296 | -1.8029 | 0.28659 | [12 empty] | 1.2716 | 2.3432 | 3.7916 | -0.70989 | 0.0042237
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1360 | 6.6789 | [20 empty] | -0.091942 | 3.7215 | 7.9418 | -11.0614 | 11.495 | -8.8538 | 5.0518 | -1.8122 | 0.28828 | [12 empty] | 1.2716 | 2.345 | 3.7972 | -0.71146 | 0.0042362
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1361 | 6.6849 | [20 empty] | -0.091942 | 3.7242 | 7.9535 | -11.0858 | 11.5288 | -8.8864 | 5.0741 | -1.8216 | 0.28998 | [12 empty] | 1.2716 | 2.3467 | 3.8028 | -0.71303 | 0.0042487
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1362 | 6.6908 | [20 empty] | -0.091942 | 3.727 | 7.9652 | -11.1103 | 11.5627 | -8.9191 | 5.0965 | -1.831 | 0.29169 | [12 empty] | 1.2716 | 2.3484 | 3.8084 | -0.7146 | 0.0042612
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1363 | 6.6967 | [20 empty] | -0.091942 | 3.7297 | 7.9769 | -11.1348 | 11.5967 | -8.9519 | 5.119 | -1.8404 | 0.2934 | [12 empty] | 1.2716 | 2.3501 | 3.814 | -0.71617 | 0.0042737
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1364 | 6.7027 | [20 empty] | -0.091942 | 3.7324 | 7.9886 | -11.1593 | 11.6308 | -8.9848 | 5.1416 | -1.8499 | 0.29513 | [12 empty] | 1.2716 | 2.3519 | 3.8196 | -0.71775 | 0.0042862
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1365 | 6.7086 | [20 empty] | -0.091942 | 3.7352 | 8.0003 | -11.1838 | 11.6649 | -9.0178 | 5.1643 | -1.8594 | 0.29687 | [12 empty] | 1.2716 | 2.3536 | 3.8252 | -0.71933 | 0.0042988
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1366 | 6.7145 | [20 empty] | -0.091942 | 3.7379 | 8.0121 | -11.2084 | 11.6992 | -9.0509 | 5.187 | -1.869 | 0.29861 | [12 empty] | 1.2716 | 2.3553 | 3.8308 | -0.72091 | 0.0043114
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1367 | 6.7205 | [20 empty] | -0.091942 | 3.7406 | 8.0238 | -11.2331 | 11.7335 | -9.0841 | 5.2098 | -1.8786 | 0.30036 | [12 empty] | 1.2716 | 2.357 | 3.8364 | -0.7225 | 0.0043241
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1368 | 6.7264 | [20 empty] | -0.091942 | 3.7434 | 8.0355 | -11.2577 | 11.7678 | -9.1173 | 5.2327 | -1.8882 | 0.30213 | [12 empty] | 1.2716 | 2.3588 | 3.842 | -0.72409 | 0.0043367
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1369 | 6.7323 | [20 empty] | -0.091942 | 3.7461 | 8.0473 | -11.2824 | 11.8023 | -9.1507 | 5.2557 | -1.8979 | 0.3039 | [12 empty] | 1.2716 | 2.3605 | 3.8476 | -0.72567 | 0.0043494
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1370 | 6.7383 | [20 empty] | -0.091942 | 3.7489 | 8.0591 | -11.3072 | 11.8368 | -9.1842 | 5.2788 | -1.9076 | 0.30568 | [12 empty] | 1.2716 | 2.3622 | 3.8532 | -0.72727 | 0.0043622
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1371 | 6.7442 | [20 empty] | -0.091942 | 3.7516 | 8.0708 | -11.332 | 11.8714 | -9.2177 | 5.302 | -1.9174 | 0.30747 | [12 empty] | 1.2716 | 2.3639 | 3.8589 | -0.72886 | 0.0043749
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1372 | 6.7502 | [20 empty] | -0.091942 | 3.7543 | 8.0826 | -11.3568 | 11.9061 | -9.2514 | 5.3252 | -1.9272 | 0.30927 | [12 empty] | 1.2716 | 2.3657 | 3.8645 | -0.73046 | 0.0043877
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1373 | 6.7561 | [20 empty] | -0.091942 | 3.7571 | 8.0944 | -11.3816 | 11.9408 | -9.2852 | 5.3485 | -1.937 | 0.31107 | [12 empty] | 1.2716 | 2.3674 | 3.8701 | -0.73205 | 0.0044005
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1374 | 6.7621 | [20 empty] | -0.091942 | 3.7598 | 8.1062 | -11.4065 | 11.9756 | -9.319 | 5.372 | -1.9469 | 0.31289 | [12 empty] | 1.2716 | 2.3691 | 3.8758 | -0.73365 | 0.0044133
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1375 | 6.768 | [20 empty] | -0.091942 | 3.7625 | 8.118 | -11.4314 | 12.0105 | -9.353 | 5.3955 | -1.9569 | 0.31472 | [12 empty] | 1.2716 | 2.3708 | 3.8814 | -0.73526 | 0.0044262
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1376 | 6.7739 | [20 empty] | -0.091942 | 3.7653 | 8.1298 | -11.4564 | 12.0455 | -9.3871 | 5.419 | -1.9669 | 0.31655 | [12 empty] | 1.2716 | 2.3725 | 3.8871 | -0.73686 | 0.0044391
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1377 | 6.7799 | [20 empty] | -0.091942 | 3.768 | 8.1416 | -11.4814 | 12.0806 | -9.4212 | 5.4427 | -1.9769 | 0.3184 | [12 empty] | 1.2716 | 2.3743 | 3.8927 | -0.73847 | 0.004452
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1378 | 6.7858 | [20 empty] | -0.091942 | 3.7707 | 8.1534 | -11.5064 | 12.1157 | -9.4555 | 5.4665 | -1.987 | 0.32025 | [12 empty] | 1.2716 | 2.376 | 3.8984 | -0.74008 | 0.0044649
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1379 | 6.7918 | [20 empty] | -0.091942 | 3.7735 | 8.1653 | -11.5315 | 12.1509 | -9.4898 | 5.4903 | -1.9971 | 0.32212 | [12 empty] | 1.2716 | 2.3777 | 3.904 | -0.74169 | 0.0044779
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1380 | 6.7977 | [20 empty] | -0.091942 | 3.7762 | 8.1771 | -11.5566 | 12.1862 | -9.5243 | 5.5143 | -2.0072 | 0.32399 | [12 empty] | 1.2716 | 2.3794 | 3.9097 | -0.74331 | 0.0044909
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1381 | 6.8037 | [20 empty] | -0.091942 | 3.779 | 8.189 | -11.5817 | 12.2216 | -9.5588 | 5.5383 | -2.0174 | 0.32587 | [12 empty] | 1.2716 | 2.3812 | 3.9154 | -0.74492 | 0.0045039
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1382 | 6.8096 | [20 empty] | -0.091942 | 3.7817 | 8.2009 | -11.6069 | 12.257 | -9.5935 | 5.5624 | -2.0277 | 0.32777 | [12 empty] | 1.2716 | 2.3829 | 3.921 | -0.74654 | 0.004517
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1383 | 6.8156 | [20 empty] | -0.091942 | 3.7844 | 8.2127 | -11.6321 | 12.2925 | -9.6283 | 5.5866 | -2.038 | 0.32967 | [12 empty] | 1.2716 | 2.3846 | 3.9267 | -0.74817 | 0.0045301
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1384 | 6.8215 | [20 empty] | -0.091942 | 3.7872 | 8.2246 | -11.6574 | 12.3281 | -9.6631 | 5.6109 | -2.0483 | 0.33158 | [12 empty] | 1.2716 | 2.3863 | 3.9324 | -0.74979 | 0.0045432
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1385 | 6.8275 | [20 empty] | -0.091942 | 3.7899 | 8.2365 | -11.6827 | 12.3638 | -9.6981 | 5.6352 | -2.0587 | 0.3335 | [12 empty] | 1.2716 | 2.3881 | 3.9381 | -0.75142 | 0.0045564
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1386 | 6.8335 | [20 empty] | -0.091942 | 3.7926 | 8.2484 | -11.708 | 12.3995 | -9.7331 | 5.6597 | -2.0691 | 0.33543 | [12 empty] | 1.2716 | 2.3898 | 3.9438 | -0.75305 | 0.0045695
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1387 | 6.8394 | [20 empty] | -0.091942 | 3.7954 | 8.2603 | -11.7334 | 12.4353 | -9.7683 | 5.6842 | -2.0796 | 0.33737 | [12 empty] | 1.2716 | 2.3915 | 3.9495 | -0.75468 | 0.0045827
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1388 | 6.8454 | [20 empty] | -0.091942 | 3.7981 | 8.2722 | -11.7588 | 12.4712 | -9.8036 | 5.7089 | -2.0901 | 0.33932 | [12 empty] | 1.2716 | 2.3932 | 3.9552 | -0.75631 | 0.004596
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1389 | 6.8513 | [20 empty] | -0.091942 | 3.8008 | 8.2841 | -11.7842 | 12.5072 | -9.8389 | 5.7336 | -2.1007 | 0.34128 | [12 empty] | 1.2716 | 2.395 | 3.9609 | -0.75795 | 0.0046092
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1390 | 6.8573 | [20 empty] | -0.091942 | 3.8036 | 8.2961 | -11.8097 | 12.5433 | -9.8744 | 5.7584 | -2.1113 | 0.34326 | [12 empty] | 1.2716 | 2.3967 | 3.9666 | -0.75958 | 0.0046225
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1391 | 6.8632 | [20 empty] | -0.091942 | 3.8063 | 8.308 | -11.8352 | 12.5794 | -9.91 | 5.7833 | -2.1219 | 0.34524 | [12 empty] | 1.2716 | 2.3984 | 3.9723 | -0.76122 | 0.0046358
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1392 | 6.8692 | [20 empty] | -0.091942 | 3.8091 | 8.32 | -11.8607 | 12.6156 | -9.9456 | 5.8083 | -2.1326 | 0.34723 | [12 empty] | 1.2716 | 2.4001 | 3.978 | -0.76287 | 0.0046492
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1393 | 6.8752 | [20 empty] | -0.091942 | 3.8118 | 8.3319 | -11.8863 | 12.6519 | -9.9814 | 5.8334 | -2.1434 | 0.34923 | [12 empty] | 1.2716 | 2.4019 | 3.9837 | -0.76451 | 0.0046625
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1394 | 6.8811 | [20 empty] | -0.091942 | 3.8145 | 8.3439 | -11.9119 | 12.6883 | -10.0173 | 5.8585 | -2.1542 | 0.35124 | [12 empty] | 1.2716 | 2.4036 | 3.9894 | -0.76616 | 0.0046759
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1395 | 6.8871 | [20 empty] | -0.091942 | 3.8173 | 8.3559 | -11.9376 | 12.7247 | -10.0533 | 5.8838 | -2.165 | 0.35326 | [12 empty] | 1.2716 | 2.4053 | 3.9952 | -0.76781 | 0.0046894
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1396 | 6.8931 | [20 empty] | -0.091942 | 3.82 | 8.3678 | -11.9632 | 12.7613 | -10.0894 | 5.9092 | -2.1759 | 0.35529 | [12 empty] | 1.2716 | 2.407 | 4.0009 | -0.76946 | 0.0047028
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1397 | 6.899 | [20 empty] | -0.091942 | 3.8227 | 8.3798 | -11.989 | 12.7979 | -10.1256 | 5.9346 | -2.1869 | 0.35733 | [12 empty] | 1.2716 | 2.4088 | 4.0066 | -0.77112 | 0.0047163
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1398 | 6.905 | [20 empty] | -0.091942 | 3.8255 | 8.3918 | -12.0147 | 12.8345 | -10.1619 | 5.9601 | -2.1978 | 0.35938 | [12 empty] | 1.2716 | 2.4105 | 4.0124 | -0.77277 | 0.0047298
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1399 | 6.911 | [20 empty] | -0.091942 | 3.8282 | 8.4039 | -12.0405 | 12.8713 | -10.1982 | 5.9858 | -2.2089 | 0.36144 | [12 empty] | 1.2716 | 2.4122 | 4.0181 | -0.77443 | 0.0047434
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1400 | 6.9169 | [20 empty] | -0.091942 | 3.8309 | 8.4159 | -12.0664 | 12.9082 | -10.2347 | 6.0115 | -2.2199 | 0.36352 | [12 empty] | 1.2716 | 2.4139 | 4.0238 | -0.7761 | 0.004757
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1401 | 6.9229 | [20 empty] | -0.091942 | 3.8337 | 8.4279 | -12.0923 | 12.9451 | -10.2714 | 6.0373 | -2.2311 | 0.3656 | [12 empty] | 1.2716 | 2.4157 | 4.0296 | -0.77776 | 0.0047706
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1402 | 6.9289 | [20 empty] | -0.091942 | 3.8364 | 8.4399 | -12.1182 | 12.9821 | -10.3081 | 6.0632 | -2.2422 | 0.36769 | [12 empty] | 1.2716 | 2.4174 | 4.0354 | -0.77943 | 0.0047842
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1403 | 6.9349 | [20 empty] | -0.091942 | 3.8392 | 8.452 | -12.1441 | 13.0191 | -10.3449 | 6.0892 | -2.2535 | 0.36979 | [12 empty] | 1.2716 | 2.4191 | 4.0411 | -0.7811 | 0.0047979
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1404 | 6.9408 | [20 empty] | -0.091942 | 3.8419 | 8.464 | -12.1701 | 13.0563 | -10.3818 | 6.1153 | -2.2647 | 0.37191 | [12 empty] | 1.2716 | 2.4208 | 4.0469 | -0.78277 | 0.0048116
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1405 | 6.9468 | [20 empty] | -0.091942 | 3.8446 | 8.4761 | -12.1961 | 13.0935 | -10.4188 | 6.1414 | -2.276 | 0.37403 | [12 empty] | 1.2716 | 2.4226 | 4.0526 | -0.78444 | 0.0048253
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1406 | 6.9528 | [20 empty] | -0.091942 | 3.8474 | 8.4882 | -12.2222 | 13.1309 | -10.456 | 6.1677 | -2.2874 | 0.37617 | [12 empty] | 1.2716 | 2.4243 | 4.0584 | -0.78612 | 0.004839
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1407 | 6.9588 | [20 empty] | -0.091942 | 3.8501 | 8.5002 | -12.2483 | 13.1683 | -10.4932 | 6.1941 | -2.2988 | 0.37831 | [12 empty] | 1.2716 | 2.426 | 4.0642 | -0.7878 | 0.0048528
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1408 | 6.9647 | [20 empty] | -0.091942 | 3.8528 | 8.5123 | -12.2744 | 13.2057 | -10.5305 | 6.2205 | -2.3103 | 0.38047 | [12 empty] | 1.2716 | 2.4277 | 4.07 | -0.78948 | 0.0048666
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1409 | 6.9707 | [20 empty] | -0.091942 | 3.8556 | 8.5244 | -12.3006 | 13.2433 | -10.568 | 6.2471 | -2.3218 | 0.38264 | [12 empty] | 1.2716 | 2.4294 | 4.0758 | -0.79116 | 0.0048805
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1410 | 6.9767 | [20 empty] | -0.091942 | 3.8583 | 8.5365 | -12.3268 | 13.2809 | -10.6055 | 6.2737 | -2.3333 | 0.38481 | [12 empty] | 1.2716 | 2.4312 | 4.0815 | -0.79285 | 0.0048943
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1411 | 6.9827 | [20 empty] | -0.091942 | 3.861 | 8.5486 | -12.353 | 13.3186 | -10.6432 | 6.3005 | -2.3449 | 0.387 | [12 empty] | 1.2716 | 2.4329 | 4.0873 | -0.79453 | 0.0049082
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1412 | 6.9887 | [20 empty] | -0.091942 | 3.8638 | 8.5608 | -12.3793 | 13.3564 | -10.681 | 6.3273 | -2.3566 | 0.3892 | [12 empty] | 1.2716 | 2.4346 | 4.0931 | -0.79622 | 0.0049222
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1413 | 6.9947 | [20 empty] | -0.091942 | 3.8665 | 8.5729 | -12.4056 | 13.3943 | -10.7188 | 6.3543 | -2.3683 | 0.39141 | [12 empty] | 1.2716 | 2.4363 | 4.0989 | -0.79792 | 0.0049361
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1414 | 7.0006 | [20 empty] | -0.091942 | 3.8693 | 8.585 | -12.432 | 13.4323 | -10.7568 | 6.3813 | -2.3801 | 0.39364 | [12 empty] | 1.2716 | 2.4381 | 4.1047 | -0.79961 | 0.0049501
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1415 | 7.0066 | [20 empty] | -0.091942 | 3.872 | 8.5972 | -12.4584 | 13.4703 | -10.7949 | 6.4084 | -2.3919 | 0.39587 | [12 empty] | 1.2716 | 2.4398 | 4.1105 | -0.80131 | 0.0049641
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1416 | 7.0126 | [20 empty] | -0.091942 | 3.8747 | 8.6093 | -12.4848 | 13.5084 | -10.8331 | 6.4356 | -2.4037 | 0.39811 | [12 empty] | 1.2716 | 2.4415 | 4.1163 | -0.80301 | 0.0049782
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1417 | 7.0186 | [20 empty] | -0.091942 | 3.8775 | 8.6215 | -12.5113 | 13.5466 | -10.8714 | 6.463 | -2.4156 | 0.40037 | [12 empty] | 1.2716 | 2.4432 | 4.1222 | -0.80471 | 0.0049923
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1418 | 7.0246 | [20 empty] | -0.091942 | 3.8802 | 8.6337 | -12.5378 | 13.5849 | -10.9098 | 6.4904 | -2.4276 | 0.40263 | [12 empty] | 1.2716 | 2.445 | 4.128 | -0.80642 | 0.0050064
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1419 | 7.0306 | [20 empty] | -0.091942 | 3.8829 | 8.6459 | -12.5644 | 13.6233 | -10.9484 | 6.5179 | -2.4396 | 0.40491 | [12 empty] | 1.2716 | 2.4467 | 4.1338 | -0.80812 | 0.0050205
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1420 | 7.0366 | [20 empty] | -0.091942 | 3.8857 | 8.658 | -12.5909 | 13.6617 | -10.987 | 6.5455 | -2.4517 | 0.4072 | [12 empty] | 1.2716 | 2.4484 | 4.1396 | -0.80983 | 0.0050347
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1421 | 7.0426 | [20 empty] | -0.091942 | 3.8884 | 8.6702 | -12.6176 | 13.7002 | -11.0257 | 6.5732 | -2.4638 | 0.4095 | [12 empty] | 1.2716 | 2.4501 | 4.1455 | -0.81155 | 0.0050489
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1422 | 7.0486 | [20 empty] | -0.091942 | 3.8911 | 8.6824 | -12.6442 | 13.7388 | -11.0646 | 6.601 | -2.4759 | 0.41181 | [12 empty] | 1.2716 | 2.4519 | 4.1513 | -0.81326 | 0.0050631
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1423 | 7.0546 | [20 empty] | -0.091942 | 3.8939 | 8.6947 | -12.6709 | 13.7775 | -11.1035 | 6.6289 | -2.4882 | 0.41413 | [12 empty] | 1.2716 | 2.4536 | 4.1571 | -0.81498 | 0.0050774
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1424 | 7.0606 | [20 empty] | -0.091942 | 3.8966 | 8.7069 | -12.6976 | 13.8163 | -11.1426 | 6.6569 | -2.5004 | 0.41647 | [12 empty] | 1.2716 | 2.4553 | 4.163 | -0.8167 | 0.0050916
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1425 | 7.0666 | [20 empty] | -0.091942 | 3.8994 | 8.7191 | -12.7244 | 13.8552 | -11.1818 | 6.685 | -2.5127 | 0.41881 | [12 empty] | 1.2716 | 2.457 | 4.1688 | -0.81842 | 0.005106
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1426 | 7.0726 | [20 empty] | -0.091942 | 3.9021 | 8.7314 | -12.7512 | 13.8941 | -11.2211 | 6.7132 | -2.5251 | 0.42117 | [12 empty] | 1.2716 | 2.4588 | 4.1747 | -0.82014 | 0.0051203
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1427 | 7.0786 | [20 empty] | -0.091942 | 3.9048 | 8.7436 | -12.7781 | 13.9331 | -11.2605 | 6.7415 | -2.5375 | 0.42354 | [12 empty] | 1.2716 | 2.4605 | 4.1806 | -0.82187 | 0.0051347
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1428 | 7.0846 | [20 empty] | -0.091942 | 3.9076 | 8.7559 | -12.8049 | 13.9722 | -11.3 | 6.7699 | -2.55 | 0.42592 | [12 empty] | 1.2716 | 2.4622 | 4.1864 | -0.8236 | 0.0051491
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1429 | 7.0906 | [20 empty] | -0.091942 | 3.9103 | 8.7681 | -12.8319 | 14.0114 | -11.3396 | 6.7984 | -2.5625 | 0.42831 | [12 empty] | 1.2716 | 2.4639 | 4.1923 | -0.82533 | 0.0051635
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1430 | 7.0966 | [20 empty] | -0.091942 | 3.913 | 8.7804 | -12.8588 | 14.0506 | -11.3793 | 6.827 | -2.5751 | 0.43071 | [12 empty] | 1.2716 | 2.4657 | 4.1981 | -0.82706 | 0.005178
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1431 | 7.1026 | [20 empty] | -0.091942 | 3.9158 | 8.7927 | -12.8858 | 14.09 | -11.4192 | 6.8557 | -2.5877 | 0.43313 | [12 empty] | 1.2716 | 2.4674 | 4.204 | -0.8288 | 0.0051925
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1432 | 7.1086 | [20 empty] | -0.091942 | 3.9185 | 8.805 | -12.9128 | 14.1294 | -11.4591 | 6.8845 | -2.6004 | 0.43556 | [12 empty] | 1.2716 | 2.4691 | 4.2099 | -0.83054 | 0.005207
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1433 | 7.1146 | [20 empty] | -0.091942 | 3.9212 | 8.8173 | -12.9399 | 14.1689 | -11.4992 | 6.9134 | -2.6132 | 0.43799 | [12 empty] | 1.2716 | 2.4708 | 4.2158 | -0.83228 | 0.0052216
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1434 | 7.1206 | [20 empty] | -0.091942 | 3.924 | 8.8296 | -12.967 | 14.2085 | -11.5394 | 6.9424 | -2.626 | 0.44045 | [12 empty] | 1.2716 | 2.4726 | 4.2217 | -0.83402 | 0.0052362
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1435 | 7.1266 | [20 empty] | -0.091942 | 3.9267 | 8.8419 | -12.9942 | 14.2482 | -11.5797 | 6.9715 | -2.6388 | 0.44291 | [12 empty] | 1.2716 | 2.4743 | 4.2276 | -0.83577 | 0.0052508
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1436 | 7.1326 | [20 empty] | -0.091942 | 3.9295 | 8.8543 | -13.0214 | 14.2879 | -11.6201 | 7.0007 | -2.6517 | 0.44538 | [12 empty] | 1.2716 | 2.476 | 4.2335 | -0.83752 | 0.0052655
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1437 | 7.1386 | [20 empty] | -0.091942 | 3.9322 | 8.8666 | -13.0486 | 14.3278 | -11.6606 | 7.03 | -2.6647 | 0.44787 | [12 empty] | 1.2716 | 2.4777 | 4.2393 | -0.83927 | 0.0052801
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1438 | 7.1446 | [20 empty] | -0.091942 | 3.9349 | 8.8789 | -13.0758 | 14.3677 | -11.7012 | 7.0594 | -2.6777 | 0.45037 | [12 empty] | 1.2716 | 2.4794 | 4.2453 | -0.84102 | 0.0052949
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1439 | 7.1506 | [20 empty] | -0.091942 | 3.9377 | 8.8913 | -13.1031 | 14.4077 | -11.742 | 7.0889 | -2.6907 | 0.45288 | [12 empty] | 1.2716 | 2.4812 | 4.2512 | -0.84278 | 0.0053096
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1440 | 7.1567 | [20 empty] | -0.091942 | 3.9404 | 8.9036 | -13.1305 | 14.4478 | -11.7828 | 7.1185 | -2.7038 | 0.45541 | [12 empty] | 1.2716 | 2.4829 | 4.2571 | -0.84454 | 0.0053244
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1441 | 7.1627 | [20 empty] | -0.091942 | 3.9431 | 8.916 | -13.1578 | 14.488 | -11.8238 | 7.1482 | -2.717 | 0.45794 | [12 empty] | 1.2716 | 2.4846 | 4.263 | -0.8463 | 0.0053392
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1442 | 7.1687 | [20 empty] | -0.091942 | 3.9459 | 8.9284 | -13.1853 | 14.5282 | -11.8649 | 7.178 | -2.7302 | 0.46049 | [12 empty] | 1.2716 | 2.4863 | 4.2689 | -0.84806 | 0.005354
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1443 | 7.1747 | [20 empty] | -0.091942 | 3.9486 | 8.9408 | -13.2127 | 14.5686 | -11.9061 | 7.2079 | -2.7435 | 0.46305 | [12 empty] | 1.2716 | 2.4881 | 4.2748 | -0.84983 | 0.0053689
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1444 | 7.1807 | [20 empty] | -0.091942 | 3.9513 | 8.9532 | -13.2402 | 14.609 | -11.9474 | 7.238 | -2.7569 | 0.46563 | [12 empty] | 1.2716 | 2.4898 | 4.2808 | -0.85159 | 0.0053838
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1445 | 7.1867 | [20 empty] | -0.091942 | 3.9541 | 8.9656 | -13.2677 | 14.6495 | -11.9888 | 7.2681 | -2.7703 | 0.46821 | [12 empty] | 1.2716 | 2.4915 | 4.2867 | -0.85337 | 0.0053987
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1446 | 7.1928 | [20 empty] | -0.091942 | 3.9568 | 8.978 | -13.2953 | 14.6901 | -12.0304 | 7.2983 | -2.7837 | 0.47081 | [12 empty] | 1.2716 | 2.4932 | 4.2926 | -0.85514 | 0.0054137
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1447 | 7.1988 | [20 empty] | -0.091942 | 3.9596 | 8.9904 | -13.3229 | 14.7308 | -12.072 | 7.3286 | -2.7972 | 0.47342 | [12 empty] | 1.2716 | 2.495 | 4.2986 | -0.85691 | 0.0054287
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1448 | 7.2048 | [20 empty] | -0.091942 | 3.9623 | 9.0029 | -13.3505 | 14.7716 | -12.1138 | 7.3591 | -2.8108 | 0.47604 | [12 empty] | 1.2716 | 2.4967 | 4.3045 | -0.85869 | 0.0054437
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1449 | 7.2108 | [20 empty] | -0.091942 | 3.965 | 9.0153 | -13.3782 | 14.8124 | -12.1557 | 7.3896 | -2.8244 | 0.47868 | [12 empty] | 1.2716 | 2.4984 | 4.3104 | -0.86047 | 0.0054587
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1450 | 7.2169 | [20 empty] | -0.091942 | 3.9678 | 9.0277 | -13.4059 | 14.8533 | -12.1977 | 7.4203 | -2.8381 | 0.48133 | [12 empty] | 1.2716 | 2.5001 | 4.3164 | -0.86225 | 0.0054738
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1451 | 7.2229 | [20 empty] | -0.091942 | 3.9705 | 9.0402 | -13.4337 | 14.8944 | -12.2398 | 7.451 | -2.8518 | 0.48399 | [12 empty] | 1.2716 | 2.5019 | 4.3224 | -0.86404 | 0.0054889
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1452 | 7.2289 | [20 empty] | -0.091942 | 3.9732 | 9.0527 | -13.4615 | 14.9355 | -12.282 | 7.4819 | -2.8656 | 0.48667 | [12 empty] | 1.2716 | 2.5036 | 4.3283 | -0.86583 | 0.0055041
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1453 | 7.2349 | [20 empty] | -0.091942 | 3.976 | 9.0651 | -13.4893 | 14.9766 | -12.3244 | 7.5129 | -2.8794 | 0.48935 | [12 empty] | 1.2716 | 2.5053 | 4.3343 | -0.86762 | 0.0055193
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1454 | 7.241 | [20 empty] | -0.091942 | 3.9787 | 9.0776 | -13.5172 | 15.0179 | -12.3668 | 7.544 | -2.8933 | 0.49206 | [12 empty] | 1.2716 | 2.507 | 4.3402 | -0.86941 | 0.0055345
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1455 | 7.247 | [20 empty] | -0.091942 | 3.9814 | 9.0901 | -13.5451 | 15.0593 | -12.4094 | 7.5751 | -2.9073 | 0.49477 | [12 empty] | 1.2716 | 2.5088 | 4.3462 | -0.8712 | 0.0055497
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1456 | 7.253 | [20 empty] | -0.091942 | 3.9842 | 9.1026 | -13.573 | 15.1007 | -12.4521 | 7.6064 | -2.9213 | 0.4975 | [12 empty] | 1.2716 | 2.5105 | 4.3522 | -0.873 | 0.005565
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1457 | 7.2591 | [20 empty] | -0.091942 | 3.9869 | 9.1151 | -13.601 | 15.1422 | -12.495 | 7.6378 | -2.9354 | 0.50024 | [12 empty] | 1.2716 | 2.5122 | 4.3582 | -0.8748 | 0.0055803
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1458 | 7.2651 | [20 empty] | -0.091942 | 3.9897 | 9.1276 | -13.629 | 15.1839 | -12.5379 | 7.6693 | -2.9495 | 0.50299 | [12 empty] | 1.2716 | 2.5139 | 4.3642 | -0.8766 | 0.0055956
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1459 | 7.2711 | [20 empty] | -0.091942 | 3.9924 | 9.1402 | -13.6571 | 15.2256 | -12.5809 | 7.7009 | -2.9637 | 0.50576 | [12 empty] | 1.2716 | 2.5157 | 4.3701 | -0.87841 | 0.005611
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1460 | 7.2772 | [20 empty] | -0.091942 | 3.9951 | 9.1527 | -13.6852 | 15.2673 | -12.6241 | 7.7327 | -2.9779 | 0.50854 | [12 empty] | 1.2716 | 2.5174 | 4.3761 | -0.88022 | 0.0056264
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1461 | 7.2832 | [20 empty] | -0.091942 | 3.9979 | 9.1652 | -13.7133 | 15.3092 | -12.6674 | 7.7645 | -2.9922 | 0.51133 | [12 empty] | 1.2716 | 2.5191 | 4.3821 | -0.88203 | 0.0056418
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1462 | 7.2892 | [20 empty] | -0.091942 | 4.0006 | 9.1778 | -13.7415 | 15.3512 | -12.7108 | 7.7964 | -3.0066 | 0.51414 | [12 empty] | 1.2716 | 2.5208 | 4.3881 | -0.88384 | 0.0056573
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1463 | 7.2953 | [20 empty] | -0.091942 | 4.0033 | 9.1903 | -13.7697 | 15.3932 | -12.7544 | 7.8285 | -3.021 | 0.51696 | [12 empty] | 1.2716 | 2.5226 | 4.3941 | -0.88565 | 0.0056728
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1464 | 7.3013 | [20 empty] | -0.091942 | 4.0061 | 9.2029 | -13.798 | 15.4353 | -12.798 | 7.8607 | -3.0355 | 0.51979 | [12 empty] | 1.2716 | 2.5243 | 4.4002 | -0.88747 | 0.0056883
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1465 | 7.3074 | [20 empty] | -0.091942 | 4.0088 | 9.2155 | -13.8263 | 15.4776 | -12.8418 | 7.8929 | -3.0501 | 0.52264 | [12 empty] | 1.2716 | 2.526 | 4.4062 | -0.88929 | 0.0057039
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1466 | 7.3134 | [20 empty] | -0.091942 | 4.0115 | 9.2281 | -13.8546 | 15.5199 | -12.8857 | 7.9253 | -3.0647 | 0.5255 | [12 empty] | 1.2716 | 2.5277 | 4.4122 | -0.89111 | 0.0057195
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1467 | 7.3194 | [20 empty] | -0.091942 | 4.0143 | 9.2407 | -13.883 | 15.5623 | -12.9297 | 7.9578 | -3.0793 | 0.52837 | [12 empty] | 1.2716 | 2.5295 | 4.4182 | -0.89294 | 0.0057351
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1468 | 7.3255 | [20 empty] | -0.091942 | 4.017 | 9.2533 | -13.9114 | 15.6047 | -12.9738 | 7.9904 | -3.094 | 0.53126 | [12 empty] | 1.2716 | 2.5312 | 4.4242 | -0.89477 | 0.0057507
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1469 | 7.3315 | [20 empty] | -0.091942 | 4.0198 | 9.2659 | -13.9399 | 15.6473 | -13.0181 | 8.0231 | -3.1088 | 0.53416 | [12 empty] | 1.2716 | 2.5329 | 4.4303 | -0.8966 | 0.0057664
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1470 | 7.3376 | [20 empty] | -0.091942 | 4.0225 | 9.2785 | -13.9683 | 15.6899 | -13.0624 | 8.0559 | -3.1237 | 0.53708 | [12 empty] | 1.2716 | 2.5346 | 4.4363 | -0.89843 | 0.0057821
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1471 | 7.3436 | [20 empty] | -0.091942 | 4.0252 | 9.2911 | -13.9969 | 15.7327 | -13.1069 | 8.0889 | -3.1386 | 0.54001 | [12 empty] | 1.2716 | 2.5363 | 4.4423 | -0.90026 | 0.0057979
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1472 | 7.3497 | [20 empty] | -0.091942 | 4.028 | 9.3038 | -14.0254 | 15.7755 | -13.1515 | 8.1219 | -3.1535 | 0.54295 | [12 empty] | 1.2716 | 2.5381 | 4.4484 | -0.9021 | 0.0058137
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1473 | 7.3557 | [20 empty] | -0.091942 | 4.0307 | 9.3164 | -14.054 | 15.8184 | -13.1963 | 8.1551 | -3.1686 | 0.54591 | [12 empty] | 1.2716 | 2.5398 | 4.4544 | -0.90394 | 0.0058295
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1474 | 7.3618 | [20 empty] | -0.091942 | 4.0334 | 9.3291 | -14.0827 | 15.8614 | -13.2411 | 8.1884 | -3.1837 | 0.54888 | [12 empty] | 1.2716 | 2.5415 | 4.4605 | -0.90578 | 0.0058453
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1475 | 7.3678 | [20 empty] | -0.091942 | 4.0362 | 9.3417 | -14.1114 | 15.9045 | -13.2861 | 8.2218 | -3.1988 | 0.55187 | [12 empty] | 1.2716 | 2.5432 | 4.4665 | -0.90763 | 0.0058612
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1476 | 7.3739 | [20 empty] | -0.091942 | 4.0389 | 9.3544 | -14.1401 | 15.9477 | -13.3312 | 8.2553 | -3.214 | 0.55487 | [12 empty] | 1.2716 | 2.545 | 4.4726 | -0.90947 | 0.0058771
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1477 | 7.3799 | [20 empty] | -0.091942 | 4.0416 | 9.3671 | -14.1688 | 15.9909 | -13.3764 | 8.2889 | -3.2293 | 0.55788 | [12 empty] | 1.2716 | 2.5467 | 4.4786 | -0.91132 | 0.0058931
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1478 | 7.386 | [20 empty] | -0.091942 | 4.0444 | 9.3798 | -14.1976 | 16.0343 | -13.4217 | 8.3226 | -3.2446 | 0.56091 | [12 empty] | 1.2716 | 2.5484 | 4.4847 | -0.91318 | 0.005909
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1479 | 7.392 | [20 empty] | -0.091942 | 4.0471 | 9.3925 | -14.2265 | 16.0777 | -13.4672 | 8.3564 | -3.26 | 0.56395 | [12 empty] | 1.2716 | 2.5501 | 4.4908 | -0.91503 | 0.005925
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1480 | 7.3981 | [20 empty] | -0.091942 | 4.0499 | 9.4052 | -14.2554 | 16.1213 | -13.5128 | 8.3904 | -3.2755 | 0.56701 | [12 empty] | 1.2716 | 2.5519 | 4.4969 | -0.91689 | 0.0059411
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1481 | 7.4041 | [20 empty] | -0.091942 | 4.0526 | 9.4179 | -14.2843 | 16.1649 | -13.5585 | 8.4245 | -3.291 | 0.57008 | [12 empty] | 1.2716 | 2.5536 | 4.5029 | -0.91875 | 0.0059572
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1482 | 7.4102 | [20 empty] | -0.091942 | 4.0553 | 9.4306 | -14.3132 | 16.2086 | -13.6044 | 8.4587 | -3.3066 | 0.57317 | [12 empty] | 1.2716 | 2.5553 | 4.509 | -0.92061 | 0.0059733
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1483 | 7.4163 | [20 empty] | -0.091942 | 4.0581 | 9.4433 | -14.3422 | 16.2524 | -13.6503 | 8.493 | -3.3223 | 0.57627 | [12 empty] | 1.2716 | 2.557 | 4.5151 | -0.92248 | 0.0059894
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1484 | 7.4223 | [20 empty] | -0.091942 | 4.0608 | 9.4561 | -14.3713 | 16.2962 | -13.6964 | 8.5274 | -3.338 | 0.57939 | [12 empty] | 1.2716 | 2.5588 | 4.5212 | -0.92434 | 0.0060056
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1485 | 7.4284 | [20 empty] | -0.091942 | 4.0635 | 9.4688 | -14.4003 | 16.3402 | -13.7426 | 8.5619 | -3.3537 | 0.58252 | [12 empty] | 1.2716 | 2.5605 | 4.5273 | -0.92621 | 0.0060218
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1486 | 7.4345 | [20 empty] | -0.091942 | 4.0663 | 9.4816 | -14.4294 | 16.3843 | -13.7889 | 8.5966 | -3.3696 | 0.58567 | [12 empty] | 1.2716 | 2.5622 | 4.5334 | -0.92808 | 0.006038
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1487 | 7.4405 | [20 empty] | -0.091942 | 4.069 | 9.4943 | -14.4586 | 16.4284 | -13.8354 | 8.6313 | -3.3855 | 0.58883 | [12 empty] | 1.2716 | 2.5639 | 4.5395 | -0.92996 | 0.0060543
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1488 | 7.4466 | [20 empty] | -0.091942 | 4.0717 | 9.5071 | -14.4878 | 16.4727 | -13.882 | 8.6662 | -3.4015 | 0.592 | [12 empty] | 1.2716 | 2.5657 | 4.5456 | -0.93184 | 0.0060706
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1489 | 7.4527 | [20 empty] | -0.091942 | 4.0745 | 9.5199 | -14.517 | 16.517 | -13.9287 | 8.7012 | -3.4175 | 0.59519 | [12 empty] | 1.2716 | 2.5674 | 4.5517 | -0.93372 | 0.0060869
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1490 | 7.4587 | [20 empty] | -0.091942 | 4.0772 | 9.5327 | -14.5463 | 16.5614 | -13.9755 | 8.7364 | -3.4336 | 0.5984 | [12 empty] | 1.2716 | 2.5691 | 4.5578 | -0.9356 | 0.0061033
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1491 | 7.4648 | [20 empty] | -0.091942 | 4.08 | 9.5455 | -14.5756 | 16.6059 | -14.0225 | 8.7716 | -3.4498 | 0.60162 | [12 empty] | 1.2716 | 2.5708 | 4.5639 | -0.93748 | 0.0061197
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1492 | 7.4709 | [20 empty] | -0.091942 | 4.0827 | 9.5583 | -14.6049 | 16.6505 | -14.0696 | 8.8069 | -3.466 | 0.60485 | [12 empty] | 1.2716 | 2.5726 | 4.5701 | -0.93937 | 0.0061361
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1493 | 7.4769 | [20 empty] | -0.091942 | 4.0854 | 9.5711 | -14.6343 | 16.6952 | -14.1168 | 8.8424 | -3.4823 | 0.6081 | [12 empty] | 1.2716 | 2.5743 | 4.5762 | -0.94126 | 0.0061526
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1494 | 7.483 | [20 empty] | -0.091942 | 4.0882 | 9.5839 | -14.6637 | 16.74 | -14.1641 | 8.878 | -3.4986 | 0.61137 | [12 empty] | 1.2716 | 2.576 | 4.5823 | -0.94316 | 0.0061691
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1495 | 7.4891 | [20 empty] | -0.091942 | 4.0909 | 9.5968 | -14.6932 | 16.7848 | -14.2116 | 8.9137 | -3.5151 | 0.61465 | [12 empty] | 1.2716 | 2.5777 | 4.5885 | -0.94505 | 0.0061856
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1496 | 7.4952 | [20 empty] | -0.091942 | 4.0936 | 9.6096 | -14.7227 | 16.8298 | -14.2592 | 8.9496 | -3.5316 | 0.61795 | [12 empty] | 1.2716 | 2.5795 | 4.5946 | -0.94695 | 0.0062022
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1497 | 7.5012 | [20 empty] | -0.091942 | 4.0964 | 9.6225 | -14.7523 | 16.8748 | -14.3069 | 8.9855 | -3.5481 | 0.62126 | [12 empty] | 1.2716 | 2.5812 | 4.6008 | -0.94885 | 0.0062188
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1498 | 7.5073 | [20 empty] | -0.091942 | 4.0991 | 9.6353 | -14.7818 | 16.92 | -14.3548 | 9.0216 | -3.5647 | 0.62459 | [12 empty] | 1.2716 | 2.5829 | 4.6069 | -0.95075 | 0.0062354
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1499 | 7.5134 | [20 empty] | -0.091942 | 4.1018 | 9.6482 | -14.8115 | 16.9652 | -14.4027 | 9.0578 | -3.5814 | 0.62793 | [12 empty] | 1.2716 | 2.5846 | 4.6131 | -0.95266 | 0.0062521
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1500 | 7.5195 | [20 empty] | -0.091942 | 4.1046 | 9.6611 | -14.8411 | 17.0105 | -14.4508 | 9.0941 | -3.5982 | 0.63129 | [12 empty] | 1.2716 | 2.5864 | 4.6192 | -0.95456 | 0.0062688
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1501 | 7.5255 | [20 empty] | -0.091942 | 4.1073 | 9.674 | -14.8708 | 17.0559 | -14.4991 | 9.1305 | -3.615 | 0.63466 | [12 empty] | 1.2716 | 2.5881 | 4.6254 | -0.95647 | 0.0062855
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1502 | 7.5316 | [20 empty] | -0.091942 | 4.1101 | 9.6869 | -14.9006 | 17.1014 | -14.5474 | 9.1671 | -3.6319 | 0.63806 | [12 empty] | 1.2716 | 2.5898 | 4.6315 | -0.95839 | 0.0063023
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1503 | 7.5377 | [20 empty] | -0.091942 | 4.1128 | 9.6998 | -14.9303 | 17.147 | -14.5959 | 9.2038 | -3.6489 | 0.64146 | [12 empty] | 1.2716 | 2.5915 | 4.6377 | -0.9603 | 0.0063191
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1504 | 7.5438 | [20 empty] | -0.091942 | 4.1155 | 9.7127 | -14.9602 | 17.1927 | -14.6446 | 9.2406 | -3.6659 | 0.64488 | [12 empty] | 1.2716 | 2.5932 | 4.6439 | -0.96222 | 0.0063359
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1505 | 7.5499 | [20 empty] | -0.091942 | 4.1183 | 9.7256 | -14.99 | 17.2384 | -14.6933 | 9.2775 | -3.683 | 0.64832 | [12 empty] | 1.2716 | 2.595 | 4.6501 | -0.96414 | 0.0063528
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1506 | 7.556 | [20 empty] | -0.091942 | 4.121 | 9.7385 | -15.0199 | 17.2843 | -14.7422 | 9.3146 | -3.7001 | 0.65178 | [12 empty] | 1.2716 | 2.5967 | 4.6562 | -0.96606 | 0.0063697
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1507 | 7.562 | [20 empty] | -0.091942 | 4.1237 | 9.7515 | -15.0499 | 17.3303 | -14.7912 | 9.3517 | -3.7174 | 0.65525 | [12 empty] | 1.2716 | 2.5984 | 4.6624 | -0.96799 | 0.0063866
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1508 | 7.5681 | [20 empty] | -0.091942 | 4.1265 | 9.7644 | -15.0798 | 17.3763 | -14.8403 | 9.389 | -3.7347 | 0.65873 | [12 empty] | 1.2716 | 2.6001 | 4.6686 | -0.96992 | 0.0064036
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1509 | 7.5742 | [20 empty] | -0.091942 | 4.1292 | 9.7774 | -15.1099 | 17.4224 | -14.8896 | 9.4264 | -3.7521 | 0.66224 | [12 empty] | 1.2716 | 2.6019 | 4.6748 | -0.97185 | 0.0064206
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1510 | 7.5803 | [20 empty] | -0.091942 | 4.1319 | 9.7903 | -15.1399 | 17.4687 | -14.939 | 9.464 | -3.7695 | 0.66576 | [12 empty] | 1.2716 | 2.6036 | 4.681 | -0.97378 | 0.0064376
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1511 | 7.5864 | [20 empty] | -0.091942 | 4.1347 | 9.8033 | -15.17 | 17.515 | -14.9885 | 9.5017 | -3.787 | 0.66929 | [12 empty] | 1.2716 | 2.6053 | 4.6872 | -0.97572 | 0.0064547
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1512 | 7.5925 | [20 empty] | -0.091942 | 4.1374 | 9.8163 | -15.2002 | 17.5614 | -15.0382 | 9.5394 | -3.8046 | 0.67284 | [12 empty] | 1.2716 | 2.607 | 4.6934 | -0.97766 | 0.0064718
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1513 | 7.5986 | [20 empty] | -0.091942 | 4.1402 | 9.8293 | -15.2303 | 17.6079 | -15.088 | 9.5774 | -3.8222 | 0.67641 | [12 empty] | 1.2716 | 2.6088 | 4.6996 | -0.9796 | 0.0064889
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1514 | 7.6047 | [20 empty] | -0.091942 | 4.1429 | 9.8423 | -15.2606 | 17.6545 | -15.1379 | 9.6154 | -3.8399 | 0.68 | [12 empty] | 1.2716 | 2.6105 | 4.7058 | -0.98154 | 0.0065061
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1515 | 7.6108 | [20 empty] | -0.091942 | 4.1456 | 9.8553 | -15.2908 | 17.7012 | -15.188 | 9.6536 | -3.8577 | 0.6836 | [12 empty] | 1.2716 | 2.6122 | 4.7121 | -0.98349 | 0.0065233
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1516 | 7.6169 | [20 empty] | -0.091942 | 4.1484 | 9.8683 | -15.3211 | 17.748 | -15.2382 | 9.6919 | -3.8756 | 0.68721 | [12 empty] | 1.2716 | 2.6139 | 4.7183 | -0.98544 | 0.0065406
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1517 | 7.623 | [20 empty] | -0.091942 | 4.1511 | 9.8813 | -15.3515 | 17.7948 | -15.2885 | 9.7303 | -3.8935 | 0.69085 | [12 empty] | 1.2716 | 2.6157 | 4.7245 | -0.98739 | 0.0065578
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1518 | 7.6291 | [20 empty] | -0.091942 | 4.1538 | 9.8943 | -15.3818 | 17.8418 | -15.339 | 9.7688 | -3.9115 | 0.6945 | [12 empty] | 1.2716 | 2.6174 | 4.7307 | -0.98934 | 0.0065751
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1519 | 7.6352 | [20 empty] | -0.091942 | 4.1566 | 9.9074 | -15.4123 | 17.8889 | -15.3895 | 9.8075 | -3.9296 | 0.69817 | [12 empty] | 1.2716 | 2.6191 | 4.737 | -0.9913 | 0.0065925
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1520 | 7.6413 | [20 empty] | -0.091942 | 4.1593 | 9.9204 | -15.4427 | 17.936 | -15.4403 | 9.8463 | -3.9477 | 0.70186 | [12 empty] | 1.2716 | 2.6208 | 4.7432 | -0.99326 | 0.0066099
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1521 | 7.6474 | [20 empty] | -0.091942 | 4.162 | 9.9335 | -15.4732 | 17.9833 | -15.4911 | 9.8853 | -3.966 | 0.70556 | [12 empty] | 1.2716 | 2.6226 | 4.7495 | -0.99522 | 0.0066273
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1522 | 7.6535 | [20 empty] | -0.091942 | 4.1648 | 9.9465 | -15.5038 | 18.0306 | -15.5421 | 9.9243 | -3.9843 | 0.70928 | [12 empty] | 1.2716 | 2.6243 | 4.7557 | -0.99718 | 0.0066447
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1523 | 7.6596 | [20 empty] | -0.091942 | 4.1675 | 9.9596 | -15.5343 | 18.078 | -15.5932 | 9.9635 | -4.0026 | 0.71301 | [12 empty] | 1.2716 | 2.626 | 4.762 | -0.99915 | 0.0066622
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1524 | 7.6657 | [20 empty] | -0.091942 | 4.1703 | 9.9727 | -15.565 | 18.1256 | -15.6445 | 10.0028 | -4.021 | 0.71677 | [12 empty] | 1.2716 | 2.6277 | 4.7682 | -1.0011 | 0.0066797
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1525 | 7.6718 | [20 empty] | -0.091942 | 4.173 | 9.9858 | -15.5956 | 18.1732 | -15.6959 | 10.0423 | -4.0396 | 0.72054 | [12 empty] | 1.2716 | 2.6295 | 4.7745 | -1.0031 | 0.0066973
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1526 | 7.6779 | [20 empty] | -0.091942 | 4.1757 | 9.9989 | -15.6263 | 18.2209 | -15.7474 | 10.0818 | -4.0581 | 0.72433 | [12 empty] | 1.2716 | 2.6312 | 4.7807 | -1.0051 | 0.0067149
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1527 | 7.684 | [20 empty] | -0.091942 | 4.1785 | 10.012 | -15.6571 | 18.2687 | -15.7991 | 10.1215 | -4.0768 | 0.72813 | [12 empty] | 1.2716 | 2.6329 | 4.787 | -1.007 | 0.0067325
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1528 | 7.6901 | [20 empty] | -0.091942 | 4.1812 | 10.0251 | -15.6878 | 18.3166 | -15.8509 | 10.1614 | -4.0955 | 0.73196 | [12 empty] | 1.2716 | 2.6346 | 4.7933 | -1.009 | 0.0067501
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1529 | 7.6962 | [20 empty] | -0.091942 | 4.1839 | 10.0382 | -15.7187 | 18.3646 | -15.9028 | 10.2013 | -4.1143 | 0.7358 | [12 empty] | 1.2716 | 2.6364 | 4.7996 | -1.011 | 0.0067678
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1530 | 7.7023 | [20 empty] | -0.091942 | 4.1867 | 10.0514 | -15.7495 | 18.4127 | -15.9549 | 10.2414 | -4.1332 | 0.73966 | [12 empty] | 1.2716 | 2.6381 | 4.8058 | -1.013 | 0.0067855
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1531 | 7.7084 | [20 empty] | -0.091942 | 4.1894 | 10.0645 | -15.7804 | 18.4609 | -16.0071 | 10.2817 | -4.1521 | 0.74353 | [12 empty] | 1.2716 | 2.6398 | 4.8121 | -1.015 | 0.0068033
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1532 | 7.7145 | [20 empty] | -0.091942 | 4.1921 | 10.0777 | -15.8114 | 18.5092 | -16.0595 | 10.322 | -4.1711 | 0.74743 | [12 empty] | 1.2716 | 2.6415 | 4.8184 | -1.017 | 0.0068211
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1533 | 7.7207 | [20 empty] | -0.091942 | 4.1949 | 10.0908 | -15.8423 | 18.5575 | -16.1119 | 10.3625 | -4.1902 | 0.75134 | [12 empty] | 1.2716 | 2.6433 | 4.8247 | -1.019 | 0.0068389
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1534 | 7.7268 | [20 empty] | -0.091942 | 4.1976 | 10.104 | -15.8734 | 18.606 | -16.1646 | 10.4031 | -4.2094 | 0.75527 | [12 empty] | 1.2716 | 2.645 | 4.831 | -1.021 | 0.0068568
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1535 | 7.7329 | [20 empty] | -0.091942 | 4.2004 | 10.1172 | -15.9044 | 18.6546 | -16.2173 | 10.4439 | -4.2287 | 0.75922 | [12 empty] | 1.2716 | 2.6467 | 4.8373 | -1.023 | 0.0068747
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1536 | 7.739 | [20 empty] | -0.091942 | 4.2031 | 10.1304 | -15.9355 | 18.7032 | -16.2702 | 10.4848 | -4.248 | 0.76318 | [12 empty] | 1.2716 | 2.6484 | 4.8436 | -1.025 | 0.0068926
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1537 | 7.7451 | [20 empty] | -0.091942 | 4.2058 | 10.1436 | -15.9667 | 18.752 | -16.3232 | 10.5258 | -4.2674 | 0.76717 | [12 empty] | 1.2716 | 2.6501 | 4.8499 | -1.027 | 0.0069106
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1538 | 7.7513 | [20 empty] | -0.091942 | 4.2086 | 10.1568 | -15.9979 | 18.8008 | -16.3764 | 10.567 | -4.2869 | 0.77117 | [12 empty] | 1.2716 | 2.6519 | 4.8562 | -1.029 | 0.0069286
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1539 | 7.7574 | [20 empty] | -0.091942 | 4.2113 | 10.17 | -16.0291 | 18.8498 | -16.4297 | 10.6083 | -4.3064 | 0.77519 | [12 empty] | 1.2716 | 2.6536 | 4.8625 | -1.031 | 0.0069466
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1540 | 7.7635 | [20 empty] | -0.091942 | 4.214 | 10.1832 | -16.0604 | 18.8988 | -16.4832 | 10.6497 | -4.326 | 0.77923 | [12 empty] | 1.2716 | 2.6553 | 4.8689 | -1.033 | 0.0069647
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1541 | 7.7696 | [20 empty] | -0.091942 | 4.2168 | 10.1964 | -16.0917 | 18.948 | -16.5367 | 10.6912 | -4.3457 | 0.78329 | [12 empty] | 1.2716 | 2.657 | 4.8752 | -1.035 | 0.0069828
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1542 | 7.7757 | [20 empty] | -0.091942 | 4.2195 | 10.2097 | -16.123 | 18.9972 | -16.5905 | 10.7329 | -4.3655 | 0.78736 | [12 empty] | 1.2716 | 2.6588 | 4.8815 | -1.037 | 0.0070009
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1543 | 7.7819 | [20 empty] | -0.091942 | 4.2222 | 10.2229 | -16.1544 | 19.0465 | -16.6443 | 10.7748 | -4.3854 | 0.79146 | [12 empty] | 1.2716 | 2.6605 | 4.8878 | -1.039 | 0.0070191
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1544 | 7.788 | [20 empty] | -0.091942 | 4.225 | 10.2362 | -16.1858 | 19.0959 | -16.6983 | 10.8167 | -4.4053 | 0.79557 | [12 empty] | 1.2716 | 2.6622 | 4.8942 | -1.0411 | 0.0070373
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1545 | 7.7941 | [20 empty] | -0.091942 | 4.2277 | 10.2494 | -16.2173 | 19.1455 | -16.7525 | 10.8588 | -4.4253 | 0.7997 | [12 empty] | 1.2716 | 2.6639 | 4.9005 | -1.0431 | 0.0070556
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1546 | 7.8002 | [20 empty] | -0.091942 | 4.2305 | 10.2627 | -16.2488 | 19.1951 | -16.8068 | 10.9011 | -4.4454 | 0.80385 | [12 empty] | 1.2716 | 2.6657 | 4.9069 | -1.0451 | 0.0070739
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1547 | 7.8064 | [20 empty] | -0.091942 | 4.2332 | 10.276 | -16.2804 | 19.2448 | -16.8612 | 10.9435 | -4.4656 | 0.80802 | [12 empty] | 1.2716 | 2.6674 | 4.9132 | -1.0471 | 0.0070922
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1548 | 7.8125 | [20 empty] | -0.091942 | 4.2359 | 10.2893 | -16.3119 | 19.2946 | -16.9158 | 10.986 | -4.4858 | 0.81221 | [12 empty] | 1.2716 | 2.6691 | 4.9196 | -1.0492 | 0.0071105
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1549 | 7.8186 | [20 empty] | -0.091942 | 4.2387 | 10.3026 | -16.3436 | 19.3445 | -16.9705 | 11.0286 | -4.5061 | 0.81641 | [12 empty] | 1.2716 | 2.6708 | 4.9259 | -1.0512 | 0.0071289
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1550 | 7.8248 | [20 empty] | -0.091942 | 4.2414 | 10.3159 | -16.3753 | 19.3945 | -17.0253 | 11.0714 | -4.5265 | 0.82064 | [12 empty] | 1.2716 | 2.6726 | 4.9323 | -1.0532 | 0.0071474
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1551 | 7.8309 | [20 empty] | -0.091942 | 4.2441 | 10.3292 | -16.407 | 19.4446 | -17.0803 | 11.1143 | -4.547 | 0.82488 | [12 empty] | 1.2716 | 2.6743 | 4.9387 | -1.0553 | 0.0071658
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1552 | 7.837 | [20 empty] | -0.091942 | 4.2469 | 10.3425 | -16.4387 | 19.4948 | -17.1355 | 11.1574 | -4.5676 | 0.82915 | [12 empty] | 1.2716 | 2.676 | 4.945 | -1.0573 | 0.0071843
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1553 | 7.8432 | [20 empty] | -0.091942 | 4.2496 | 10.3559 | -16.4705 | 19.5451 | -17.1907 | 11.2006 | -4.5882 | 0.83343 | [12 empty] | 1.2716 | 2.6777 | 4.9514 | -1.0594 | 0.0072028
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1554 | 7.8493 | [20 empty] | -0.091942 | 4.2523 | 10.3692 | -16.5024 | 19.5955 | -17.2461 | 11.2439 | -4.6089 | 0.83774 | [12 empty] | 1.2716 | 2.6795 | 4.9578 | -1.0614 | 0.0072214
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1555 | 7.8555 | [20 empty] | -0.091942 | 4.2551 | 10.3825 | -16.5342 | 19.646 | -17.3017 | 11.2874 | -4.6297 | 0.84206 | [12 empty] | 1.2716 | 2.6812 | 4.9642 | -1.0635 | 0.00724
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1556 | 7.8616 | [20 empty] | -0.091942 | 4.2578 | 10.3959 | -16.5662 | 19.6966 | -17.3574 | 11.331 | -4.6506 | 0.8464 | [12 empty] | 1.2716 | 2.6829 | 4.9706 | -1.0655 | 0.0072587
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1557 | 7.8677 | [20 empty] | -0.091942 | 4.2606 | 10.4093 | -16.5981 | 19.7472 | -17.4133 | 11.3748 | -4.6716 | 0.85076 | [12 empty] | 1.2716 | 2.6846 | 4.9769 | -1.0676 | 0.0072773
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1558 | 7.8739 | [20 empty] | -0.091942 | 4.2633 | 10.4226 | -16.6301 | 19.798 | -17.4692 | 11.4187 | -4.6926 | 0.85514 | [12 empty] | 1.2716 | 2.6864 | 4.9833 | -1.0696 | 0.0072961
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1559 | 7.88 | [20 empty] | -0.091942 | 4.266 | 10.436 | -16.6622 | 19.8489 | -17.5254 | 11.4628 | -4.7138 | 0.85954 | [12 empty] | 1.2716 | 2.6881 | 4.9897 | -1.0717 | 0.0073148
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1560 | 7.8862 | [20 empty] | -0.091942 | 4.2688 | 10.4494 | -16.6942 | 19.8999 | -17.5817 | 11.5069 | -4.735 | 0.86396 | [12 empty] | 1.2716 | 2.6898 | 4.9961 | -1.0738 | 0.0073336
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1561 | 7.8923 | [20 empty] | -0.091942 | 4.2715 | 10.4628 | -16.7264 | 19.9509 | -17.6381 | 11.5513 | -4.7562 | 0.8684 | [12 empty] | 1.2716 | 2.6915 | 5.0025 | -1.0758 | 0.0073524
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1562 | 7.8984 | [20 empty] | -0.091942 | 4.2742 | 10.4762 | -16.7585 | 20.0021 | -17.6947 | 11.5957 | -4.7776 | 0.87287 | [12 empty] | 1.2716 | 2.6933 | 5.009 | -1.0779 | 0.0073713
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1563 | 7.9046 | [20 empty] | -0.091942 | 4.277 | 10.4896 | -16.7907 | 20.0534 | -17.7514 | 11.6404 | -4.7991 | 0.87735 | [12 empty] | 1.2716 | 2.695 | 5.0154 | -1.08 | 0.0073902
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1564 | 7.9107 | [20 empty] | -0.091942 | 4.2797 | 10.5031 | -16.823 | 20.1048 | -17.8082 | 11.6851 | -4.8206 | 0.88185 | [12 empty] | 1.2716 | 2.6967 | 5.0218 | -1.082 | 0.0074091
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1565 | 7.9169 | [20 empty] | -0.091942 | 4.2824 | 10.5165 | -16.8553 | 20.1562 | -17.8652 | 11.73 | -4.8422 | 0.88637 | [12 empty] | 1.2716 | 2.6984 | 5.0282 | -1.0841 | 0.0074281
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1566 | 7.923 | [20 empty] | -0.091942 | 4.2852 | 10.53 | -16.8876 | 20.2078 | -17.9224 | 11.7751 | -4.8639 | 0.89091 | [12 empty] | 1.2716 | 2.7002 | 5.0346 | -1.0862 | 0.0074471
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1567 | 7.9292 | [20 empty] | -0.091942 | 4.2879 | 10.5434 | -16.92 | 20.2595 | -17.9797 | 11.8202 | -4.8857 | 0.89547 | [12 empty] | 1.2716 | 2.7019 | 5.0411 | -1.0883 | 0.0074661
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1568 | 7.9353 | [20 empty] | -0.091942 | 4.2907 | 10.5569 | -16.9524 | 20.3112 | -18.0371 | 11.8656 | -4.9076 | 0.90005 | [12 empty] | 1.2716 | 2.7036 | 5.0475 | -1.0904 | 0.0074852
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1569 | 7.9415 | [20 empty] | -0.091942 | 4.2934 | 10.5703 | -16.9849 | 20.3631 | -18.0947 | 11.9111 | -4.9295 | 0.90465 | [12 empty] | 1.2716 | 2.7053 | 5.054 | -1.0924 | 0.0075043
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1570 | 7.9477 | [20 empty] | -0.091942 | 4.2961 | 10.5838 | -17.0173 | 20.4151 | -18.1524 | 11.9567 | -4.9516 | 0.90928 | [12 empty] | 1.2716 | 2.707 | 5.0604 | -1.0945 | 0.0075235
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1571 | 7.9538 | [20 empty] | -0.091942 | 4.2989 | 10.5973 | -17.0499 | 20.4671 | -18.2103 | 12.0024 | -4.9737 | 0.91392 | [12 empty] | 1.2716 | 2.7088 | 5.0668 | -1.0966 | 0.0075426
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1572 | 7.96 | [20 empty] | -0.091942 | 4.3016 | 10.6108 | -17.0825 | 20.5193 | -18.2684 | 12.0484 | -4.9959 | 0.91858 | [12 empty] | 1.2716 | 2.7105 | 5.0733 | -1.0987 | 0.0075619
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1573 | 7.9661 | [20 empty] | -0.091942 | 4.3043 | 10.6243 | -17.1151 | 20.5715 | -18.3265 | 12.0944 | -5.0182 | 0.92327 | [12 empty] | 1.2716 | 2.7122 | 5.0798 | -1.1008 | 0.0075811
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1574 | 7.9723 | [20 empty] | -0.091942 | 4.3071 | 10.6378 | -17.1477 | 20.6239 | -18.3849 | 12.1406 | -5.0405 | 0.92798 | [12 empty] | 1.2716 | 2.7139 | 5.0862 | -1.1029 | 0.0076004
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1575 | 7.9784 | [20 empty] | -0.091942 | 4.3098 | 10.6513 | -17.1805 | 20.6764 | -18.4433 | 12.187 | -5.063 | 0.9327 | [12 empty] | 1.2716 | 2.7157 | 5.0927 | -1.105 | 0.0076198
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1576 | 7.9846 | [20 empty] | -0.091942 | 4.3125 | 10.6649 | -17.2132 | 20.7289 | -18.502 | 12.2335 | -5.0855 | 0.93745 | [12 empty] | 1.2716 | 2.7174 | 5.0992 | -1.1071 | 0.0076391
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1577 | 7.9908 | [20 empty] | -0.091942 | 4.3153 | 10.6784 | -17.246 | 20.7816 | -18.5607 | 12.2801 | -5.1082 | 0.94222 | [12 empty] | 1.2716 | 2.7191 | 5.1056 | -1.1092 | 0.0076585
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1578 | 7.9969 | [20 empty] | -0.091942 | 4.318 | 10.692 | -17.2788 | 20.8343 | -18.6197 | 12.3269 | -5.1309 | 0.94701 | [12 empty] | 1.2716 | 2.7208 | 5.1121 | -1.1114 | 0.007678
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1579 | 8.0031 | [20 empty] | -0.091942 | 4.3208 | 10.7055 | -17.3117 | 20.8872 | -18.6787 | 12.3739 | -5.1537 | 0.95182 | [12 empty] | 1.2716 | 2.7226 | 5.1186 | -1.1135 | 0.0076975
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1580 | 8.0093 | [20 empty] | -0.091942 | 4.3235 | 10.7191 | -17.3446 | 20.9402 | -18.738 | 12.421 | -5.1766 | 0.95666 | [12 empty] | 1.2716 | 2.7243 | 5.1251 | -1.1156 | 0.007717
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1581 | 8.0154 | [20 empty] | -0.091942 | 4.3262 | 10.7326 | -17.3775 | 20.9932 | -18.7973 | 12.4682 | -5.1996 | 0.96151 | [12 empty] | 1.2716 | 2.726 | 5.1316 | -1.1177 | 0.0077365
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1582 | 8.0216 | [20 empty] | -0.091942 | 4.329 | 10.7462 | -17.4105 | 21.0464 | -18.8569 | 12.5156 | -5.2226 | 0.96639 | [12 empty] | 1.2716 | 2.7277 | 5.1381 | -1.1198 | 0.0077561
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1583 | 8.0278 | [20 empty] | -0.091942 | 4.3317 | 10.7598 | -17.4436 | 21.0997 | -18.9165 | 12.5631 | -5.2458 | 0.97128 | [12 empty] | 1.2716 | 2.7295 | 5.1446 | -1.122 | 0.0077757
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1584 | 8.0339 | [20 empty] | -0.091942 | 4.3344 | 10.7734 | -17.4767 | 21.153 | -18.9764 | 12.6108 | -5.269 | 0.9762 | [12 empty] | 1.2716 | 2.7312 | 5.1511 | -1.1241 | 0.0077954
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1585 | 8.0401 | [20 empty] | -0.091942 | 4.3372 | 10.787 | -17.5098 | 21.2065 | -19.0363 | 12.6587 | -5.2924 | 0.98114 | [12 empty] | 1.2716 | 2.7329 | 5.1576 | -1.1262 | 0.0078151
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1586 | 8.0463 | [20 empty] | -0.091942 | 4.3399 | 10.8006 | -17.5429 | 21.2601 | -19.0965 | 12.7067 | -5.3158 | 0.98611 | [12 empty] | 1.2716 | 2.7346 | 5.1641 | -1.1283 | 0.0078349
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1587 | 8.0524 | [20 empty] | -0.091942 | 4.3426 | 10.8143 | -17.5761 | 21.3137 | -19.1567 | 12.7548 | -5.3393 | 0.99109 | [12 empty] | 1.2716 | 2.7364 | 5.1706 | -1.1305 | 0.0078546
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1588 | 8.0586 | [20 empty] | -0.091942 | 4.3454 | 10.8279 | -17.6094 | 21.3675 | -19.2172 | 12.8031 | -5.3629 | 0.9961 | [12 empty] | 1.2716 | 2.7381 | 5.1771 | -1.1326 | 0.0078745
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1589 | 8.0648 | [20 empty] | -0.091942 | 4.3481 | 10.8415 | -17.6427 | 21.4214 | -19.2778 | 12.8516 | -5.3866 | 1.0011 | [12 empty] | 1.2716 | 2.7398 | 5.1836 | -1.1348 | 0.0078943
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1590 | 8.071 | [20 empty] | -0.091942 | 4.3509 | 10.8552 | -17.676 | 21.4754 | -19.3385 | 12.9002 | -5.4103 | 1.0062 | [12 empty] | 1.2716 | 2.7415 | 5.1901 | -1.1369 | 0.0079142
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1591 | 8.0771 | [20 empty] | -0.091942 | 4.3536 | 10.8688 | -17.7094 | 21.5294 | -19.3994 | 12.9489 | -5.4342 | 1.0113 | [12 empty] | 1.2716 | 2.7433 | 5.1967 | -1.139 | 0.0079341
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1592 | 8.0833 | [20 empty] | -0.091942 | 4.3563 | 10.8825 | -17.7428 | 21.5836 | -19.4604 | 12.9978 | -5.4582 | 1.0164 | [12 empty] | 1.2716 | 2.745 | 5.2032 | -1.1412 | 0.0079541
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1593 | 8.0895 | [20 empty] | -0.091942 | 4.3591 | 10.8962 | -17.7763 | 21.6379 | -19.5216 | 13.0469 | -5.4822 | 1.0215 | [12 empty] | 1.2716 | 2.7467 | 5.2098 | -1.1433 | 0.0079741
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1594 | 8.0957 | [20 empty] | -0.091942 | 4.3618 | 10.9099 | -17.8098 | 21.6923 | -19.583 | 13.0961 | -5.5063 | 1.0266 | [12 empty] | 1.2716 | 2.7484 | 5.2163 | -1.1455 | 0.0079941
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1595 | 8.1019 | [20 empty] | -0.091942 | 4.3645 | 10.9236 | -17.8433 | 21.7468 | -19.6445 | 13.1455 | -5.5306 | 1.0318 | [12 empty] | 1.2716 | 2.7502 | 5.2228 | -1.1477 | 0.0080142
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1596 | 8.108 | [20 empty] | -0.091942 | 4.3673 | 10.9373 | -17.8769 | 21.8014 | -19.7061 | 13.195 | -5.5549 | 1.037 | [12 empty] | 1.2716 | 2.7519 | 5.2294 | -1.1498 | 0.0080343
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1597 | 8.1142 | [20 empty] | -0.091942 | 4.37 | 10.951 | -17.9105 | 21.856 | -19.7679 | 13.2447 | -5.5793 | 1.0422 | [12 empty] | 1.2716 | 2.7536 | 5.2359 | -1.152 | 0.0080545
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1598 | 8.1204 | [20 empty] | -0.091942 | 4.3727 | 10.9647 | -17.9442 | 21.9108 | -19.8299 | 13.2945 | -5.6038 | 1.0474 | [12 empty] | 1.2716 | 2.7553 | 5.2425 | -1.1541 | 0.0080747
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1599 | 8.1266 | [20 empty] | -0.091942 | 4.3755 | 10.9784 | -17.9779 | 21.9657 | -19.892 | 13.3445 | -5.6284 | 1.0527 | [12 empty] | 1.2716 | 2.7571 | 5.2491 | -1.1563 | 0.0080949
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1600 | 8.1328 | [20 empty] | -0.091942 | 4.3782 | 10.9922 | -18.0116 | 22.0207 | -19.9543 | 13.3947 | -5.6531 | 1.0579 | [12 empty] | 1.2716 | 2.7588 | 5.2556 | -1.1585 | 0.0081152
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1601 | 8.139 | [20 empty] | -0.091942 | 4.381 | 11.0059 | -18.0454 | 22.0758 | -20.0168 | 13.445 | -5.6778 | 1.0632 | [12 empty] | 1.2716 | 2.7605 | 5.2622 | -1.1607 | 0.0081355
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1602 | 8.1452 | [20 empty] | -0.091942 | 4.3837 | 11.0197 | -18.0793 | 22.131 | -20.0793 | 13.4955 | -5.7027 | 1.0686 | [12 empty] | 1.2716 | 2.7622 | 5.2688 | -1.1628 | 0.0081558
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1603 | 8.1513 | [20 empty] | -0.091942 | 4.3864 | 11.0334 | -18.1131 | 22.1864 | -20.1421 | 13.5461 | -5.7277 | 1.0739 | [12 empty] | 1.2716 | 2.7639 | 5.2754 | -1.165 | 0.0081762
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1604 | 8.1575 | [20 empty] | -0.091942 | 4.3892 | 11.0472 | -18.147 | 22.2418 | -20.205 | 13.5969 | -5.7527 | 1.0793 | [12 empty] | 1.2716 | 2.7657 | 5.282 | -1.1672 | 0.0081966
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1605 | 8.1637 | [20 empty] | -0.091942 | 4.3919 | 11.061 | -18.181 | 22.2973 | -20.2681 | 13.6478 | -5.7779 | 1.0847 | [12 empty] | 1.2716 | 2.7674 | 5.2885 | -1.1694 | 0.0082171
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1606 | 8.1699 | [20 empty] | -0.091942 | 4.3946 | 11.0748 | -18.215 | 22.3529 | -20.3313 | 13.6989 | -5.8031 | 1.0901 | [12 empty] | 1.2716 | 2.7691 | 5.2951 | -1.1716 | 0.0082376
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1607 | 8.1761 | [20 empty] | -0.091942 | 4.3974 | 11.0885 | -18.2491 | 22.4086 | -20.3947 | 13.7502 | -5.8285 | 1.0955 | [12 empty] | 1.2716 | 2.7708 | 5.3017 | -1.1738 | 0.0082581
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1608 | 8.1823 | [20 empty] | -0.091942 | 4.4001 | 11.1024 | -18.2832 | 22.4645 | -20.4582 | 13.8016 | -5.8539 | 1.101 | [12 empty] | 1.2716 | 2.7726 | 5.3083 | -1.176 | 0.0082787
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1609 | 8.1885 | [20 empty] | -0.091942 | 4.4028 | 11.1162 | -18.3173 | 22.5204 | -20.5219 | 13.8532 | -5.8794 | 1.1065 | [12 empty] | 1.2716 | 2.7743 | 5.3149 | -1.1781 | 0.0082993
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1610 | 8.1947 | [20 empty] | -0.091942 | 4.4056 | 11.13 | -18.3515 | 22.5764 | -20.5857 | 13.9049 | -5.9051 | 1.112 | [12 empty] | 1.2716 | 2.776 | 5.3215 | -1.1803 | 0.00832
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1611 | 8.2009 | [20 empty] | -0.091942 | 4.4083 | 11.1438 | -18.3857 | 22.6326 | -20.6497 | 13.9568 | -5.9308 | 1.1175 | [12 empty] | 1.2716 | 2.7777 | 5.3282 | -1.1825 | 0.0083407
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1612 | 8.2071 | [20 empty] | -0.091942 | 4.4111 | 11.1577 | -18.4199 | 22.6888 | -20.7139 | 14.0089 | -5.9566 | 1.1231 | [12 empty] | 1.2716 | 2.7795 | 5.3348 | -1.1847 | 0.0083614
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1613 | 8.2133 | [20 empty] | -0.091942 | 4.4138 | 11.1715 | -18.4542 | 22.7452 | -20.7782 | 14.0611 | -5.9825 | 1.1287 | [12 empty] | 1.2716 | 2.7812 | 5.3414 | -1.187 | 0.0083822
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1614 | 8.2195 | [20 empty] | -0.091942 | 4.4165 | 11.1854 | -18.4886 | 22.8016 | -20.8427 | 14.1135 | -6.0085 | 1.1343 | [12 empty] | 1.2716 | 2.7829 | 5.348 | -1.1892 | 0.008403
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1615 | 8.2257 | [20 empty] | -0.091942 | 4.4193 | 11.1992 | -18.523 | 22.8582 | -20.9074 | 14.166 | -6.0346 | 1.1399 | [12 empty] | 1.2716 | 2.7846 | 5.3546 | -1.1914 | 0.0084238
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1616 | 8.2319 | [20 empty] | -0.091942 | 4.422 | 11.2131 | -18.5574 | 22.9149 | -20.9722 | 14.2187 | -6.0609 | 1.1456 | [12 empty] | 1.2716 | 2.7864 | 5.3613 | -1.1936 | 0.0084447
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1617 | 8.2381 | [20 empty] | -0.091942 | 4.4247 | 11.227 | -18.5919 | 22.9716 | -21.0372 | 14.2716 | -6.0872 | 1.1513 | [12 empty] | 1.2716 | 2.7881 | 5.3679 | -1.1958 | 0.0084656
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1618 | 8.2443 | [20 empty] | -0.091942 | 4.4275 | 11.2409 | -18.6264 | 23.0285 | -21.1023 | 14.3246 | -6.1136 | 1.157 | [12 empty] | 1.2716 | 2.7898 | 5.3746 | -1.198 | 0.0084866
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1619 | 8.2505 | [20 empty] | -0.091942 | 4.4302 | 11.2548 | -18.6609 | 23.0855 | -21.1676 | 14.3778 | -6.1401 | 1.1627 | [12 empty] | 1.2716 | 2.7915 | 5.3812 | -1.2003 | 0.0085076
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1620 | 8.2567 | [20 empty] | -0.091942 | 4.4329 | 11.2687 | -18.6955 | 23.1426 | -21.233 | 14.4312 | -6.1667 | 1.1685 | [12 empty] | 1.2716 | 2.7933 | 5.3879 | -1.2025 | 0.0085286
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1621 | 8.2629 | [20 empty] | -0.091942 | 4.4357 | 11.2826 | -18.7302 | 23.1998 | -21.2986 | 14.4847 | -6.1933 | 1.1743 | [12 empty] | 1.2716 | 2.795 | 5.3945 | -1.2047 | 0.0085497
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1622 | 8.2692 | [20 empty] | -0.091942 | 4.4384 | 11.2965 | -18.7649 | 23.2571 | -21.3644 | 14.5384 | -6.2201 | 1.1801 | [12 empty] | 1.2716 | 2.7967 | 5.4012 | -1.2069 | 0.0085708
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1623 | 8.2754 | [20 empty] | -0.091942 | 4.4412 | 11.3105 | -18.7996 | 23.3145 | -21.4304 | 14.5923 | -6.247 | 1.1859 | [12 empty] | 1.2716 | 2.7984 | 5.4078 | -1.2092 | 0.008592
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1624 | 8.2816 | [20 empty] | -0.091942 | 4.4439 | 11.3244 | -18.8344 | 23.372 | -21.4965 | 14.6463 | -6.274 | 1.1918 | [12 empty] | 1.2716 | 2.8002 | 5.4145 | -1.2114 | 0.0086132
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1625 | 8.2878 | [20 empty] | -0.091942 | 4.4466 | 11.3383 | -18.8692 | 23.4296 | -21.5627 | 14.7005 | -6.3011 | 1.1976 | [12 empty] | 1.2716 | 2.8019 | 5.4212 | -1.2136 | 0.0086344
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1626 | 8.294 | [20 empty] | -0.091942 | 4.4494 | 11.3523 | -18.904 | 23.4874 | -21.6292 | 14.7549 | -6.3283 | 1.2035 | [12 empty] | 1.2716 | 2.8036 | 5.4278 | -1.2159 | 0.0086557
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1627 | 8.3002 | [20 empty] | -0.091942 | 4.4521 | 11.3663 | -18.9389 | 23.5452 | -21.6958 | 14.8094 | -6.3556 | 1.2095 | [12 empty] | 1.2716 | 2.8053 | 5.4345 | -1.2181 | 0.008677
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1628 | 8.3064 | [20 empty] | -0.091942 | 4.4548 | 11.3802 | -18.9739 | 23.6031 | -21.7625 | 14.8641 | -6.383 | 1.2154 | [12 empty] | 1.2716 | 2.8071 | 5.4412 | -1.2204 | 0.0086983
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1629 | 8.3126 | [20 empty] | -0.091942 | 4.4576 | 11.3942 | -19.0089 | 23.6612 | -21.8294 | 14.919 | -6.4105 | 1.2214 | [12 empty] | 1.2716 | 2.8088 | 5.4479 | -1.2226 | 0.0087197
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1630 | 8.3189 | [20 empty] | -0.091942 | 4.4603 | 11.4082 | -19.0439 | 23.7193 | -21.8965 | 14.974 | -6.4381 | 1.2274 | [12 empty] | 1.2716 | 2.8105 | 5.4546 | -1.2249 | 0.0087412
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1631 | 8.3251 | [20 empty] | -0.091942 | 4.463 | 11.4222 | -19.079 | 23.7776 | -21.9638 | 15.0292 | -6.4658 | 1.2335 | [12 empty] | 1.2716 | 2.8122 | 5.4613 | -1.2271 | 0.0087626
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1632 | 8.3313 | [20 empty] | -0.091942 | 4.4658 | 11.4362 | -19.1141 | 23.836 | -22.0312 | 15.0846 | -6.4936 | 1.2395 | [12 empty] | 1.2716 | 2.814 | 5.468 | -1.2294 | 0.0087841
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1633 | 8.3375 | [20 empty] | -0.091942 | 4.4685 | 11.4503 | -19.1492 | 23.8944 | -22.0988 | 15.1401 | -6.5215 | 1.2456 | [12 empty] | 1.2716 | 2.8157 | 5.4747 | -1.2317 | 0.0088057
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1634 | 8.3438 | [20 empty] | -0.091942 | 4.4713 | 11.4643 | -19.1844 | 23.953 | -22.1665 | 15.1958 | -6.5495 | 1.2517 | [12 empty] | 1.2716 | 2.8174 | 5.4814 | -1.2339 | 0.0088273
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1635 | 8.35 | [20 empty] | -0.091942 | 4.474 | 11.4783 | -19.2197 | 24.0117 | -22.2344 | 15.2517 | -6.5776 | 1.2579 | [12 empty] | 1.2716 | 2.8191 | 5.4881 | -1.2362 | 0.0088489
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1636 | 8.3562 | [20 empty] | -0.091942 | 4.4767 | 11.4924 | -19.255 | 24.0705 | -22.3025 | 15.3078 | -6.6058 | 1.2641 | [12 empty] | 1.2716 | 2.8208 | 5.4948 | -1.2385 | 0.0088706
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1637 | 8.3624 | [20 empty] | -0.091942 | 4.4795 | 11.5064 | -19.2903 | 24.1294 | -22.3707 | 15.364 | -6.6341 | 1.2702 | [12 empty] | 1.2716 | 2.8226 | 5.5015 | -1.2407 | 0.0088923
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1638 | 8.3687 | [20 empty] | -0.091942 | 4.4822 | 11.5205 | -19.3257 | 24.1884 | -22.4392 | 15.4204 | -6.6626 | 1.2765 | [12 empty] | 1.2716 | 2.8243 | 5.5082 | -1.243 | 0.008914
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1639 | 8.3749 | [20 empty] | -0.091942 | 4.4849 | 11.5346 | -19.3611 | 24.2475 | -22.5077 | 15.477 | -6.6911 | 1.2827 | [12 empty] | 1.2716 | 2.826 | 5.515 | -1.2453 | 0.0089358
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1640 | 8.3811 | [20 empty] | -0.091942 | 4.4877 | 11.5486 | -19.3965 | 24.3068 | -22.5765 | 15.5337 | -6.7197 | 1.289 | [12 empty] | 1.2716 | 2.8277 | 5.5217 | -1.2476 | 0.0089576
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1641 | 8.3873 | [20 empty] | -0.091942 | 4.4904 | 11.5627 | -19.4321 | 24.3661 | -22.6454 | 15.5906 | -6.7485 | 1.2953 | [12 empty] | 1.2716 | 2.8295 | 5.5284 | -1.2498 | 0.0089795
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1642 | 8.3936 | [20 empty] | -0.091942 | 4.4931 | 11.5768 | -19.4676 | 24.4256 | -22.7145 | 15.6477 | -6.7773 | 1.3016 | [12 empty] | 1.2716 | 2.8312 | 5.5352 | -1.2521 | 0.0090014
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1643 | 8.3998 | [20 empty] | -0.091942 | 4.4959 | 11.5909 | -19.5032 | 24.4851 | -22.7837 | 15.705 | -6.8062 | 1.308 | [12 empty] | 1.2716 | 2.8329 | 5.5419 | -1.2544 | 0.0090234
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1644 | 8.406 | [20 empty] | -0.091942 | 4.4986 | 11.605 | -19.5388 | 24.5448 | -22.8531 | 15.7624 | -6.8353 | 1.3144 | [12 empty] | 1.2716 | 2.8346 | 5.5487 | -1.2567 | 0.0090454
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1645 | 8.4123 | [20 empty] | -0.091942 | 4.5014 | 11.6192 | -19.5745 | 24.6046 | -22.9227 | 15.8201 | -6.8644 | 1.3208 | [12 empty] | 1.2716 | 2.8364 | 5.5554 | -1.259 | 0.0090674
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1646 | 8.4185 | [20 empty] | -0.091942 | 4.5041 | 11.6333 | -19.6102 | 24.6644 | -22.9925 | 15.8778 | -6.8937 | 1.3272 | [12 empty] | 1.2716 | 2.8381 | 5.5622 | -1.2613 | 0.0090895
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1647 | 8.4247 | [20 empty] | -0.091942 | 4.5068 | 11.6474 | -19.646 | 24.7244 | -23.0624 | 15.9358 | -6.9231 | 1.3337 | [12 empty] | 1.2716 | 2.8398 | 5.5689 | -1.2636 | 0.0091116
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1648 | 8.431 | [20 empty] | -0.091942 | 4.5096 | 11.6616 | -19.6818 | 24.7845 | -23.1325 | 15.994 | -6.9526 | 1.3402 | [12 empty] | 1.2716 | 2.8415 | 5.5757 | -1.2659 | 0.0091337
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1649 | 8.4372 | [20 empty] | -0.091942 | 4.5123 | 11.6757 | -19.7176 | 24.8447 | -23.2028 | 16.0523 | -6.9821 | 1.3467 | [12 empty] | 1.2716 | 2.8433 | 5.5825 | -1.2682 | 0.0091559
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1650 | 8.4435 | [20 empty] | -0.091942 | 4.515 | 11.6899 | -19.7535 | 24.9051 | -23.2732 | 16.1108 | -7.0118 | 1.3532 | [12 empty] | 1.2716 | 2.845 | 5.5892 | -1.2705 | 0.0091781
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1651 | 8.4497 | [20 empty] | -0.091942 | 4.5178 | 11.7041 | -19.7895 | 24.9655 | -23.3438 | 16.1694 | -7.0416 | 1.3598 | [12 empty] | 1.2716 | 2.8467 | 5.596 | -1.2728 | 0.0092004
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1652 | 8.4559 | [20 empty] | -0.091942 | 4.5205 | 11.7183 | -19.8254 | 25.026 | -23.4146 | 16.2283 | -7.0716 | 1.3664 | [12 empty] | 1.2716 | 2.8484 | 5.6028 | -1.2752 | 0.0092227
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1653 | 8.4622 | [20 empty] | -0.091942 | 4.5232 | 11.7324 | -19.8615 | 25.0867 | -23.4856 | 16.2873 | -7.1016 | 1.373 | [12 empty] | 1.2716 | 2.8502 | 5.6096 | -1.2775 | 0.0092451
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1654 | 8.4684 | [20 empty] | -0.091942 | 4.526 | 11.7466 | -19.8975 | 25.1475 | -23.5567 | 16.3465 | -7.1317 | 1.3797 | [12 empty] | 1.2716 | 2.8519 | 5.6164 | -1.2798 | 0.0092675
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1655 | 8.4747 | [20 empty] | -0.091942 | 4.5287 | 11.7609 | -19.9337 | 25.2083 | -23.628 | 16.4059 | -7.1619 | 1.3864 | [12 empty] | 1.2716 | 2.8536 | 5.6232 | -1.2821 | 0.0092899
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1656 | 8.4809 | [20 empty] | -0.091942 | 4.5315 | 11.7751 | -19.9698 | 25.2693 | -23.6995 | 16.4655 | -7.1923 | 1.3931 | [12 empty] | 1.2716 | 2.8553 | 5.63 | -1.2844 | 0.0093124
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1657 | 8.4872 | [20 empty] | -0.091942 | 4.5342 | 11.7893 | -20.006 | 25.3304 | -23.7711 | 16.5252 | -7.2227 | 1.3998 | [12 empty] | 1.2716 | 2.8571 | 5.6368 | -1.2868 | 0.0093349
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1658 | 8.4934 | [20 empty] | -0.091942 | 4.5369 | 11.8035 | -20.0423 | 25.3916 | -23.8429 | 16.5852 | -7.2533 | 1.4066 | [12 empty] | 1.2716 | 2.8588 | 5.6436 | -1.2891 | 0.0093574
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1659 | 8.4997 | [20 empty] | -0.091942 | 4.5397 | 11.8178 | -20.0785 | 25.4529 | -23.9149 | 16.6453 | -7.284 | 1.4134 | [12 empty] | 1.2716 | 2.8605 | 5.6504 | -1.2914 | 0.00938
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1660 | 8.5059 | [20 empty] | -0.091942 | 4.5424 | 11.832 | -20.1149 | 25.5143 | -23.9871 | 16.7056 | -7.3148 | 1.4202 | [12 empty] | 1.2716 | 2.8622 | 5.6572 | -1.2938 | 0.0094027
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1661 | 8.5122 | [20 empty] | -0.091942 | 4.5451 | 11.8463 | -20.1512 | 25.5759 | -24.0594 | 16.766 | -7.3457 | 1.4271 | [12 empty] | 1.2716 | 2.864 | 5.664 | -1.2961 | 0.0094253
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1662 | 8.5184 | [20 empty] | -0.091942 | 4.5479 | 11.8606 | -20.1877 | 25.6375 | -24.1319 | 16.8267 | -7.3767 | 1.434 | [12 empty] | 1.2716 | 2.8657 | 5.6708 | -1.2984 | 0.0094481
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1663 | 8.5247 | [20 empty] | -0.091942 | 4.5506 | 11.8748 | -20.2241 | 25.6993 | -24.2046 | 16.8875 | -7.4078 | 1.4409 | [12 empty] | 1.2716 | 2.8674 | 5.6777 | -1.3008 | 0.0094708
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1664 | 8.5309 | [20 empty] | -0.091942 | 4.5533 | 11.8891 | -20.2606 | 25.7611 | -24.2775 | 16.9485 | -7.4391 | 1.4479 | [12 empty] | 1.2716 | 2.8691 | 5.6845 | -1.3031 | 0.0094936
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1665 | 8.5372 | [20 empty] | -0.091942 | 4.5561 | 11.9034 | -20.2972 | 25.8231 | -24.3505 | 17.0098 | -7.4704 | 1.4548 | [12 empty] | 1.2716 | 2.8709 | 5.6913 | -1.3055 | 0.0095165
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1666 | 8.5434 | [20 empty] | -0.091942 | 4.5588 | 11.9177 | -20.3338 | 25.8852 | -24.4237 | 17.0711 | -7.5019 | 1.4618 | [12 empty] | 1.2716 | 2.8726 | 5.6982 | -1.3078 | 0.0095393
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1667 | 8.5497 | [20 empty] | -0.091942 | 4.5616 | 11.932 | -20.3704 | 25.9474 | -24.4971 | 17.1327 | -7.5334 | 1.4689 | [12 empty] | 1.2716 | 2.8743 | 5.705 | -1.3102 | 0.0095623
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1668 | 8.5559 | [20 empty] | -0.091942 | 4.5643 | 11.9463 | -20.4071 | 26.0097 | -24.5707 | 17.1945 | -7.5651 | 1.4759 | [12 empty] | 1.2716 | 2.876 | 5.7119 | -1.3126 | 0.0095852
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1669 | 8.5622 | [20 empty] | -0.091942 | 4.567 | 11.9607 | -20.4438 | 26.0722 | -24.6444 | 17.2564 | -7.5969 | 1.483 | [12 empty] | 1.2716 | 2.8777 | 5.7187 | -1.3149 | 0.0096082
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1670 | 8.5685 | [20 empty] | -0.091942 | 4.5698 | 11.975 | -20.4806 | 26.1347 | -24.7183 | 17.3185 | -7.6289 | 1.4902 | [12 empty] | 1.2716 | 2.8795 | 5.7256 | -1.3173 | 0.0096313
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1671 | 8.5747 | [20 empty] | -0.091942 | 4.5725 | 11.9894 | -20.5174 | 26.1974 | -24.7924 | 17.3809 | -7.6609 | 1.4973 | [12 empty] | 1.2716 | 2.8812 | 5.7324 | -1.3197 | 0.0096544
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1672 | 8.581 | [20 empty] | -0.091942 | 4.5752 | 12.0037 | -20.5543 | 26.2601 | -24.8667 | 17.4434 | -7.693 | 1.5045 | [12 empty] | 1.2716 | 2.8829 | 5.7393 | -1.322 | 0.0096775
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1673 | 8.5873 | [20 empty] | -0.091942 | 4.578 | 12.0181 | -20.5912 | 26.323 | -24.9412 | 17.5061 | -7.7253 | 1.5117 | [12 empty] | 1.2716 | 2.8846 | 5.7462 | -1.3244 | 0.0097007
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1674 | 8.5935 | [20 empty] | -0.091942 | 4.5807 | 12.0324 | -20.6281 | 26.386 | -25.0158 | 17.5689 | -7.7577 | 1.5189 | [12 empty] | 1.2716 | 2.8864 | 5.753 | -1.3268 | 0.0097239
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1675 | 8.5998 | [20 empty] | -0.091942 | 4.5834 | 12.0468 | -20.6651 | 26.4491 | -25.0906 | 17.632 | -7.7902 | 1.5262 | [12 empty] | 1.2716 | 2.8881 | 5.7599 | -1.3292 | 0.0097472
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1676 | 8.606 | [20 empty] | -0.091942 | 4.5862 | 12.0612 | -20.7021 | 26.5123 | -25.1656 | 17.6952 | -7.8228 | 1.5335 | [12 empty] | 1.2716 | 2.8898 | 5.7668 | -1.3315 | 0.0097704
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1677 | 8.6123 | [20 empty] | -0.091942 | 4.5889 | 12.0756 | -20.7392 | 26.5757 | -25.2408 | 17.7587 | -7.8555 | 1.5409 | [12 empty] | 1.2716 | 2.8915 | 5.7737 | -1.3339 | 0.0097938
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1678 | 8.6186 | [20 empty] | -0.091942 | 4.5917 | 12.09 | -20.7763 | 26.6391 | -25.3161 | 17.8223 | -7.8884 | 1.5482 | [12 empty] | 1.2716 | 2.8933 | 5.7806 | -1.3363 | 0.0098172
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1679 | 8.6249 | [20 empty] | -0.091942 | 4.5944 | 12.1044 | -20.8135 | 26.7027 | -25.3916 | 17.8861 | -7.9214 | 1.5556 | [12 empty] | 1.2716 | 2.895 | 5.7874 | -1.3387 | 0.0098406
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1680 | 8.6311 | [20 empty] | -0.091942 | 4.5971 | 12.1189 | -20.8507 | 26.7663 | -25.4673 | 17.9502 | -7.9544 | 1.5631 | [12 empty] | 1.2716 | 2.8967 | 5.7943 | -1.3411 | 0.0098641
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1681 | 8.6374 | [20 empty] | -0.091942 | 4.5999 | 12.1333 | -20.888 | 26.8301 | -25.5432 | 18.0144 | -7.9876 | 1.5705 | [12 empty] | 1.2716 | 2.8984 | 5.8012 | -1.3435 | 0.0098876
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1682 | 8.6437 | [20 empty] | -0.091942 | 4.6026 | 12.1477 | -20.9253 | 26.894 | -25.6193 | 18.0788 | -8.021 | 1.578 | [12 empty] | 1.2716 | 2.9002 | 5.8081 | -1.3459 | 0.0099111
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1683 | 8.6499 | [20 empty] | -0.091942 | 4.6053 | 12.1622 | -20.9626 | 26.958 | -25.6955 | 18.1433 | -8.0544 | 1.5855 | [12 empty] | 1.2716 | 2.9019 | 5.8151 | -1.3483 | 0.0099347
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1684 | 8.6562 | [20 empty] | -0.091942 | 4.6081 | 12.1766 | -21 | 27.0222 | -25.772 | 18.2081 | -8.088 | 1.5931 | [12 empty] | 1.2716 | 2.9036 | 5.822 | -1.3507 | 0.0099583
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1685 | 8.6625 | [20 empty] | -0.091942 | 4.6108 | 12.1911 | -21.0374 | 27.0864 | -25.8486 | 18.2731 | -8.1216 | 1.6007 | [12 empty] | 1.2716 | 2.9053 | 5.8289 | -1.3531 | 0.009982
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1686 | 8.6688 | [20 empty] | -0.091942 | 4.6135 | 12.2056 | -21.0749 | 27.1508 | -25.9254 | 18.3383 | -8.1554 | 1.6083 | [12 empty] | 1.2716 | 2.9071 | 5.8358 | -1.3555 | 0.010006
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1687 | 8.675 | [20 empty] | -0.091942 | 4.6163 | 12.2201 | -21.1124 | 27.2152 | -26.0023 | 18.4036 | -8.1894 | 1.6159 | [12 empty] | 1.2716 | 2.9088 | 5.8427 | -1.3579 | 0.010029
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1688 | 8.6813 | [20 empty] | -0.091942 | 4.619 | 12.2345 | -21.15 | 27.2798 | -26.0795 | 18.4692 | -8.2234 | 1.6236 | [12 empty] | 1.2716 | 2.9105 | 5.8497 | -1.3603 | 0.010053
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1689 | 8.6876 | [20 empty] | -0.091942 | 4.6218 | 12.249 | -21.1876 | 27.3445 | -26.1568 | 18.5349 | -8.2576 | 1.6313 | [12 empty] | 1.2716 | 2.9122 | 5.8566 | -1.3628 | 0.010077
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1690 | 8.6939 | [20 empty] | -0.091942 | 4.6245 | 12.2636 | -21.2253 | 27.4094 | -26.2344 | 18.6008 | -8.2918 | 1.6391 | [12 empty] | 1.2716 | 2.914 | 5.8635 | -1.3652 | 0.010101
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1691 | 8.7002 | [20 empty] | -0.091942 | 4.6272 | 12.2781 | -21.263 | 27.4743 | -26.3121 | 18.667 | -8.3263 | 1.6468 | [12 empty] | 1.2716 | 2.9157 | 5.8705 | -1.3676 | 0.010125
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1692 | 8.7065 | [20 empty] | -0.091942 | 4.63 | 12.2926 | -21.3007 | 27.5393 | -26.39 | 18.7333 | -8.3608 | 1.6546 | [12 empty] | 1.2716 | 2.9174 | 5.8774 | -1.37 | 0.010149
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1693 | 8.7127 | [20 empty] | -0.091942 | 4.6327 | 12.3071 | -21.3385 | 27.6045 | -26.468 | 18.7998 | -8.3954 | 1.6625 | [12 empty] | 1.2716 | 2.9191 | 5.8844 | -1.3725 | 0.010173
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1694 | 8.719 | [20 empty] | -0.091942 | 4.6354 | 12.3217 | -21.3763 | 27.6698 | -26.5463 | 18.8666 | -8.4302 | 1.6703 | [12 empty] | 1.2716 | 2.9209 | 5.8913 | -1.3749 | 0.010197
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1695 | 8.7253 | [20 empty] | -0.091942 | 4.6382 | 12.3362 | -21.4142 | 27.7352 | -26.6247 | 18.9335 | -8.4651 | 1.6782 | [12 empty] | 1.2716 | 2.9226 | 5.8983 | -1.3773 | 0.010221
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1696 | 8.7316 | [20 empty] | -0.091942 | 4.6409 | 12.3508 | -21.4521 | 27.8007 | -26.7034 | 19.0006 | -8.5001 | 1.6862 | [12 empty] | 1.2716 | 2.9243 | 5.9052 | -1.3798 | 0.010245
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1697 | 8.7379 | [20 empty] | -0.091942 | 4.6436 | 12.3654 | -21.4901 | 27.8663 | -26.7822 | 19.0679 | -8.5353 | 1.6942 | [12 empty] | 1.2716 | 2.926 | 5.9122 | -1.3822 | 0.010269
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1698 | 8.7442 | [20 empty] | -0.091942 | 4.6464 | 12.3799 | -21.5281 | 27.932 | -26.8612 | 19.1354 | -8.5705 | 1.7022 | [12 empty] | 1.2716 | 2.9278 | 5.9192 | -1.3847 | 0.010294
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1699 | 8.7505 | [20 empty] | -0.091942 | 4.6491 | 12.3945 | -21.5662 | 27.9979 | -26.9404 | 19.2032 | -8.6059 | 1.7102 | [12 empty] | 1.2716 | 2.9295 | 5.9261 | -1.3871 | 0.010318
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1700 | 8.7568 | [20 empty] | -0.091942 | 4.6519 | 12.4091 | -21.6043 | 28.0639 | -27.0198 | 19.2711 | -8.6415 | 1.7183 | [12 empty] | 1.2716 | 2.9312 | 5.9331 | -1.3896 | 0.010342
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1701 | 8.7631 | [20 empty] | -0.091942 | 4.6546 | 12.4237 | -21.6424 | 28.13 | -27.0993 | 19.3392 | -8.6771 | 1.7264 | [12 empty] | 1.2716 | 2.9329 | 5.9401 | -1.392 | 0.010367
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1702 | 8.7693 | [20 empty] | -0.091942 | 4.6573 | 12.4383 | -21.6806 | 28.1962 | -27.1791 | 19.4075 | -8.7129 | 1.7345 | [12 empty] | 1.2716 | 2.9346 | 5.9471 | -1.3945 | 0.010391
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1703 | 8.7756 | [20 empty] | -0.091942 | 4.6601 | 12.4529 | -21.7189 | 28.2625 | -27.259 | 19.476 | -8.7488 | 1.7427 | [12 empty] | 1.2716 | 2.9364 | 5.9541 | -1.3969 | 0.010415
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1704 | 8.7819 | [20 empty] | -0.091942 | 4.6628 | 12.4676 | -21.7571 | 28.3289 | -27.3391 | 19.5447 | -8.7848 | 1.7509 | [12 empty] | 1.2716 | 2.9381 | 5.9611 | -1.3994 | 0.01044
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1705 | 8.7882 | [20 empty] | -0.091942 | 4.6655 | 12.4822 | -21.7955 | 28.3955 | -27.4195 | 19.6137 | -8.8209 | 1.7591 | [12 empty] | 1.2716 | 2.9398 | 5.9681 | -1.4019 | 0.010464
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1706 | 8.7945 | [20 empty] | -0.091942 | 4.6683 | 12.4969 | -21.8338 | 28.4622 | -27.5 | 19.6828 | -8.8572 | 1.7674 | [12 empty] | 1.2716 | 2.9415 | 5.9751 | -1.4043 | 0.010489
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1707 | 8.8008 | [20 empty] | -0.091942 | 4.671 | 12.5115 | -21.8723 | 28.529 | -27.5807 | 19.7521 | -8.8936 | 1.7757 | [12 empty] | 1.2716 | 2.9433 | 5.9821 | -1.4068 | 0.010514
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1708 | 8.8071 | [20 empty] | -0.091942 | 4.6737 | 12.5262 | -21.9107 | 28.5959 | -27.6615 | 19.8216 | -8.9302 | 1.784 | [12 empty] | 1.2716 | 2.945 | 5.9891 | -1.4093 | 0.010538
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1709 | 8.8134 | [20 empty] | -0.091942 | 4.6765 | 12.5409 | -21.9492 | 28.6629 | -27.7426 | 19.8914 | -8.9668 | 1.7924 | [12 empty] | 1.2716 | 2.9467 | 5.9961 | -1.4117 | 0.010563
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1710 | 8.8197 | [20 empty] | -0.091942 | 4.6792 | 12.5555 | -21.9878 | 28.7301 | -27.8239 | 19.9613 | -9.0036 | 1.8008 | [12 empty] | 1.2716 | 2.9484 | 6.0031 | -1.4142 | 0.010588
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1711 | 8.826 | [20 empty] | -0.091942 | 4.682 | 12.5702 | -22.0264 | 28.7973 | -27.9053 | 20.0315 | -9.0405 | 1.8093 | [12 empty] | 1.2716 | 2.9502 | 6.0102 | -1.4167 | 0.010613
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1712 | 8.8323 | [20 empty] | -0.091942 | 4.6847 | 12.5849 | -22.065 | 28.8647 | -27.987 | 20.1018 | -9.0776 | 1.8177 | [12 empty] | 1.2716 | 2.9519 | 6.0172 | -1.4192 | 0.010637
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1713 | 8.8386 | [20 empty] | -0.091942 | 4.6874 | 12.5996 | -22.1037 | 28.9322 | -28.0688 | 20.1724 | -9.1148 | 1.8262 | [12 empty] | 1.2716 | 2.9536 | 6.0242 | -1.4217 | 0.010662
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1714 | 8.8449 | [20 empty] | -0.091942 | 4.6902 | 12.6143 | -22.1424 | 28.9998 | -28.1508 | 20.2431 | -9.1521 | 1.8348 | [12 empty] | 1.2716 | 2.9553 | 6.0312 | -1.4242 | 0.010687
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1715 | 8.8513 | [20 empty] | -0.091942 | 4.6929 | 12.6291 | -22.1812 | 29.0676 | -28.233 | 20.3141 | -9.1895 | 1.8434 | [12 empty] | 1.2716 | 2.9571 | 6.0383 | -1.4267 | 0.010712
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1716 | 8.8576 | [20 empty] | -0.091942 | 4.6956 | 12.6438 | -22.22 | 29.1354 | -28.3154 | 20.3853 | -9.2271 | 1.852 | [12 empty] | 1.2716 | 2.9588 | 6.0453 | -1.4292 | 0.010737
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1717 | 8.8639 | [20 empty] | -0.091942 | 4.6984 | 12.6585 | -22.2589 | 29.2034 | -28.398 | 20.4566 | -9.2648 | 1.8606 | [12 empty] | 1.2716 | 2.9605 | 6.0524 | -1.4317 | 0.010762
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1718 | 8.8702 | [20 empty] | -0.091942 | 4.7011 | 12.6733 | -22.2978 | 29.2715 | -28.4808 | 20.5282 | -9.3026 | 1.8693 | [12 empty] | 1.2716 | 2.9622 | 6.0594 | -1.4342 | 0.010787
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1719 | 8.8765 | [20 empty] | -0.091942 | 4.7038 | 12.688 | -22.3368 | 29.3397 | -28.5638 | 20.6 | -9.3406 | 1.878 | [12 empty] | 1.2716 | 2.964 | 6.0665 | -1.4367 | 0.010812
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1720 | 8.8828 | [20 empty] | -0.091942 | 4.7066 | 12.7028 | -22.3758 | 29.408 | -28.647 | 20.672 | -9.3787 | 1.8868 | [12 empty] | 1.2716 | 2.9657 | 6.0735 | -1.4392 | 0.010838
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1721 | 8.8891 | [20 empty] | -0.091942 | 4.7093 | 12.7176 | -22.4148 | 29.4765 | -28.7304 | 20.7443 | -9.417 | 1.8956 | [12 empty] | 1.2716 | 2.9674 | 6.0806 | -1.4417 | 0.010863
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1722 | 8.8954 | [20 empty] | -0.091942 | 4.7121 | 12.7324 | -22.4539 | 29.545 | -28.8139 | 20.8167 | -9.4553 | 1.9044 | [12 empty] | 1.2716 | 2.9691 | 6.0877 | -1.4442 | 0.010888
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1723 | 8.9017 | [20 empty] | -0.091942 | 4.7148 | 12.7472 | -22.4931 | 29.6137 | -28.8977 | 20.8893 | -9.4938 | 1.9133 | [12 empty] | 1.2716 | 2.9709 | 6.0948 | -1.4467 | 0.010913
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1724 | 8.9081 | [20 empty] | -0.091942 | 4.7175 | 12.762 | -22.5323 | 29.6825 | -28.9817 | 20.9622 | -9.5325 | 1.9222 | [12 empty] | 1.2716 | 2.9726 | 6.1018 | -1.4492 | 0.010939
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1725 | 8.9144 | [20 empty] | -0.091942 | 4.7203 | 12.7768 | -22.5715 | 29.7515 | -29.0658 | 21.0352 | -9.5712 | 1.9311 | [12 empty] | 1.2716 | 2.9743 | 6.1089 | -1.4518 | 0.010964
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1726 | 8.9207 | [20 empty] | -0.091942 | 4.723 | 12.7916 | -22.6108 | 29.8205 | -29.1502 | 21.1085 | -9.6101 | 1.9401 | [12 empty] | 1.2716 | 2.976 | 6.116 | -1.4543 | 0.01099
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1727 | 8.927 | [20 empty] | -0.091942 | 4.7257 | 12.8064 | -22.6501 | 29.8897 | -29.2347 | 21.182 | -9.6492 | 1.9491 | [12 empty] | 1.2716 | 2.9778 | 6.1231 | -1.4568 | 0.011015
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1728 | 8.9333 | [20 empty] | -0.091942 | 4.7285 | 12.8213 | -22.6895 | 29.959 | -29.3194 | 21.2557 | -9.6884 | 1.9582 | [12 empty] | 1.2716 | 2.9795 | 6.1302 | -1.4594 | 0.011041
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1729 | 8.9397 | [20 empty] | -0.091942 | 4.7312 | 12.8361 | -22.7289 | 30.0284 | -29.4044 | 21.3296 | -9.7277 | 1.9672 | [12 empty] | 1.2716 | 2.9812 | 6.1373 | -1.4619 | 0.011066
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1730 | 8.946 | [20 empty] | -0.091942 | 4.7339 | 12.8509 | -22.7683 | 30.0979 | -29.4895 | 21.4037 | -9.7671 | 1.9764 | [12 empty] | 1.2716 | 2.9829 | 6.1444 | -1.4644 | 0.011092
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1731 | 8.9523 | [20 empty] | -0.091942 | 4.7367 | 12.8658 | -22.8078 | 30.1676 | -29.5748 | 21.4781 | -9.8067 | 1.9855 | [12 empty] | 1.2716 | 2.9847 | 6.1515 | -1.467 | 0.011117
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1732 | 8.9586 | [20 empty] | -0.091942 | 4.7394 | 12.8807 | -22.8474 | 30.2373 | -29.6604 | 21.5526 | -9.8464 | 1.9947 | [12 empty] | 1.2716 | 2.9864 | 6.1586 | -1.4695 | 0.011143
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1733 | 8.965 | [20 empty] | -0.091942 | 4.7422 | 12.8956 | -22.887 | 30.3072 | -29.7461 | 21.6274 | -9.8863 | 2.004 | [12 empty] | 1.2716 | 2.9881 | 6.1657 | -1.4721 | 0.011169
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1734 | 8.9713 | [20 empty] | -0.091942 | 4.7449 | 12.9104 | -22.9266 | 30.3772 | -29.832 | 21.7024 | -9.9263 | 2.0132 | [12 empty] | 1.2716 | 2.9898 | 6.1728 | -1.4746 | 0.011195
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1735 | 8.9776 | [20 empty] | -0.091942 | 4.7476 | 12.9253 | -22.9663 | 30.4474 | -29.9181 | 21.7776 | -9.9665 | 2.0225 | [12 empty] | 1.2716 | 2.9915 | 6.1799 | -1.4772 | 0.011221
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1736 | 8.9839 | [20 empty] | -0.091942 | 4.7504 | 12.9402 | -23.0061 | 30.5176 | -30.0044 | 21.853 | -10.0067 | 2.0319 | [12 empty] | 1.2716 | 2.9933 | 6.1871 | -1.4797 | 0.011247
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1737 | 8.9903 | [20 empty] | -0.091942 | 4.7531 | 12.9552 | -23.0458 | 30.588 | -30.091 | 21.9286 | -10.0472 | 2.0413 | [12 empty] | 1.2716 | 2.995 | 6.1942 | -1.4823 | 0.011272
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1738 | 8.9966 | [20 empty] | -0.091942 | 4.7558 | 12.9701 | -23.0857 | 30.6585 | -30.1777 | 22.0045 | -10.0877 | 2.0507 | [12 empty] | 1.2716 | 2.9967 | 6.2013 | -1.4848 | 0.011298
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1739 | 9.0029 | [20 empty] | -0.091942 | 4.7586 | 12.985 | -23.1255 | 30.7291 | -30.2646 | 22.0806 | -10.1284 | 2.0601 | [12 empty] | 1.2716 | 2.9984 | 6.2085 | -1.4874 | 0.011324
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1740 | 9.0093 | [20 empty] | -0.091942 | 4.7613 | 12.9999 | -23.1654 | 30.7999 | -30.3517 | 22.1569 | -10.1693 | 2.0696 | [12 empty] | 1.2716 | 3.0002 | 6.2156 | -1.49 | 0.011351
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1741 | 9.0156 | [20 empty] | -0.091942 | 4.764 | 13.0149 | -23.2054 | 30.8707 | -30.439 | 22.2334 | -10.2102 | 2.0792 | [12 empty] | 1.2716 | 3.0019 | 6.2228 | -1.4925 | 0.011377
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1742 | 9.0219 | [20 empty] | -0.091942 | 4.7668 | 13.0298 | -23.2454 | 30.9417 | -30.5265 | 22.3101 | -10.2514 | 2.0887 | [12 empty] | 1.2716 | 3.0036 | 6.2299 | -1.4951 | 0.011403
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1743 | 9.0283 | [20 empty] | -0.091942 | 4.7695 | 13.0448 | -23.2855 | 31.0128 | -30.6143 | 22.387 | -10.2926 | 2.0984 | [12 empty] | 1.2716 | 3.0053 | 6.2371 | -1.4977 | 0.011429
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1744 | 9.0346 | [20 empty] | -0.091942 | 4.7723 | 13.0598 | -23.3256 | 31.0841 | -30.7022 | 22.4642 | -10.334 | 2.108 | [12 empty] | 1.2716 | 3.0071 | 6.2442 | -1.5003 | 0.011455
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1745 | 9.0409 | [20 empty] | -0.091942 | 4.775 | 13.0748 | -23.3657 | 31.1554 | -30.7903 | 22.5416 | -10.3756 | 2.1177 | [12 empty] | 1.2716 | 3.0088 | 6.2514 | -1.5029 | 0.011482
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1746 | 9.0473 | [20 empty] | -0.091942 | 4.7777 | 13.0898 | -23.4059 | 31.2269 | -30.8786 | 22.6192 | -10.4173 | 2.1274 | [12 empty] | 1.2716 | 3.0105 | 6.2586 | -1.5054 | 0.011508
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1747 | 9.0536 | [20 empty] | -0.091942 | 4.7805 | 13.1047 | -23.4462 | 31.2985 | -30.9672 | 22.6971 | -10.4591 | 2.1372 | [12 empty] | 1.2716 | 3.0122 | 6.2657 | -1.508 | 0.011534
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1748 | 9.06 | [20 empty] | -0.091942 | 4.7832 | 13.1198 | -23.4864 | 31.3702 | -31.0559 | 22.7751 | -10.5011 | 2.147 | [12 empty] | 1.2716 | 3.014 | 6.2729 | -1.5106 | 0.011561
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1749 | 9.0663 | [20 empty] | -0.091942 | 4.7859 | 13.1348 | -23.5268 | 31.4421 | -31.1448 | 22.8534 | -10.5432 | 2.1568 | [12 empty] | 1.2716 | 3.0157 | 6.2801 | -1.5132 | 0.011587
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1750 | 9.0726 | [20 empty] | -0.091942 | 4.7887 | 13.1498 | -23.5671 | 31.514 | -31.234 | 22.9319 | -10.5855 | 2.1667 | [12 empty] | 1.2716 | 3.0174 | 6.2873 | -1.5158 | 0.011614
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1751 | 9.079 | [20 empty] | -0.091942 | 4.7914 | 13.1648 | -23.6076 | 31.5861 | -31.3233 | 23.0107 | -10.6279 | 2.1766 | [12 empty] | 1.2716 | 3.0191 | 6.2945 | -1.5184 | 0.01164
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1752 | 9.0853 | [20 empty] | -0.091942 | 4.7942 | 13.1799 | -23.648 | 31.6584 | -31.4129 | 23.0896 | -10.6705 | 2.1866 | [12 empty] | 1.2716 | 3.0209 | 6.3016 | -1.521 | 0.011667
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1753 | 9.0917 | [20 empty] | -0.091942 | 4.7969 | 13.1949 | -23.6886 | 31.7307 | -31.5026 | 23.1688 | -10.7132 | 2.1966 | [12 empty] | 1.2716 | 3.0226 | 6.3088 | -1.5236 | 0.011694
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1754 | 9.098 | [20 empty] | -0.091942 | 4.7996 | 13.21 | -23.7291 | 31.8032 | -31.5926 | 23.2482 | -10.756 | 2.2067 | [12 empty] | 1.2716 | 3.0243 | 6.316 | -1.5262 | 0.01172
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1755 | 9.1044 | [20 empty] | -0.091942 | 4.8024 | 13.225 | -23.7697 | 31.8758 | -31.6827 | 23.3279 | -10.799 | 2.2167 | [12 empty] | 1.2716 | 3.026 | 6.3232 | -1.5288 | 0.011747
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1756 | 9.1107 | [20 empty] | -0.091942 | 4.8051 | 13.2401 | -23.8104 | 31.9485 | -31.7731 | 23.4077 | -10.8422 | 2.2269 | [12 empty] | 1.2716 | 3.0278 | 6.3305 | -1.5315 | 0.011774
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1757 | 9.1171 | [20 empty] | -0.091942 | 4.8078 | 13.2552 | -23.8511 | 32.0213 | -31.8637 | 23.4878 | -10.8855 | 2.237 | [12 empty] | 1.2716 | 3.0295 | 6.3377 | -1.5341 | 0.011801
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1758 | 9.1234 | [20 empty] | -0.091942 | 4.8106 | 13.2703 | -23.8918 | 32.0943 | -31.9544 | 23.5682 | -10.9289 | 2.2472 | [12 empty] | 1.2716 | 3.0312 | 6.3449 | -1.5367 | 0.011828
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1759 | 9.1298 | [20 empty] | -0.091942 | 4.8133 | 13.2854 | -23.9326 | 32.1674 | -32.0454 | 23.6487 | -10.9725 | 2.2575 | [12 empty] | 1.2716 | 3.0329 | 6.3521 | -1.5393 | 0.011854
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1760 | 9.1361 | [20 empty] | -0.091942 | 4.816 | 13.3005 | -23.9735 | 32.2406 | -32.1366 | 23.7295 | -11.0162 | 2.2678 | [12 empty] | 1.2716 | 3.0347 | 6.3593 | -1.5419 | 0.011881
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1761 | 9.1425 | [20 empty] | -0.091942 | 4.8188 | 13.3156 | -24.0144 | 32.3139 | -32.228 | 23.8105 | -11.0601 | 2.2781 | [12 empty] | 1.2716 | 3.0364 | 6.3666 | -1.5446 | 0.011908
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1762 | 9.1488 | [20 empty] | -0.091942 | 4.8215 | 13.3308 | -24.0553 | 32.3874 | -32.3196 | 23.8917 | -11.1042 | 2.2885 | [12 empty] | 1.2716 | 3.0381 | 6.3738 | -1.5472 | 0.011936
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1763 | 9.1552 | [20 empty] | -0.091942 | 4.8243 | 13.3459 | -24.0963 | 32.4609 | -32.4115 | 23.9732 | -11.1483 | 2.2989 | [12 empty] | 1.2716 | 3.0398 | 6.381 | -1.5498 | 0.011963
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1764 | 9.1615 | [20 empty] | -0.091942 | 4.827 | 13.361 | -24.1373 | 32.5347 | -32.5035 | 24.0549 | -11.1927 | 2.3093 | [12 empty] | 1.2716 | 3.0416 | 6.3883 | -1.5525 | 0.01199
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1765 | 9.1679 | [20 empty] | -0.091942 | 4.8297 | 13.3762 | -24.1784 | 32.6085 | -32.5957 | 24.1369 | -11.2372 | 2.3198 | [12 empty] | 1.2716 | 3.0433 | 6.3955 | -1.5551 | 0.012017
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1766 | 9.1743 | [20 empty] | -0.091942 | 4.8325 | 13.3913 | -24.2195 | 32.6825 | -32.6882 | 24.219 | -11.2818 | 2.3304 | [12 empty] | 1.2716 | 3.045 | 6.4028 | -1.5578 | 0.012044
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1767 | 9.1806 | [20 empty] | -0.091942 | 4.8352 | 13.4065 | -24.2607 | 32.7565 | -32.7808 | 24.3014 | -11.3266 | 2.3409 | [12 empty] | 1.2716 | 3.0467 | 6.41 | -1.5604 | 0.012072
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1768 | 9.187 | [20 empty] | -0.091942 | 4.8379 | 13.4217 | -24.3019 | 32.8308 | -32.8737 | 24.3841 | -11.3716 | 2.3516 | [12 empty] | 1.2716 | 3.0484 | 6.4173 | -1.5631 | 0.012099
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1769 | 9.1933 | [20 empty] | -0.091942 | 4.8407 | 13.4369 | -24.3431 | 32.9051 | -32.9667 | 24.4669 | -11.4167 | 2.3622 | [12 empty] | 1.2716 | 3.0502 | 6.4245 | -1.5657 | 0.012126
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1770 | 9.1997 | [20 empty] | -0.091942 | 4.8434 | 13.4521 | -24.3844 | 32.9796 | -33.06 | 24.55 | -11.4619 | 2.3729 | [12 empty] | 1.2716 | 3.0519 | 6.4318 | -1.5684 | 0.012154
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1771 | 9.2061 | [20 empty] | -0.091942 | 4.8461 | 13.4673 | -24.4258 | 33.0542 | -33.1535 | 24.6334 | -11.5073 | 2.3837 | [12 empty] | 1.2716 | 3.0536 | 6.4391 | -1.571 | 0.012181
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1772 | 9.2124 | [20 empty] | -0.091942 | 4.8489 | 13.4825 | -24.4672 | 33.1289 | -33.2472 | 24.7169 | -11.5529 | 2.3945 | [12 empty] | 1.2716 | 3.0553 | 6.4463 | -1.5737 | 0.012209
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1773 | 9.2188 | [20 empty] | -0.091942 | 4.8516 | 13.4977 | -24.5086 | 33.2037 | -33.3412 | 24.8008 | -11.5986 | 2.4053 | [12 empty] | 1.2716 | 3.0571 | 6.4536 | -1.5764 | 0.012236
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1774 | 9.2252 | [20 empty] | -0.091942 | 4.8544 | 13.5129 | -24.5501 | 33.2787 | -33.4353 | 24.8848 | -11.6445 | 2.4162 | [12 empty] | 1.2716 | 3.0588 | 6.4609 | -1.579 | 0.012264
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1775 | 9.2315 | [20 empty] | -0.091942 | 4.8571 | 13.5282 | -24.5917 | 33.3538 | -33.5296 | 24.9691 | -11.6905 | 2.4271 | [12 empty] | 1.2716 | 3.0605 | 6.4682 | -1.5817 | 0.012292
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1776 | 9.2379 | [20 empty] | -0.091942 | 4.8598 | 13.5434 | -24.6333 | 33.429 | -33.6242 | 25.0536 | -11.7367 | 2.4381 | [12 empty] | 1.2716 | 3.0622 | 6.4755 | -1.5844 | 0.012319
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1777 | 9.2443 | [20 empty] | -0.091942 | 4.8626 | 13.5587 | -24.6749 | 33.5044 | -33.719 | 25.1384 | -11.783 | 2.4491 | [12 empty] | 1.2716 | 3.064 | 6.4828 | -1.5871 | 0.012347
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1778 | 9.2506 | [20 empty] | -0.091942 | 4.8653 | 13.574 | -24.7166 | 33.5799 | -33.8139 | 25.2234 | -11.8295 | 2.4601 | [12 empty] | 1.2716 | 3.0657 | 6.4901 | -1.5897 | 0.012375
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1779 | 9.257 | [20 empty] | -0.091942 | 4.868 | 13.5892 | -24.7583 | 33.6555 | -33.9091 | 25.3086 | -11.8762 | 2.4712 | [12 empty] | 1.2716 | 3.0674 | 6.4974 | -1.5924 | 0.012403
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1780 | 9.2634 | [20 empty] | -0.091942 | 4.8708 | 13.6045 | -24.8001 | 33.7312 | -34.0045 | 25.3941 | -11.923 | 2.4823 | [12 empty] | 1.2716 | 3.0691 | 6.5047 | -1.5951 | 0.012431
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1781 | 9.2698 | [20 empty] | -0.091942 | 4.8735 | 13.6198 | -24.8419 | 33.8071 | -34.1002 | 25.4798 | -11.9699 | 2.4935 | [12 empty] | 1.2716 | 3.0709 | 6.512 | -1.5978 | 0.012459
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1782 | 9.2761 | [20 empty] | -0.091942 | 4.8762 | 13.6351 | -24.8838 | 33.8831 | -34.196 | 25.5658 | -12.0171 | 2.5047 | [12 empty] | 1.2716 | 3.0726 | 6.5193 | -1.6005 | 0.012487
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1783 | 9.2825 | [20 empty] | -0.091942 | 4.879 | 13.6504 | -24.9257 | 33.9592 | -34.2921 | 25.652 | -12.0643 | 2.516 | [12 empty] | 1.2716 | 3.0743 | 6.5266 | -1.6032 | 0.012515
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1784 | 9.2889 | [20 empty] | -0.091942 | 4.8817 | 13.6657 | -24.9676 | 34.0354 | -34.3883 | 25.7384 | -12.1118 | 2.5273 | [12 empty] | 1.2716 | 3.076 | 6.5339 | -1.6059 | 0.012543
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1785 | 9.2953 | [20 empty] | -0.091942 | 4.8845 | 13.681 | -25.0096 | 34.1118 | -34.4848 | 25.8251 | -12.1594 | 2.5387 | [12 empty] | 1.2716 | 3.0778 | 6.5413 | -1.6086 | 0.012571
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1786 | 9.3016 | [20 empty] | -0.091942 | 4.8872 | 13.6964 | -25.0517 | 34.1883 | -34.5815 | 25.912 | -12.2072 | 2.5501 | [12 empty] | 1.2716 | 3.0795 | 6.5486 | -1.6113 | 0.012599
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1787 | 9.308 | [20 empty] | -0.091942 | 4.8899 | 13.7117 | -25.0938 | 34.265 | -34.6785 | 25.9992 | -12.2551 | 2.5615 | [12 empty] | 1.2716 | 3.0812 | 6.5559 | -1.614 | 0.012627
7 | 1413.8919 | 1382 | 6.9947 | 7.0006 | 1413 | 1414 | 1788 | 9.3144 | [20 empty] | -0.091942 | 4.8927 | 13.7271 | -25.136 | 34.3417 | -34.7756 | 26.0866 | -12.3032 | 2.573 | [12 empty] | 1.2716 | 3.0829 | 6.5633 | -1.6167 | 0.012656
[7 empty] | 1789 | 9.3208 | [20 empty] | -0.091942 | 4.8954 | 13.7424 | -25.1782 | 34.4186 | -34.8729 | 26.1743 | -12.3514 | 2.5845 | [12 empty] | 1.2716 | 3.0847 | 6.5706 | -1.6194 | 0.012684
[7 empty] | 1790 | 9.3272 | [20 empty] | -0.091942 | 4.8981 | 13.7578 | -25.2204 | 34.4956 | -34.9705 | 26.2622 | -12.3998 | 2.5961 | [12 empty] | 1.2716 | 3.0864 | 6.578 | -1.6221 | 0.012712
[7 empty] | 1791 | 9.3335 | [20 empty] | -0.091942 | 4.9009 | 13.7732 | -25.2627 | 34.5728 | -35.0683 | 26.3503 | -12.4484 | 2.6077 | [12 empty] | 1.2716 | 3.0881 | 6.5853 | -1.6249 | 0.012741
[7 empty] | 1792 | 9.3399 | [20 empty] | -0.091942 | 4.9036 | 13.7886 | -25.305 | 34.6501 | -35.1663 | 26.4387 | -12.4971 | 2.6194 | [12 empty] | 1.2716 | 3.0898 | 6.5927 | -1.6276 | 0.012769
[7 empty] | 1793 | 9.3463 | [20 empty] | -0.091942 | 4.9063 | 13.804 | -25.3474 | 34.7275 | -35.2646 | 26.5274 | -12.546 | 2.6311 | [12 empty] | 1.2716 | 3.0916 | 6.6 | -1.6303 | 0.012798
[7 empty] | 1794 | 9.3527 | [20 empty] | -0.091942 | 4.9091 | 13.8194 | -25.3899 | 34.805 | -35.363 | 26.6163 | -12.5951 | 2.6429 | [12 empty] | 1.2716 | 3.0933 | 6.6074 | -1.633 | 0.012827
[7 empty] | 1795 | 9.3591 | [20 empty] | -0.091942 | 4.9118 | 13.8348 | -25.4323 | 34.8827 | -35.4617 | 26.7054 | -12.6443 | 2.6547 | [12 empty] | 1.2716 | 3.095 | 6.6148 | -1.6358 | 0.012855
[7 empty] | 1796 | 9.3655 | [20 empty] | -0.091942 | 4.9146 | 13.8502 | -25.4749 | 34.9605 | -35.5606 | 26.7948 | -12.6937 | 2.6665 | [12 empty] | 1.2716 | 3.0967 | 6.6221 | -1.6385 | 0.012884
[7 empty] | 1797 | 9.3719 | [20 empty] | -0.091942 | 4.9173 | 13.8656 | -25.5174 | 35.0384 | -35.6597 | 26.8844 | -12.7433 | 2.6785 | [12 empty] | 1.2716 | 3.0985 | 6.6295 | -1.6413 | 0.012913
[7 empty] | 1798 | 9.3783 | [20 empty] | -0.091942 | 4.92 | 13.881 | -25.5601 | 35.1165 | -35.759 | 26.9743 | -12.793 | 2.6904 | [12 empty] | 1.2716 | 3.1002 | 6.6369 | -1.644 | 0.012941
[7 empty] | 1799 | 9.3846 | [20 empty] | -0.091942 | 4.9228 | 13.8965 | -25.6027 | 35.1946 | -35.8586 | 27.0645 | -12.8429 | 2.7024 | [12 empty] | 1.2716 | 3.1019 | 6.6443 | -1.6467 | 0.01297
[7 empty] | 1800 | 9.391 | [20 empty] | -0.091942 | 4.9255 | 13.9119 | -25.6455 | 35.273 | -35.9583 | 27.1549 | -12.8929 | 2.7144 | [12 empty] | 1.2716 | 3.1036 | 6.6517 | -1.6495 | 0.012999
[7 empty] | 1801 | 9.3974 | [20 empty] | -0.091942 | 4.9282 | 13.9274 | -25.6882 | 35.3514 | -36.0583 | 27.2455 | -12.9432 | 2.7265 | [12 empty] | 1.2716 | 3.1053 | 6.6591 | -1.6522 | 0.013028
[7 empty] | 1802 | 9.4038 | [20 empty] | -0.091942 | 4.931 | 13.9429 | -25.731 | 35.43 | -36.1585 | 27.3364 | -12.9936 | 2.7387 | [12 empty] | 1.2716 | 3.1071 | 6.6665 | -1.655 | 0.013057
[7 empty] | 1803 | 9.4102 | [20 empty] | -0.091942 | 4.9337 | 13.9584 | -25.7739 | 35.5087 | -36.259 | 27.4275 | -13.0441 | 2.7508 | [12 empty] | 1.2716 | 3.1088 | 6.6739 | -1.6577 | 0.013086
[7 empty] | 1804 | 9.4166 | [20 empty] | -0.091942 | 4.9364 | 13.9738 | -25.8168 | 35.5875 | -36.3596 | 27.5189 | -13.0948 | 2.7631 | [12 empty] | 1.2716 | 3.1105 | 6.6813 | -1.6605 | 0.013115
[7 empty] | 1805 | 9.423 | [20 empty] | -0.091942 | 4.9392 | 13.9893 | -25.8598 | 35.6665 | -36.4605 | 27.6106 | -13.1457 | 2.7753 | [12 empty] | 1.2716 | 3.1122 | 6.6887 | -1.6633 | 0.013144
[7 empty] | 1806 | 9.4294 | [20 empty] | -0.091942 | 4.9419 | 14.0048 | -25.9028 | 35.7456 | -36.5616 | 27.7025 | -13.1968 | 2.7877 | [12 empty] | 1.2716 | 3.114 | 6.6961 | -1.666 | 0.013173
[7 empty] | 1807 | 9.4358 | [20 empty] | -0.091942 | 4.9447 | 14.0204 | -25.9458 | 35.8249 | -36.663 | 27.7947 | -13.248 | 2.8 | [12 empty] | 1.2716 | 3.1157 | 6.7035 | -1.6688 | 0.013202
[7 empty] | 1808 | 9.4422 | [20 empty] | -0.091942 | 4.9474 | 14.0359 | -25.9889 | 35.9042 | -36.7645 | 27.8871 | -13.2994 | 2.8125 | [12 empty] | 1.2716 | 3.1174 | 6.7109 | -1.6716 | 0.013232
[7 empty] | 1809 | 9.4486 | [20 empty] | -0.091942 | 4.9501 | 14.0514 | -26.0321 | 35.9837 | -36.8663 | 27.9798 | -13.351 | 2.8249 | [12 empty] | 1.2716 | 3.1191 | 6.7184 | -1.6744 | 0.013261
[7 empty] | 1810 | 9.455 | [20 empty] | -0.091942 | 4.9529 | 14.067 | -26.0753 | 36.0634 | -36.9683 | 28.0727 | -13.4028 | 2.8374 | [12 empty] | 1.2716 | 3.1209 | 6.7258 | -1.6771 | 0.01329
[7 empty] | 1811 | 9.4614 | [20 empty] | -0.091942 | 4.9556 | 14.0825 | -26.1185 | 36.1431 | -37.0706 | 28.1659 | -13.4547 | 2.85 | [12 empty] | 1.2716 | 3.1226 | 6.7332 | -1.6799 | 0.01332
[7 empty] | 1812 | 9.4678 | [20 empty] | -0.091942 | 4.9583 | 14.0981 | -26.1618 | 36.223 | -37.173 | 28.2593 | -13.5068 | 2.8626 | [12 empty] | 1.2716 | 3.1243 | 6.7407 | -1.6827 | 0.013349
[7 empty] | 1813 | 9.4742 | [20 empty] | -0.091942 | 4.9611 | 14.1136 | -26.2051 | 36.3031 | -37.2757 | 28.353 | -13.559 | 2.8753 | [12 empty] | 1.2716 | 3.126 | 6.7481 | -1.6855 | 0.013379
[7 empty] | 1814 | 9.4806 | [20 empty] | -0.091942 | 4.9638 | 14.1292 | -26.2485 | 36.3832 | -37.3786 | 28.447 | -13.6115 | 2.888 | [12 empty] | 1.2716 | 3.1278 | 6.7555 | -1.6883 | 0.013408
[7 empty] | 1815 | 9.487 | [20 empty] | -0.091942 | 4.9665 | 14.1448 | -26.292 | 36.4635 | -37.4818 | 28.5412 | -13.6641 | 2.9008 | [12 empty] | 1.2716 | 3.1295 | 6.763 | -1.6911 | 0.013438
[7 empty] | 1816 | 9.4935 | [20 empty] | -0.091942 | 4.9693 | 14.1604 | -26.3354 | 36.5439 | -37.5851 | 28.6357 | -13.7169 | 2.9136 | [12 empty] | 1.2716 | 3.1312 | 6.7704 | -1.6939 | 0.013467
[7 empty] | 1817 | 9.4999 | [20 empty] | -0.091942 | 4.972 | 14.176 | -26.379 | 36.6245 | -37.6887 | 28.7304 | -13.7698 | 2.9264 | [12 empty] | 1.2716 | 3.1329 | 6.7779 | -1.6967 | 0.013497
[7 empty] | 1818 | 9.5063 | [20 empty] | -0.091942 | 4.9748 | 14.1916 | -26.4225 | 36.7052 | -37.7926 | 28.8254 | -13.823 | 2.9393 | [12 empty] | 1.2716 | 3.1347 | 6.7854 | -1.6995 | 0.013527
[7 empty] | 1819 | 9.5127 | [20 empty] | -0.091942 | 4.9775 | 14.2072 | -26.4662 | 36.786 | -37.8966 | 28.9207 | -13.8763 | 2.9523 | [12 empty] | 1.2716 | 3.1364 | 6.7928 | -1.7023 | 0.013557
[7 empty] | 1820 | 9.5191 | [20 empty] | -0.091942 | 4.9802 | 14.2228 | -26.5098 | 36.867 | -38.0009 | 29.0162 | -13.9298 | 2.9653 | [12 empty] | 1.2716 | 3.1381 | 6.8003 | -1.7051 | 0.013586
[7 empty] | 1821 | 9.5255 | [20 empty] | -0.091942 | 4.983 | 14.2385 | -26.5536 | 36.9481 | -38.1054 | 29.112 | -13.9834 | 2.9784 | [12 empty] | 1.2716 | 3.1398 | 6.8078 | -1.7079 | 0.013616
[7 empty] | 1822 | 9.5319 | [20 empty] | -0.091942 | 4.9857 | 14.2541 | -26.5973 | 37.0293 | -38.2102 | 29.2081 | -14.0373 | 2.9915 | [12 empty] | 1.2716 | 3.1416 | 6.8153 | -1.7107 | 0.013646
[7 empty] | 1823 | 9.5383 | [20 empty] | -0.091942 | 4.9884 | 14.2697 | -26.6411 | 37.1107 | -38.3151 | 29.3044 | -14.0913 | 3.0046 | [12 empty] | 1.2716 | 3.1433 | 6.8227 | -1.7135 | 0.013676
[7 empty] | 1824 | 9.5448 | [20 empty] | -0.091942 | 4.9912 | 14.2854 | -26.685 | 37.1921 | -38.4203 | 29.401 | -14.1455 | 3.0178 | [12 empty] | 1.2716 | 3.145 | 6.8302 | -1.7163 | 0.013706
[7 empty] | 1825 | 9.5512 | [20 empty] | -0.091942 | 4.9939 | 14.3011 | -26.7289 | 37.2738 | -38.5258 | 29.4978 | -14.1999 | 3.0311 | [12 empty] | 1.2716 | 3.1467 | 6.8377 | -1.7192 | 0.013736
[7 empty] | 1826 | 9.5576 | [20 empty] | -0.091942 | 4.9966 | 14.3168 | -26.7729 | 37.3555 | -38.6314 | 29.5949 | -14.2544 | 3.0444 | [12 empty] | 1.2716 | 3.1485 | 6.8452 | -1.722 | 0.013766
[7 empty] | 1827 | 9.564 | [20 empty] | -0.091942 | 4.9994 | 14.3324 | -26.8169 | 37.4374 | -38.7373 | 29.6923 | -14.3092 | 3.0578 | [12 empty] | 1.2716 | 3.1502 | 6.8527 | -1.7248 | 0.013797
[7 empty] | 1828 | 9.5704 | [20 empty] | -0.091942 | 5.0021 | 14.3481 | -26.861 | 37.5195 | -38.8435 | 29.7899 | -14.3641 | 3.0712 | [12 empty] | 1.2716 | 3.1519 | 6.8602 | -1.7277 | 0.013827
[7 empty] | 1829 | 9.5768 | [20 empty] | -0.091942 | 5.0049 | 14.3638 | -26.9051 | 37.6016 | -38.9498 | 29.8879 | -14.4192 | 3.0847 | [12 empty] | 1.2716 | 3.1536 | 6.8677 | -1.7305 | 0.013857
[7 empty] | 1830 | 9.5833 | [20 empty] | -0.091942 | 5.0076 | 14.3795 | -26.9492 | 37.6839 | -39.0564 | 29.986 | -14.4744 | 3.0982 | [12 empty] | 1.2716 | 3.1554 | 6.8752 | -1.7333 | 0.013887
[7 empty] | 1831 | 9.5897 | [20 empty] | -0.091942 | 5.0103 | 14.3953 | -26.9934 | 37.7664 | -39.1632 | 30.0845 | -14.5299 | 3.1118 | [12 empty] | 1.2716 | 3.1571 | 6.8828 | -1.7362 | 0.013918
[7 empty] | 1832 | 9.5961 | [20 empty] | -0.091942 | 5.0131 | 14.411 | -27.0377 | 37.8489 | -39.2703 | 30.1832 | -14.5855 | 3.1254 | [12 empty] | 1.2716 | 3.1588 | 6.8903 | -1.739 | 0.013948
[7 empty] | 1833 | 9.6025 | [20 empty] | -0.091942 | 5.0158 | 14.4267 | -27.082 | 37.9317 | -39.3776 | 30.2822 | -14.6414 | 3.1391 | [12 empty] | 1.2716 | 3.1605 | 6.8978 | -1.7419 | 0.013979
[7 empty] | 1834 | 9.609 | [20 empty] | -0.091942 | 5.0185 | 14.4425 | -27.1263 | 38.0145 | -39.4851 | 30.3815 | -14.6974 | 3.1528 | [12 empty] | 1.2716 | 3.1622 | 6.9053 | -1.7447 | 0.014009
[7 empty] | 1835 | 9.6154 | [20 empty] | -0.091942 | 5.0213 | 14.4582 | -27.1707 | 38.0975 | -39.5929 | 30.481 | -14.7536 | 3.1666 | [12 empty] | 1.2716 | 3.164 | 6.9129 | -1.7476 | 0.01404
[7 empty] | 1836 | 9.6218 | [20 empty] | -0.091942 | 5.024 | 14.474 | -27.2152 | 38.1806 | -39.7009 | 30.5808 | -14.8099 | 3.1804 | [12 empty] | 1.2716 | 3.1657 | 6.9204 | -1.7504 | 0.01407
[7 empty] | 1837 | 9.6283 | [20 empty] | -0.091942 | 5.0267 | 14.4898 | -27.2597 | 38.2638 | -39.8091 | 30.6809 | -14.8665 | 3.1943 | [12 empty] | 1.2716 | 3.1674 | 6.9279 | -1.7533 | 0.014101
[7 empty] | 1838 | 9.6347 | [20 empty] | -0.091942 | 5.0295 | 14.5055 | -27.3042 | 38.3472 | -39.9176 | 30.7812 | -14.9232 | 3.2082 | [12 empty] | 1.2716 | 3.1691 | 6.9355 | -1.7562 | 0.014132
[7 empty] | 1839 | 9.6411 | [20 empty] | -0.091942 | 5.0322 | 14.5213 | -27.3488 | 38.4308 | -40.0263 | 30.8818 | -14.9802 | 3.2222 | [12 empty] | 1.2716 | 3.1709 | 6.943 | -1.759 | 0.014163
[7 empty] | 1840 | 9.6475 | [20 empty] | -0.091942 | 5.035 | 14.5371 | -27.3934 | 38.5144 | -40.1353 | 30.9827 | -15.0373 | 3.2362 | [12 empty] | 1.2716 | 3.1726 | 6.9506 | -1.7619 | 0.014194
[7 empty] | 1841 | 9.654 | [20 empty] | -0.091942 | 5.0377 | 14.5529 | -27.4381 | 38.5982 | -40.2444 | 31.0839 | -15.0946 | 3.2503 | [12 empty] | 1.2716 | 3.1743 | 6.9581 | -1.7648 | 0.014224
[7 empty] | 1842 | 9.6604 | [20 empty] | -0.091942 | 5.0404 | 14.5687 | -27.4829 | 38.6821 | -40.3539 | 31.1853 | -15.1521 | 3.2645 | [12 empty] | 1.2716 | 3.176 | 6.9657 | -1.7677 | 0.014255
[7 empty] | 1843 | 9.6668 | [20 empty] | -0.091942 | 5.0432 | 14.5846 | -27.5276 | 38.7662 | -40.4635 | 31.287 | -15.2097 | 3.2787 | [12 empty] | 1.2716 | 3.1778 | 6.9733 | -1.7705 | 0.014286
[7 empty] | 1844 | 9.6733 | [20 empty] | -0.091942 | 5.0459 | 14.6004 | -27.5725 | 38.8504 | -40.5734 | 31.389 | -15.2676 | 3.293 | [12 empty] | 1.2716 | 3.1795 | 6.9808 | -1.7734 | 0.014317
[7 empty] | 1845 | 9.6797 | [20 empty] | -0.091942 | 5.0486 | 14.6162 | -27.6174 | 38.9348 | -40.6835 | 31.4913 | -15.3257 | 3.3073 | [12 empty] | 1.2716 | 3.1812 | 6.9884 | -1.7763 | 0.014348
[7 empty] | 1846 | 9.6862 | [20 empty] | -0.091942 | 5.0514 | 14.6321 | -27.6623 | 39.0192 | -40.7939 | 31.5939 | -15.3839 | 3.3216 | [12 empty] | 1.2716 | 3.1829 | 6.996 | -1.7792 | 0.01438
[7 empty] | 1847 | 9.6926 | [20 empty] | -0.091942 | 5.0541 | 14.6479 | -27.7073 | 39.1039 | -40.9045 | 31.6967 | -15.4423 | 3.3361 | [12 empty] | 1.2716 | 3.1847 | 7.0036 | -1.7821 | 0.014411
[7 empty] | 1848 | 9.699 | [20 empty] | -0.091942 | 5.0568 | 14.6638 | -27.7523 | 39.1886 | -41.0154 | 31.7998 | -15.5009 | 3.3505 | [12 empty] | 1.2716 | 3.1864 | 7.0112 | -1.785 | 0.014442
[7 empty] | 1849 | 9.7055 | [20 empty] | -0.091942 | 5.0596 | 14.6797 | -27.7974 | 39.2735 | -41.1265 | 31.9032 | -15.5598 | 3.3651 | [12 empty] | 1.2716 | 3.1881 | 7.0187 | -1.7879 | 0.014473
[7 empty] | 1850 | 9.7119 | [20 empty] | -0.091942 | 5.0623 | 14.6956 | -27.8425 | 39.3585 | -41.2378 | 32.0068 | -15.6188 | 3.3797 | [12 empty] | 1.2716 | 3.1898 | 7.0263 | -1.7908 | 0.014505
[7 empty] | 1851 | 9.7184 | [20 empty] | -0.091942 | 5.0651 | 14.7115 | -27.8877 | 39.4437 | -41.3494 | 32.1108 | -15.678 | 3.3943 | [12 empty] | 1.2716 | 3.1916 | 7.0339 | -1.7937 | 0.014536
[7 empty] | 1852 | 9.7248 | [20 empty] | -0.091942 | 5.0678 | 14.7274 | -27.9329 | 39.529 | -41.4612 | 32.215 | -15.7373 | 3.409 | [12 empty] | 1.2716 | 3.1933 | 7.0415 | -1.7966 | 0.014567
[7 empty] | 1853 | 9.7312 | [20 empty] | -0.091942 | 5.0705 | 14.7433 | -27.9782 | 39.6144 | -41.5733 | 32.3195 | -15.7969 | 3.4238 | [12 empty] | 1.2716 | 3.195 | 7.0491 | -1.7995 | 0.014599
[7 empty] | 1854 | 9.7377 | [20 empty] | -0.091942 | 5.0733 | 14.7592 | -28.0235 | 39.7 | -41.6856 | 32.4243 | -15.8567 | 3.4386 | [12 empty] | 1.2716 | 3.1967 | 7.0568 | -1.8024 | 0.01463
[7 empty] | 1855 | 9.7441 | [20 empty] | -0.091942 | 5.076 | 14.7751 | -28.0689 | 39.7858 | -41.7981 | 32.5294 | -15.9167 | 3.4534 | [12 empty] | 1.2716 | 3.1985 | 7.0644 | -1.8054 | 0.014662
[7 empty] | 1856 | 9.7506 | [20 empty] | -0.091942 | 5.0787 | 14.791 | -28.1143 | 39.8716 | -41.9109 | 32.6348 | -15.9768 | 3.4683 | [12 empty] | 1.2716 | 3.2002 | 7.072 | -1.8083 | 0.014694
[7 empty] | 1857 | 9.757 | [20 empty] | -0.091942 | 5.0815 | 14.807 | -28.1597 | 39.9576 | -42.0239 | 32.7404 | -16.0372 | 3.4833 | [12 empty] | 1.2716 | 3.2019 | 7.0796 | -1.8112 | 0.014725
[7 empty] | 1858 | 9.7635 | [20 empty] | -0.091942 | 5.0842 | 14.8229 | -28.2053 | 40.0438 | -42.1372 | 32.8463 | -16.0977 | 3.4984 | [12 empty] | 1.2716 | 3.2036 | 7.0872 | -1.8141 | 0.014757
[7 empty] | 1859 | 9.7699 | [20 empty] | -0.091942 | 5.0869 | 14.8389 | -28.2508 | 40.13 | -42.2507 | 32.9525 | -16.1585 | 3.5135 | [12 empty] | 1.2716 | 3.2054 | 7.0949 | -1.8171 | 0.014789
[7 empty] | 1860 | 9.7764 | [20 empty] | -0.091942 | 5.0897 | 14.8549 | -28.2964 | 40.2164 | -42.3645 | 33.059 | -16.2194 | 3.5286 | [12 empty] | 1.2716 | 3.2071 | 7.1025 | -1.82 | 0.014821
[7 empty] | 1861 | 9.7828 | [20 empty] | -0.091942 | 5.0924 | 14.8708 | -28.3421 | 40.303 | -42.4785 | 33.1658 | -16.2805 | 3.5438 | [12 empty] | 1.2716 | 3.2088 | 7.1101 | -1.8229 | 0.014853
[7 empty] | 1862 | 9.7893 | [20 empty] | -0.091942 | 5.0952 | 14.8868 | -28.3878 | 40.3897 | -42.5927 | 33.2729 | -16.3419 | 3.5591 | [12 empty] | 1.2716 | 3.2105 | 7.1178 | -1.8259 | 0.014885
[7 empty] | 1863 | 9.7958 | [20 empty] | -0.091942 | 5.0979 | 14.9028 | -28.4336 | 40.4765 | -42.7072 | 33.3803 | -16.4034 | 3.5744 | [12 empty] | 1.2716 | 3.2122 | 7.1254 | -1.8288 | 0.014917
[7 empty] | 1864 | 9.8022 | [20 empty] | -0.091942 | 5.1006 | 14.9188 | -28.4794 | 40.5635 | -42.822 | 33.4879 | -16.4652 | 3.5898 | [12 empty] | 1.2716 | 3.214 | 7.1331 | -1.8318 | 0.014949
[7 empty] | 1865 | 9.8087 | [20 empty] | -0.091942 | 5.1034 | 14.9348 | -28.5252 | 40.6506 | -42.9369 | 33.5958 | -16.5271 | 3.6052 | [12 empty] | 1.2716 | 3.2157 | 7.1407 | -1.8347 | 0.014981
[7 empty] | 1866 | 9.8151 | [20 empty] | -0.091942 | 5.1061 | 14.9509 | -28.5712 | 40.7379 | -43.0522 | 33.7041 | -16.5892 | 3.6207 | [12 empty] | 1.2716 | 3.2174 | 7.1484 | -1.8377 | 0.015013
[7 empty] | 1867 | 9.8216 | [20 empty] | -0.091942 | 5.1088 | 14.9669 | -28.6171 | 40.8253 | -43.1677 | 33.8126 | -16.6515 | 3.6362 | [12 empty] | 1.2716 | 3.2191 | 7.1561 | -1.8406 | 0.015045
[7 empty] | 1868 | 9.828 | [20 empty] | -0.091942 | 5.1116 | 14.9829 | -28.6631 | 40.9128 | -43.2834 | 33.9214 | -16.7141 | 3.6519 | [12 empty] | 1.2716 | 3.2209 | 7.1637 | -1.8436 | 0.015077
[7 empty] | 1869 | 9.8345 | [20 empty] | -0.091942 | 5.1143 | 14.999 | -28.7092 | 41.0005 | -43.3994 | 34.0305 | -16.7768 | 3.6675 | [12 empty] | 1.2716 | 3.2226 | 7.1714 | -1.8465 | 0.01511
[7 empty] | 1870 | 9.841 | [20 empty] | -0.091942 | 5.117 | 15.015 | -28.7553 | 41.0883 | -43.5156 | 34.1399 | -16.8397 | 3.6833 | [12 empty] | 1.2716 | 3.2243 | 7.1791 | -1.8495 | 0.015142
[7 empty] | 1871 | 9.8474 | [20 empty] | -0.091942 | 5.1198 | 15.0311 | -28.8014 | 41.1763 | -43.6321 | 34.2496 | -16.9029 | 3.699 | [12 empty] | 1.2716 | 3.226 | 7.1868 | -1.8525 | 0.015174
[7 empty] | 1872 | 9.8539 | [20 empty] | -0.091942 | 5.1225 | 15.0472 | -28.8477 | 41.2644 | -43.7488 | 34.3596 | -16.9662 | 3.7149 | [12 empty] | 1.2716 | 3.2278 | 7.1944 | -1.8554 | 0.015207
[7 empty] | 1873 | 9.8603 | [20 empty] | -0.091942 | 5.1253 | 15.0632 | -28.8939 | 41.3526 | -43.8658 | 34.4698 | -17.0298 | 3.7308 | [12 empty] | 1.2716 | 3.2295 | 7.2021 | -1.8584 | 0.015239
[7 empty] | 1874 | 9.8668 | [20 empty] | -0.091942 | 5.128 | 15.0793 | -28.9402 | 41.441 | -43.983 | 34.5804 | -17.0935 | 3.7468 | [12 empty] | 1.2716 | 3.2312 | 7.2098 | -1.8614 | 0.015272
[7 empty] | 1875 | 9.8733 | [20 empty] | -0.091942 | 5.1307 | 15.0954 | -28.9866 | 41.5295 | -44.1005 | 34.6913 | -17.1575 | 3.7628 | [12 empty] | 1.2716 | 3.2329 | 7.2175 | -1.8644 | 0.015305
[7 empty] | 1876 | 9.8797 | [20 empty] | -0.091942 | 5.1335 | 15.1115 | -29.033 | 41.6182 | -44.2182 | 34.8024 | -17.2216 | 3.7789 | [12 empty] | 1.2716 | 3.2347 | 7.2252 | -1.8674 | 0.015337
[7 empty] | 1877 | 9.8862 | [20 empty] | -0.091942 | 5.1362 | 15.1277 | -29.0794 | 41.707 | -44.3362 | 34.9139 | -17.286 | 3.795 | [12 empty] | 1.2716 | 3.2364 | 7.2329 | -1.8704 | 0.01537
[7 empty] | 1878 | 9.8927 | [20 empty] | -0.091942 | 5.1389 | 15.1438 | -29.1259 | 41.796 | -44.4544 | 35.0256 | -17.3506 | 3.8112 | [12 empty] | 1.2716 | 3.2381 | 7.2406 | -1.8733 | 0.015403
[7 empty] | 1879 | 9.8992 | [20 empty] | -0.091942 | 5.1417 | 15.1599 | -29.1725 | 41.8851 | -44.5729 | 35.1377 | -17.4153 | 3.8275 | [12 empty] | 1.2716 | 3.2398 | 7.2483 | -1.8763 | 0.015436
[7 empty] | 1880 | 9.9056 | [20 empty] | -0.091942 | 5.1444 | 15.176 | -29.2191 | 41.9743 | -44.6916 | 35.25 | -17.4803 | 3.8438 | [12 empty] | 1.2716 | 3.2416 | 7.2561 | -1.8793 | 0.015469
[7 empty] | 1881 | 9.9121 | [20 empty] | -0.091942 | 5.1471 | 15.1922 | -29.2657 | 42.0637 | -44.8106 | 35.3627 | -17.5455 | 3.8602 | [12 empty] | 1.2716 | 3.2433 | 7.2638 | -1.8823 | 0.015501
[7 empty] | 1882 | 9.9186 | [20 empty] | -0.091942 | 5.1499 | 15.2084 | -29.3124 | 42.1532 | -44.9299 | 35.4756 | -17.6109 | 3.8766 | [12 empty] | 1.2716 | 3.245 | 7.2715 | -1.8853 | 0.015534
[7 empty] | 1883 | 9.925 | [20 empty] | -0.091942 | 5.1526 | 15.2245 | -29.3592 | 42.2429 | -45.0494 | 35.5889 | -17.6765 | 3.8932 | [12 empty] | 1.2716 | 3.2467 | 7.2792 | -1.8883 | 0.015568
[7 empty] | 1884 | 9.9315 | [20 empty] | -0.091942 | 5.1554 | 15.2407 | -29.406 | 42.3327 | -45.1691 | 35.7024 | -17.7423 | 3.9097 | [12 empty] | 1.2716 | 3.2485 | 7.287 | -1.8914 | 0.015601
[7 empty] | 1885 | 9.938 | [20 empty] | -0.091942 | 5.1581 | 15.2569 | -29.4528 | 42.4226 | -45.2891 | 35.8163 | -17.8084 | 3.9264 | [12 empty] | 1.2716 | 3.2502 | 7.2947 | -1.8944 | 0.015634
[7 empty] | 1886 | 9.9445 | [20 empty] | -0.091942 | 5.1608 | 15.2731 | -29.4997 | 42.5127 | -45.4094 | 35.9305 | -17.8746 | 3.9431 | [12 empty] | 1.2716 | 3.2519 | 7.3025 | -1.8974 | 0.015667
[7 empty] | 1887 | 9.951 | [20 empty] | -0.091942 | 5.1636 | 15.2893 | -29.5467 | 42.6029 | -45.5299 | 36.0449 | -17.941 | 3.9598 | [12 empty] | 1.2716 | 3.2536 | 7.3102 | -1.9004 | 0.0157
[7 empty] | 1888 | 9.9574 | [20 empty] | -0.091942 | 5.1663 | 15.3055 | -29.5937 | 42.6933 | -45.6506 | 36.1597 | -18.0077 | 3.9766 | [12 empty] | 1.2716 | 3.2554 | 7.318 | -1.9034 | 0.015734
[7 empty] | 1889 | 9.9639 | [20 empty] | -0.091942 | 5.169 | 15.3217 | -29.6407 | 42.7838 | -45.7717 | 36.2747 | -18.0746 | 3.9935 | [12 empty] | 1.2716 | 3.2571 | 7.3257 | -1.9065 | 0.015767
[7 empty] | 1890 | 9.9704 | [20 empty] | -0.091942 | 5.1718 | 15.3379 | -29.6878 | 42.8745 | -45.893 | 36.3901 | -18.1417 | 4.0105 | [12 empty] | 1.2716 | 3.2588 | 7.3335 | -1.9095 | 0.0158
[7 empty] | 1891 | 9.9769 | [20 empty] | -0.091942 | 5.1745 | 15.3542 | -29.735 | 42.9653 | -46.0145 | 36.5058 | -18.209 | 4.0275 | [12 empty] | 1.2716 | 3.2605 | 7.3412 | -1.9125 | 0.015834
[7 empty] | 1892 | 9.9834 | [20 empty] | -0.091942 | 5.1772 | 15.3704 | -29.7822 | 43.0563 | -46.1363 | 36.6218 | -18.2765 | 4.0445 | [12 empty] | 1.2716 | 3.2623 | 7.349 | -1.9156 | 0.015867
[7 empty] | 1893 | 9.9898 | [20 empty] | -0.091942 | 5.18 | 15.3867 | -29.8294 | 43.1474 | -46.2583 | 36.7381 | -18.3442 | 4.0617 | [12 empty] | 1.2716 | 3.264 | 7.3568 | -1.9186 | 0.015901
[7 empty] | 1894 | 9.9963 | [20 empty] | -0.091942 | 5.1827 | 15.4029 | -29.8767 | 43.2386 | -46.3807 | 36.8547 | -18.4121 | 4.0789 | [12 empty] | 1.2716 | 3.2657 | 7.3645 | -1.9216 | 0.015934
[7 empty] | 1895 | 10.0028 | [20 empty] | -0.091942 | 5.1855 | 15.4192 | -29.9241 | 43.33 | -46.5032 | 36.9716 | -18.4803 | 4.0961 | [12 empty] | 1.2716 | 3.2674 | 7.3723 | -1.9247 | 0.015968
[7 empty] | 1896 | 10.0093 | [20 empty] | -0.091942 | 5.1882 | 15.4355 | -29.9715 | 43.4215 | -46.6261 | 37.0888 | -18.5487 | 4.1134 | [12 empty] | 1.2716 | 3.2691 | 7.3801 | -1.9277 | 0.016002
[7 empty] | 1897 | 10.0158 | [20 empty] | -0.091942 | 5.1909 | 15.4517 | -30.0189 | 43.5132 | -46.7491 | 37.2063 | -18.6173 | 4.1308 | [12 empty] | 1.2716 | 3.2709 | 7.3879 | -1.9308 | 0.016036
[7 empty] | 1898 | 10.0223 | [20 empty] | -0.091942 | 5.1937 | 15.468 | -30.0664 | 43.605 | -46.8725 | 37.3241 | -18.6861 | 4.1483 | [12 empty] | 1.2716 | 3.2726 | 7.3957 | -1.9338 | 0.01607
[7 empty] | 1899 | 10.0288 | [20 empty] | -0.091942 | 5.1964 | 15.4843 | -30.114 | 43.697 | -46.9961 | 37.4423 | -18.7551 | 4.1658 | [12 empty] | 1.2716 | 3.2743 | 7.4035 | -1.9369 | 0.016103
[7 empty] | 1900 | 10.0353 | [20 empty] | -0.091942 | 5.1991 | 15.5007 | -30.1616 | 43.7891 | -47.12 | 37.5607 | -18.8243 | 4.1834 | [12 empty] | 1.2716 | 3.276 | 7.4113 | -1.94 | 0.016137
[7 empty] | 1901 | 10.0417 | [20 empty] | -0.091942 | 5.2019 | 15.517 | -30.2092 | 43.8814 | -47.2441 | 37.6795 | -18.8938 | 4.201 | [12 empty] | 1.2716 | 3.2778 | 7.4191 | -1.943 | 0.016171
[7 empty] | 1902 | 10.0482 | [20 empty] | -0.091942 | 5.2046 | 15.5333 | -30.2569 | 43.9738 | -47.3685 | 37.7986 | -18.9635 | 4.2187 | [12 empty] | 1.2716 | 3.2795 | 7.4269 | -1.9461 | 0.016205
[7 empty] | 1903 | 10.0547 | [20 empty] | -0.091942 | 5.2073 | 15.5496 | -30.3047 | 44.0664 | -47.4931 | 37.918 | -19.0334 | 4.2365 | [12 empty] | 1.2716 | 3.2812 | 7.4347 | -1.9492 | 0.01624
[7 empty] | 1904 | 10.0612 | [20 empty] | -0.091942 | 5.2101 | 15.566 | -30.3524 | 44.159 | -47.6181 | 38.0377 | -19.1035 | 4.2544 | [12 empty] | 1.2716 | 3.2829 | 7.4425 | -1.9522 | 0.016274
[7 empty] | 1905 | 10.0677 | [20 empty] | -0.091942 | 5.2128 | 15.5823 | -30.4003 | 44.2519 | -47.7432 | 38.1577 | -19.1738 | 4.2723 | [12 empty] | 1.2716 | 3.2847 | 7.4503 | -1.9553 | 0.016308
[7 empty] | 1906 | 10.0742 | [20 empty] | -0.091942 | 5.2156 | 15.5987 | -30.4482 | 44.3449 | -47.8687 | 38.2781 | -19.2444 | 4.2902 | [12 empty] | 1.2716 | 3.2864 | 7.4582 | -1.9584 | 0.016342
[7 empty] | 1907 | 10.0807 | [20 empty] | -0.091942 | 5.2183 | 15.6151 | -30.4961 | 44.438 | -47.9944 | 38.3987 | -19.3152 | 4.3083 | [12 empty] | 1.2716 | 3.2881 | 7.466 | -1.9615 | 0.016377
[7 empty] | 1908 | 10.0872 | [20 empty] | -0.091942 | 5.221 | 15.6315 | -30.5441 | 44.5313 | -48.1204 | 38.5197 | -19.3862 | 4.3264 | [12 empty] | 1.2716 | 3.2898 | 7.4738 | -1.9646 | 0.016411
[7 empty] | 1909 | 10.0937 | [20 empty] | -0.091942 | 5.2238 | 15.6479 | -30.5922 | 44.6247 | -48.2466 | 38.641 | -19.4574 | 4.3446 | [12 empty] | 1.2716 | 3.2916 | 7.4817 | -1.9677 | 0.016445
[7 empty] | 1910 | 10.1002 | [20 empty] | -0.091942 | 5.2265 | 15.6643 | -30.6403 | 44.7183 | -48.3731 | 38.7626 | -19.5289 | 4.3628 | [12 empty] | 1.2716 | 3.2933 | 7.4895 | -1.9707 | 0.01648
[7 empty] | 1911 | 10.1067 | [20 empty] | -0.091942 | 5.2292 | 15.6807 | -30.6885 | 44.812 | -48.4999 | 38.8845 | -19.6006 | 4.3811 | [12 empty] | 1.2716 | 3.295 | 7.4973 | -1.9738 | 0.016514
[7 empty] | 1912 | 10.1132 | [20 empty] | -0.091942 | 5.232 | 15.6971 | -30.7367 | 44.9059 | -48.6269 | 39.0068 | -19.6725 | 4.3995 | [12 empty] | 1.2716 | 3.2967 | 7.5052 | -1.9769 | 0.016549
[7 empty] | 1913 | 10.1197 | [20 empty] | -0.091942 | 5.2347 | 15.7135 | -30.7849 | 44.9999 | -48.7542 | 39.1293 | -19.7446 | 4.4179 | [12 empty] | 1.2716 | 3.2985 | 7.513 | -1.9801 | 0.016584
[7 empty] | 1914 | 10.1262 | [20 empty] | -0.091942 | 5.2374 | 15.7299 | -30.8332 | 45.0941 | -48.8817 | 39.2522 | -19.817 | 4.4364 | [12 empty] | 1.2716 | 3.3002 | 7.5209 | -1.9832 | 0.016618
[7 empty] | 1915 | 10.1327 | [20 empty] | -0.091942 | 5.2402 | 15.7464 | -30.8816 | 45.1884 | -49.0096 | 39.3754 | -19.8896 | 4.455 | [12 empty] | 1.2716 | 3.3019 | 7.5288 | -1.9863 | 0.016653
[7 empty] | 1916 | 10.1392 | [20 empty] | -0.091942 | 5.2429 | 15.7628 | -30.93 | 45.2829 | -49.1377 | 39.499 | -19.9624 | 4.4737 | [12 empty] | 1.2716 | 3.3036 | 7.5366 | -1.9894 | 0.016688
[7 empty] | 1917 | 10.1458 | [20 empty] | -0.091942 | 5.2457 | 15.7793 | -30.9784 | 45.3775 | -49.266 | 39.6228 | -20.0355 | 4.4924 | [12 empty] | 1.2716 | 3.3054 | 7.5445 | -1.9925 | 0.016723
[7 empty] | 1918 | 10.1523 | [20 empty] | -0.091942 | 5.2484 | 15.7957 | -31.0269 | 45.4722 | -49.3947 | 39.747 | -20.1087 | 4.5112 | [12 empty] | 1.2716 | 3.3071 | 7.5524 | -1.9956 | 0.016758
[7 empty] | 1919 | 10.1588 | [20 empty] | -0.091942 | 5.2511 | 15.8122 | -31.0755 | 45.5671 | -49.5236 | 39.8715 | -20.1822 | 4.53 | [12 empty] | 1.2716 | 3.3088 | 7.5602 | -1.9987 | 0.016793
[7 empty] | 1920 | 10.1653 | [20 empty] | -0.091942 | 5.2539 | 15.8287 | -31.1241 | 45.6622 | -49.6527 | 39.9963 | -20.256 | 4.5489 | [12 empty] | 1.2716 | 3.3105 | 7.5681 | -2.0019 | 0.016828
[7 empty] | 1921 | 10.1718 | [20 empty] | -0.091942 | 5.2566 | 15.8452 | -31.1727 | 45.7574 | -49.7822 | 40.1215 | -20.3299 | 4.5679 | [12 empty] | 1.2716 | 3.3123 | 7.576 | -2.005 | 0.016863
[7 empty] | 1922 | 10.1783 | [20 empty] | -0.091942 | 5.2593 | 15.8617 | -31.2214 | 45.8528 | -49.9119 | 40.2469 | -20.4041 | 4.587 | [12 empty] | 1.2716 | 3.314 | 7.5839 | -2.0081 | 0.016898
[7 empty] | 1923 | 10.1848 | [20 empty] | -0.091942 | 5.2621 | 15.8782 | -31.2702 | 45.9483 | -50.0419 | 40.3727 | -20.4786 | 4.6061 | [12 empty] | 1.2716 | 3.3157 | 7.5918 | -2.0113 | 0.016933
[7 empty] | 1924 | 10.1913 | [20 empty] | -0.091942 | 5.2648 | 15.8947 | -31.319 | 46.0439 | -50.1721 | 40.4989 | -20.5532 | 4.6253 | [12 empty] | 1.2716 | 3.3174 | 7.5997 | -2.0144 | 0.016968
[7 empty] | 1925 | 10.1979 | [20 empty] | -0.091942 | 5.2675 | 15.9113 | -31.3679 | 46.1397 | -50.3026 | 40.6253 | -20.6281 | 4.6446 | [12 empty] | 1.2716 | 3.3192 | 7.6076 | -2.0175 | 0.017004
[7 empty] | 1926 | 10.2044 | [20 empty] | -0.091942 | 5.2703 | 15.9278 | -31.4168 | 46.2357 | -50.4334 | 40.7521 | -20.7032 | 4.6639 | [12 empty] | 1.2716 | 3.3209 | 7.6155 | -2.0207 | 0.017039
[7 empty] | 1927 | 10.2109 | [20 empty] | -0.091942 | 5.273 | 15.9443 | -31.4657 | 46.3318 | -50.5645 | 40.8792 | -20.7786 | 4.6833 | [12 empty] | 1.2716 | 3.3226 | 7.6234 | -2.0238 | 0.017074
[7 empty] | 1928 | 10.2174 | [20 empty] | -0.091942 | 5.2758 | 15.9609 | -31.5148 | 46.428 | -50.6958 | 41.0067 | -20.8542 | 4.7028 | [12 empty] | 1.2716 | 3.3243 | 7.6313 | -2.027 | 0.01711
[7 empty] | 1929 | 10.2239 | [20 empty] | -0.091942 | 5.2785 | 15.9774 | -31.5638 | 46.5244 | -50.8274 | 41.1345 | -20.93 | 4.7223 | [12 empty] | 1.2716 | 3.326 | 7.6392 | -2.0301 | 0.017145
[7 empty] | 1930 | 10.2304 | [20 empty] | -0.091942 | 5.2812 | 15.994 | -31.6129 | 46.621 | -50.9593 | 41.2626 | -21.0061 | 4.742 | [12 empty] | 1.2716 | 3.3278 | 7.6472 | -2.0333 | 0.017181
[7 empty] | 1931 | 10.237 | [20 empty] | -0.091942 | 5.284 | 16.0106 | -31.6621 | 46.7177 | -51.0915 | 41.391 | -21.0824 | 4.7616 | [12 empty] | 1.2716 | 3.3295 | 7.6551 | -2.0365 | 0.017217
[7 empty] | 1932 | 10.2435 | [20 empty] | -0.091942 | 5.2867 | 16.0272 | -31.7113 | 46.8145 | -51.2239 | 41.5198 | -21.159 | 4.7814 | [12 empty] | 1.2716 | 3.3312 | 7.663 | -2.0396 | 0.017252
[7 empty] | 1933 | 10.25 | [20 empty] | -0.091942 | 5.2894 | 16.0438 | -31.7606 | 46.9115 | -51.3566 | 41.6489 | -21.2357 | 4.8012 | [12 empty] | 1.2716 | 3.3329 | 7.671 | -2.0428 | 0.017288
[7 empty] | 1934 | 10.2565 | [20 empty] | -0.091942 | 5.2922 | 16.0604 | -31.8099 | 47.0087 | -51.4896 | 41.7784 | -21.3128 | 4.8212 | [12 empty] | 1.2716 | 3.3347 | 7.6789 | -2.046 | 0.017324
[7 empty] | 1935 | 10.2631 | [20 empty] | -0.091942 | 5.2949 | 16.077 | -31.8593 | 47.106 | -51.6228 | 41.9081 | -21.39 | 4.8411 | [12 empty] | 1.2716 | 3.3364 | 7.6868 | -2.0492 | 0.01736
[7 empty] | 1936 | 10.2696 | [20 empty] | -0.091942 | 5.2976 | 16.0936 | -31.9087 | 47.2034 | -51.7564 | 42.0383 | -21.4675 | 4.8612 | [12 empty] | 1.2716 | 3.3381 | 7.6948 | -2.0523 | 0.017396
[7 empty] | 1937 | 10.2761 | [20 empty] | -0.091942 | 5.3004 | 16.1102 | -31.9582 | 47.301 | -51.8902 | 42.1687 | -21.5453 | 4.8813 | [12 empty] | 1.2716 | 3.3398 | 7.7027 | -2.0555 | 0.017432
[7 empty] | 1938 | 10.2827 | [20 empty] | -0.091942 | 5.3031 | 16.1269 | -32.0077 | 47.3988 | -52.0242 | 42.2995 | -21.6232 | 4.9015 | [12 empty] | 1.2716 | 3.3416 | 7.7107 | -2.0587 | 0.017468
[7 empty] | 1939 | 10.2892 | [20 empty] | -0.091942 | 5.3059 | 16.1435 | -32.0573 | 47.4967 | -52.1586 | 42.4306 | -21.7015 | 4.9218 | [12 empty] | 1.2716 | 3.3433 | 7.7186 | -2.0619 | 0.017504
[7 empty] | 1940 | 10.2957 | [20 empty] | -0.091942 | 5.3086 | 16.1602 | -32.1069 | 47.5947 | -52.2932 | 42.5621 | -21.7799 | 4.9421 | [12 empty] | 1.2716 | 3.345 | 7.7266 | -2.0651 | 0.01754
[7 empty] | 1941 | 10.3022 | [20 empty] | -0.091942 | 5.3113 | 16.1769 | -32.1566 | 47.6929 | -52.4282 | 42.6939 | -21.8586 | 4.9625 | [12 empty] | 1.2716 | 3.3467 | 7.7346 | -2.0683 | 0.017576
[7 empty] | 1942 | 10.3088 | [20 empty] | -0.091942 | 5.3141 | 16.1935 | -32.2063 | 47.7913 | -52.5634 | 42.826 | -21.9376 | 4.983 | [12 empty] | 1.2716 | 3.3485 | 7.7425 | -2.0715 | 0.017612
[7 empty] | 1943 | 10.3153 | [20 empty] | -0.091942 | 5.3168 | 16.2102 | -32.2561 | 47.8898 | -52.6988 | 42.9585 | -22.0168 | 5.0036 | [12 empty] | 1.2716 | 3.3502 | 7.7505 | -2.0747 | 0.017649
[7 empty] | 1944 | 10.3218 | [20 empty] | -0.091942 | 5.3195 | 16.2269 | -32.3059 | 47.9885 | -52.8346 | 43.0913 | -22.0962 | 5.0242 | [12 empty] | 1.2716 | 3.3519 | 7.7585 | -2.0779 | 0.017685
[7 empty] | 1945 | 10.3284 | [20 empty] | -0.091942 | 5.3223 | 16.2436 | -32.3558 | 48.0873 | -52.9706 | 43.2245 | -22.1759 | 5.0449 | [12 empty] | 1.2716 | 3.3536 | 7.7665 | -2.0811 | 0.017721
[7 empty] | 1946 | 10.3349 | [20 empty] | -0.091942 | 5.325 | 16.2603 | -32.4057 | 48.1863 | -53.1069 | 43.358 | -22.2559 | 5.0657 | [12 empty] | 1.2716 | 3.3554 | 7.7745 | -2.0843 | 0.017758
[7 empty] | 1947 | 10.3415 | [20 empty] | -0.091942 | 5.3277 | 16.277 | -32.4557 | 48.2854 | -53.2435 | 43.4919 | -22.336 | 5.0866 | [12 empty] | 1.2716 | 3.3571 | 7.7825 | -2.0875 | 0.017794
[7 empty] | 1948 | 10.348 | [20 empty] | -0.091942 | 5.3305 | 16.2937 | -32.5057 | 48.3847 | -53.3804 | 43.6261 | -22.4165 | 5.1075 | [12 empty] | 1.2716 | 3.3588 | 7.7905 | -2.0907 | 0.017831
[7 empty] | 1949 | 10.3545 | [20 empty] | -0.091942 | 5.3332 | 16.3105 | -32.5558 | 48.4841 | -53.5175 | 43.7606 | -22.4971 | 5.1285 | [12 empty] | 1.2716 | 3.3605 | 7.7985 | -2.094 | 0.017868
[7 empty] | 1950 | 10.3611 | [20 empty] | -0.091942 | 5.336 | 16.3272 | -32.6059 | 48.5837 | -53.655 | 43.8955 | -22.5781 | 5.1496 | [12 empty] | 1.2716 | 3.3623 | 7.8065 | -2.0972 | 0.017904
[7 empty] | 1951 | 10.3676 | [20 empty] | -0.091942 | 5.3387 | 16.344 | -32.6561 | 48.6834 | -53.7927 | 44.0308 | -22.6592 | 5.1708 | [12 empty] | 1.2716 | 3.364 | 7.8145 | -2.1004 | 0.017941
[7 empty] | 1952 | 10.3742 | [20 empty] | -0.091942 | 5.3414 | 16.3607 | -32.7064 | 48.7833 | -53.9307 | 44.1663 | -22.7407 | 5.192 | [12 empty] | 1.2716 | 3.3657 | 7.8225 | -2.1036 | 0.017978
[7 empty] | 1953 | 10.3807 | [20 empty] | -0.091942 | 5.3442 | 16.3775 | -32.7567 | 48.8833 | -54.069 | 44.3023 | -22.8223 | 5.2134 | [12 empty] | 1.2716 | 3.3674 | 7.8305 | -2.1069 | 0.018015
[7 empty] | 1954 | 10.3872 | [20 empty] | -0.091942 | 5.3469 | 16.3943 | -32.807 | 48.9835 | -54.2075 | 44.4386 | -22.9043 | 5.2347 | [12 empty] | 1.2716 | 3.3692 | 7.8385 | -2.1101 | 0.018052
[7 empty] | 1955 | 10.3938 | [20 empty] | -0.091942 | 5.3496 | 16.4111 | -32.8574 | 49.0839 | -54.3464 | 44.5752 | -22.9864 | 5.2562 | [12 empty] | 1.2716 | 3.3709 | 7.8466 | -2.1134 | 0.018089
[7 empty] | 1956 | 10.4003 | [20 empty] | -0.091942 | 5.3524 | 16.4278 | -32.9078 | 49.1844 | -54.4855 | 44.7122 | -23.0689 | 5.2778 | [12 empty] | 1.2716 | 3.3726 | 7.8546 | -2.1166 | 0.018126
[7 empty] | 1957 | 10.4069 | [20 empty] | -0.091942 | 5.3551 | 16.4446 | -32.9583 | 49.2851 | -54.625 | 44.8495 | -23.1516 | 5.2994 | [12 empty] | 1.2716 | 3.3743 | 7.8626 | -2.1198 | 0.018163
[7 empty] | 1958 | 10.4134 | [20 empty] | -0.091942 | 5.3578 | 16.4615 | -33.0089 | 49.3859 | -54.7647 | 44.9872 | -23.2345 | 5.3211 | [12 empty] | 1.2716 | 3.3761 | 7.8707 | -2.1231 | 0.0182
[7 empty] | 1959 | 10.42 | [20 empty] | -0.091942 | 5.3606 | 16.4783 | -33.0595 | 49.4868 | -54.9047 | 45.1252 | -23.3177 | 5.3429 | [12 empty] | 1.2716 | 3.3778 | 7.8787 | -2.1263 | 0.018237
[7 empty] | 1960 | 10.4265 | [20 empty] | -0.091942 | 5.3633 | 16.4951 | -33.1101 | 49.588 | -55.0449 | 45.2636 | -23.4011 | 5.3647 | [12 empty] | 1.2716 | 3.3795 | 7.8867 | -2.1296 | 0.018274
[7 empty] | 1961 | 10.4331 | [20 empty] | -0.091942 | 5.3661 | 16.5119 | -33.1609 | 49.6892 | -55.1855 | 45.4023 | -23.4848 | 5.3867 | [12 empty] | 1.2716 | 3.3812 | 7.8948 | -2.1329 | 0.018312
[7 empty] | 1962 | 10.4396 | [20 empty] | -0.091942 | 5.3688 | 16.5288 | -33.2116 | 49.7907 | -55.3263 | 45.5414 | -23.5688 | 5.4087 | [12 empty] | 1.2716 | 3.3829 | 7.9028 | -2.1361 | 0.018349
[7 empty] | 1963 | 10.4462 | [20 empty] | -0.091942 | 5.3715 | 16.5456 | -33.2624 | 49.8923 | -55.4675 | 45.6809 | -23.653 | 5.4308 | [12 empty] | 1.2716 | 3.3847 | 7.9109 | -2.1394 | 0.018387
[7 empty] | 1964 | 10.4527 | [20 empty] | -0.091942 | 5.3743 | 16.5625 | -33.3133 | 49.994 | -55.6089 | 45.8207 | -23.7375 | 5.4529 | [12 empty] | 1.2716 | 3.3864 | 7.919 | -2.1427 | 0.018424
[7 empty] | 1965 | 10.4593 | [20 empty] | -0.091942 | 5.377 | 16.5794 | -33.3642 | 50.0959 | -55.7506 | 45.9608 | -23.8222 | 5.4752 | [12 empty] | 1.2716 | 3.3881 | 7.927 | -2.1459 | 0.018462
[7 empty] | 1966 | 10.4658 | [20 empty] | -0.091942 | 5.3797 | 16.5962 | -33.4152 | 50.198 | -55.8926 | 46.1013 | -23.9072 | 5.4975 | [12 empty] | 1.2716 | 3.3898 | 7.9351 | -2.1492 | 0.018499
[7 empty] | 1967 | 10.4724 | [20 empty] | -0.091942 | 5.3825 | 16.6131 | -33.4662 | 50.3002 | -56.0349 | 46.2422 | -23.9925 | 5.5199 | [12 empty] | 1.2716 | 3.3916 | 7.9432 | -2.1525 | 0.018537
[7 empty] | 1968 | 10.4789 | [20 empty] | -0.091942 | 5.3852 | 16.63 | -33.5172 | 50.4025 | -56.1775 | 46.3835 | -24.078 | 5.5424 | [12 empty] | 1.2716 | 3.3933 | 7.9513 | -2.1558 | 0.018575
[7 empty] | 1969 | 10.4855 | [20 empty] | -0.091942 | 5.3879 | 16.6469 | -33.5684 | 50.5051 | -56.3204 | 46.525 | -24.1638 | 5.565 | [12 empty] | 1.2716 | 3.395 | 7.9593 | -2.1591 | 0.018612
[7 empty] | 1970 | 10.4921 | [20 empty] | -0.091942 | 5.3907 | 16.6639 | -33.6195 | 50.6077 | -56.4635 | 46.667 | -24.2498 | 5.5876 | [12 empty] | 1.2716 | 3.3967 | 7.9674 | -2.1624 | 0.01865
[7 empty] | 1971 | 10.4986 | [20 empty] | -0.091942 | 5.3934 | 16.6808 | -33.6707 | 50.7106 | -56.607 | 46.8093 | -24.3361 | 5.6104 | [12 empty] | 1.2716 | 3.3985 | 7.9755 | -2.1657 | 0.018688
[7 empty] | 1972 | 10.5052 | [20 empty] | -0.091942 | 5.3962 | 16.6977 | -33.722 | 50.8136 | -56.7507 | 46.952 | -24.4226 | 5.6332 | [12 empty] | 1.2716 | 3.4002 | 7.9836 | -2.169 | 0.018726
[7 empty] | 1973 | 10.5117 | [20 empty] | -0.091942 | 5.3989 | 16.7146 | -33.7734 | 50.9167 | -56.8948 | 47.095 | -24.5095 | 5.6561 | [12 empty] | 1.2716 | 3.4019 | 7.9917 | -2.1723 | 0.018764
[7 empty] | 1974 | 10.5183 | [20 empty] | -0.091942 | 5.4016 | 16.7316 | -33.8247 | 51.02 | -57.0391 | 47.2384 | -24.5966 | 5.6791 | [12 empty] | 1.2716 | 3.4036 | 7.9998 | -2.1756 | 0.018802
[7 empty] | 1975 | 10.5249 | [20 empty] | -0.091942 | 5.4044 | 16.7485 | -33.8762 | 51.1235 | -57.1837 | 47.3822 | -24.6839 | 5.7021 | [12 empty] | 1.2716 | 3.4054 | 8.0079 | -2.1789 | 0.01884
[7 empty] | 1976 | 10.5314 | [20 empty] | -0.091942 | 5.4071 | 16.7655 | -33.9276 | 51.2271 | -57.3286 | 47.5263 | -24.7715 | 5.7253 | [12 empty] | 1.2716 | 3.4071 | 8.016 | -2.1822 | 0.018878
[7 empty] | 1977 | 10.538 | [20 empty] | -0.091942 | 5.4098 | 16.7825 | -33.9792 | 51.3309 | -57.4739 | 47.6708 | -24.8594 | 5.7485 | [12 empty] | 1.2716 | 3.4088 | 8.0241 | -2.1855 | 0.018917
[7 empty] | 1978 | 10.5446 | [20 empty] | -0.091942 | 5.4126 | 16.7995 | -34.0308 | 51.4348 | -57.6194 | 47.8157 | -24.9476 | 5.7718 | [12 empty] | 1.2716 | 3.4105 | 8.0323 | -2.1888 | 0.018955
[7 empty] | 1979 | 10.5511 | [20 empty] | -0.091942 | 5.4153 | 16.8165 | -34.0824 | 51.5389 | -57.7652 | 47.9609 | -25.036 | 5.7952 | [12 empty] | 1.2716 | 3.4123 | 8.0404 | -2.1921 | 0.018993
[7 empty] | 1980 | 10.5577 | [20 empty] | -0.091942 | 5.418 | 16.8335 | -34.1341 | 51.6431 | -57.9112 | 48.1065 | -25.1247 | 5.8186 | [12 empty] | 1.2716 | 3.414 | 8.0485 | -2.1955 | 0.019032
[7 empty] | 1981 | 10.5643 | [20 empty] | -0.091942 | 5.4208 | 16.8505 | -34.1858 | 51.7476 | -58.0576 | 48.2525 | -25.2136 | 5.8422 | [12 empty] | 1.2716 | 3.4157 | 8.0567 | -2.1988 | 0.01907
[7 empty] | 1982 | 10.5708 | [20 empty] | -0.091942 | 5.4235 | 16.8675 | -34.2376 | 51.8521 | -58.2043 | 48.3988 | -25.3029 | 5.8658 | [12 empty] | 1.2716 | 3.4174 | 8.0648 | -2.2021 | 0.019109
[7 empty] | 1983 | 10.5774 | [20 empty] | -0.091942 | 5.4263 | 16.8845 | -34.2895 | 51.9568 | -58.3513 | 48.5455 | -25.3924 | 5.8895 | [12 empty] | 1.2716 | 3.4192 | 8.0729 | -2.2055 | 0.019147
[7 empty] | 1984 | 10.584 | [20 empty] | -0.091942 | 5.429 | 16.9015 | -34.3414 | 52.0617 | -58.4986 | 48.6926 | -25.4821 | 5.9133 | [12 empty] | 1.2716 | 3.4209 | 8.0811 | -2.2088 | 0.019186
[7 empty] | 1985 | 10.5905 | [20 empty] | -0.091942 | 5.4317 | 16.9186 | -34.3933 | 52.1668 | -58.6462 | 48.84 | -25.5722 | 5.9372 | [12 empty] | 1.2716 | 3.4226 | 8.0892 | -2.2121 | 0.019225
[7 empty] | 1986 | 10.5971 | [20 empty] | -0.091942 | 5.4345 | 16.9356 | -34.4454 | 52.272 | -58.794 | 48.9878 | -25.6625 | 5.9612 | [12 empty] | 1.2716 | 3.4243 | 8.0974 | -2.2155 | 0.019264
[7 empty] | 1987 | 10.6037 | [20 empty] | -0.091942 | 5.4372 | 16.9527 | -34.4974 | 52.3773 | -58.9422 | 49.136 | -25.7531 | 5.9852 | [12 empty] | 1.2716 | 3.4261 | 8.1055 | -2.2188 | 0.019302
[7 empty] | 1988 | 10.6102 | [20 empty] | -0.091942 | 5.4399 | 16.9698 | -34.5495 | 52.4829 | -59.0907 | 49.2846 | -25.844 | 6.0094 | [12 empty] | 1.2716 | 3.4278 | 8.1137 | -2.2222 | 0.019341
[7 empty] | 1989 | 10.6168 | [20 empty] | -0.091942 | 5.4427 | 16.9868 | -34.6017 | 52.5885 | -59.2394 | 49.4335 | -25.9351 | 6.0336 | [12 empty] | 1.2716 | 3.4295 | 8.1219 | -2.2255 | 0.01938
[7 empty] | 1990 | 10.6234 | [20 empty] | -0.091942 | 5.4454 | 17.0039 | -34.6539 | 52.6944 | -59.3885 | 49.5828 | -26.0265 | 6.0579 | [12 empty] | 1.2716 | 3.4312 | 8.13 | -2.2289 | 0.019419
[7 empty] | 1991 | 10.63 | [20 empty] | -0.091942 | 5.4481 | 17.021 | -34.7062 | 52.8004 | -59.5379 | 49.7325 | -26.1182 | 6.0823 | [12 empty] | 1.2716 | 3.433 | 8.1382 | -2.2323 | 0.019458
[7 empty] | 1992 | 10.6365 | [20 empty] | -0.091942 | 5.4509 | 17.0381 | -34.7585 | 52.9065 | -59.6875 | 49.8825 | -26.2102 | 6.1068 | [12 empty] | 1.2716 | 3.4347 | 8.1464 | -2.2356 | 0.019497
[7 empty] | 1993 | 10.6431 | [20 empty] | -0.091942 | 5.4536 | 17.0552 | -34.8109 | 53.0128 | -59.8375 | 50.033 | -26.3024 | 6.1314 | [12 empty] | 1.2716 | 3.4364 | 8.1546 | -2.239 | 0.019537
[7 empty] | 1994 | 10.6497 | [20 empty] | -0.091942 | 5.4564 | 17.0723 | -34.8633 | 53.1193 | -59.9878 | 50.1838 | -26.3949 | 6.156 | [12 empty] | 1.2716 | 3.4381 | 8.1627 | -2.2424 | 0.019576
[7 empty] | 1995 | 10.6563 | [20 empty] | -0.091942 | 5.4591 | 17.0895 | -34.9158 | 53.226 | -60.1383 | 50.335 | -26.4877 | 6.1808 | [12 empty] | 1.2716 | 3.4398 | 8.1709 | -2.2457 | 0.019615
[7 empty] | 1996 | 10.6628 | [20 empty] | -0.091942 | 5.4618 | 17.1066 | -34.9683 | 53.3328 | -60.2892 | 50.4866 | -26.5808 | 6.2056 | [12 empty] | 1.2716 | 3.4416 | 8.1791 | -2.2491 | 0.019654
[7 empty] | 1997 | 10.6694 | [20 empty] | -0.091942 | 5.4646 | 17.1238 | -35.0209 | 53.4397 | -60.4404 | 50.6385 | -26.6742 | 6.2305 | [12 empty] | 1.2716 | 3.4433 | 8.1873 | -2.2525 | 0.019694
[7 empty] | 1998 | 10.676 | [20 empty] | -0.091942 | 5.4673 | 17.1409 | -35.0735 | 53.5468 | -60.5919 | 50.7908 | -26.7678 | 6.2555 | [12 empty] | 1.2716 | 3.445 | 8.1955 | -2.2559 | 0.019733
[7 empty] | 1999 | 10.6826 | [20 empty] | -0.091942 | 5.47 | 17.1581 | -35.1262 | 53.6541 | -60.7437 | 50.9436 | -26.8617 | 6.2806 | [12 empty] | 1.2716 | 3.4467 | 8.2037 | -2.2593 | 0.019773
[7 empty] | 2000 | 10.6892 | [20 empty] | -0.091942 | 5.4728 | 17.1752 | -35.179 | 53.7616 | -60.8957 | 51.0967 | -26.9559 | 6.3058 | [12 empty] | 1.2716 | 3.4485 | 8.2119 | -2.2627 | 0.019812
[7 empty] | 2001 | 10.6957 | [20 empty] | -0.091942 | 5.4755 | 17.1924 | -35.2317 | 53.8692 | -61.0481 | 51.2501 | -27.0504 | 6.331 | [12 empty] | 1.2716 | 3.4502 | 8.2201 | -2.2661 | 0.019852
[7 empty] | 2002 | 10.7023 | [20 empty] | -0.091942 | 5.4782 | 17.2096 | -35.2846 | 53.9769 | -61.2008 | 51.404 | -27.1452 | 6.3564 | [12 empty] | 1.2716 | 3.4519 | 8.2284 | -2.2695 | 0.019892
[7 empty] | 2003 | 10.7089 | [20 empty] | -0.091942 | 5.481 | 17.2268 | -35.3375 | 54.0849 | -61.3538 | 51.5583 | -27.2402 | 6.3818 | [12 empty] | 1.2716 | 3.4536 | 8.2366 | -2.2729 | 0.019932
[7 empty] | 2004 | 10.7155 | [20 empty] | -0.091942 | 5.4837 | 17.244 | -35.3904 | 54.193 | -61.5071 | 51.7129 | -27.3356 | 6.4074 | [12 empty] | 1.2716 | 3.4554 | 8.2448 | -2.2763 | 0.019971
[7 empty] | 2005 | 10.7221 | [20 empty] | -0.091942 | 5.4865 | 17.2612 | -35.4435 | 54.3012 | -61.6608 | 51.8679 | -27.4312 | 6.433 | [12 empty] | 1.2716 | 3.4571 | 8.253 | -2.2797 | 0.020011
[7 empty] | 2006 | 10.7287 | [20 empty] | -0.091942 | 5.4892 | 17.2785 | -35.4965 | 54.4096 | -61.8147 | 52.0233 | -27.5271 | 6.4587 | [12 empty] | 1.2716 | 3.4588 | 8.2613 | -2.2831 | 0.020051
[7 empty] | 2007 | 10.7353 | [20 empty] | -0.091942 | 5.4919 | 17.2957 | -35.5496 | 54.5182 | -61.9689 | 52.1791 | -27.6233 | 6.4845 | [12 empty] | 1.2716 | 3.4605 | 8.2695 | -2.2865 | 0.020091
[7 empty] | 2008 | 10.7418 | [20 empty] | -0.091942 | 5.4947 | 17.3129 | -35.6028 | 54.6269 | -62.1234 | 52.3353 | -27.7198 | 6.5104 | [12 empty] | 1.2716 | 3.4623 | 8.2778 | -2.2899 | 0.020131
[7 empty] | 2009 | 10.7484 | [20 empty] | -0.091942 | 5.4974 | 17.3302 | -35.656 | 54.7358 | -62.2783 | 52.4919 | -27.8166 | 6.5364 | [12 empty] | 1.2716 | 3.464 | 8.286 | -2.2934 | 0.020172
[7 empty] | 2010 | 10.755 | [20 empty] | -0.091942 | 5.5001 | 17.3474 | -35.7093 | 54.8449 | -62.4334 | 52.6488 | -27.9137 | 6.5625 | [12 empty] | 1.2716 | 3.4657 | 8.2943 | -2.2968 | 0.020212
[7 empty] | 2011 | 10.7616 | [20 empty] | -0.091942 | 5.5029 | 17.3647 | -35.7626 | 54.9541 | -62.5889 | 52.8062 | -28.011 | 6.5886 | [12 empty] | 1.2716 | 3.4674 | 8.3025 | -2.3002 | 0.020252
[7 empty] | 2012 | 10.7682 | [20 empty] | -0.091942 | 5.5056 | 17.382 | -35.816 | 55.0635 | -62.7447 | 52.964 | -28.1087 | 6.6149 | [12 empty] | 1.2716 | 3.4692 | 8.3108 | -2.3036 | 0.020292
[7 empty] | 2013 | 10.7748 | [20 empty] | -0.091942 | 5.5083 | 17.3992 | -35.8694 | 55.1731 | -62.9008 | 53.1221 | -28.2066 | 6.6412 | [12 empty] | 1.2716 | 3.4709 | 8.319 | -2.3071 | 0.020333
[7 empty] | 2014 | 10.7814 | [20 empty] | -0.091942 | 5.5111 | 17.4165 | -35.9229 | 55.2828 | -63.0571 | 53.2806 | -28.3048 | 6.6677 | [12 empty] | 1.2716 | 3.4726 | 8.3273 | -2.3105 | 0.020373
[7 empty] | 2015 | 10.788 | [20 empty] | -0.091942 | 5.5138 | 17.4338 | -35.9764 | 55.3927 | -63.2138 | 53.4396 | -28.4034 | 6.6942 | [12 empty] | 1.2716 | 3.4743 | 8.3356 | -2.314 | 0.020414
[7 empty] | 2016 | 10.7946 | [20 empty] | -0.091942 | 5.5166 | 17.4511 | -36.03 | 55.5027 | -63.3709 | 53.5989 | -28.5022 | 6.7208 | [12 empty] | 1.2716 | 3.4761 | 8.3439 | -2.3174 | 0.020454
[7 empty] | 2017 | 10.8012 | [20 empty] | -0.091942 | 5.5193 | 17.4685 | -36.0837 | 55.6129 | -63.5282 | 53.7586 | -28.6013 | 6.7476 | [12 empty] | 1.2716 | 3.4778 | 8.3521 | -2.3209 | 0.020495
[7 empty] | 2018 | 10.8078 | [20 empty] | -0.091942 | 5.522 | 17.4858 | -36.1374 | 55.7233 | -63.6858 | 53.9187 | -28.7007 | 6.7744 | [12 empty] | 1.2716 | 3.4795 | 8.3604 | -2.3243 | 0.020535
[7 empty] | 2019 | 10.8144 | [20 empty] | -0.091942 | 5.5248 | 17.5031 | -36.1911 | 55.8338 | -63.8438 | 54.0792 | -28.8004 | 6.8013 | [12 empty] | 1.2716 | 3.4812 | 8.3687 | -2.3278 | 0.020576
[7 empty] | 2020 | 10.821 | [20 empty] | -0.091942 | 5.5275 | 17.5205 | -36.2449 | 55.9445 | -64.002 | 54.2401 | -28.9004 | 6.8283 | [12 empty] | 1.2716 | 3.483 | 8.377 | -2.3312 | 0.020617
[7 empty] | 2021 | 10.8275 | [20 empty] | -0.091942 | 5.5302 | 17.5378 | -36.2988 | 56.0554 | -64.1606 | 54.4014 | -29.0007 | 6.8553 | [12 empty] | 1.2716 | 3.4847 | 8.3853 | -2.3347 | 0.020658
[7 empty] | 2022 | 10.8341 | [20 empty] | -0.091942 | 5.533 | 17.5552 | -36.3527 | 56.1664 | -64.3195 | 54.5632 | -29.1013 | 6.8825 | [12 empty] | 1.2716 | 3.4864 | 8.3936 | -2.3382 | 0.020699
[7 empty] | 2023 | 10.8407 | [20 empty] | -0.091942 | 5.5357 | 17.5725 | -36.4066 | 56.2776 | -64.4787 | 54.7253 | -29.2022 | 6.9098 | [12 empty] | 1.2716 | 3.4881 | 8.4019 | -2.3416 | 0.02074
[7 empty] | 2024 | 10.8473 | [20 empty] | -0.091942 | 5.5384 | 17.5899 | -36.4607 | 56.3889 | -64.6382 | 54.8878 | -29.3034 | 6.9372 | [12 empty] | 1.2716 | 3.4899 | 8.4102 | -2.3451 | 0.020781
[7 empty] | 2025 | 10.8539 | [20 empty] | -0.091942 | 5.5412 | 17.6073 | -36.5147 | 56.5005 | -64.7981 | 55.0507 | -29.4049 | 6.9646 | [12 empty] | 1.2716 | 3.4916 | 8.4185 | -2.3486 | 0.020822
[7 empty] | 2026 | 10.8605 | [20 empty] | -0.091942 | 5.5439 | 17.6247 | -36.5688 | 56.6122 | -64.9582 | 55.214 | -29.5067 | 6.9922 | [12 empty] | 1.2716 | 3.4933 | 8.4268 | -2.3521 | 0.020863
[7 empty] | 2027 | 10.8671 | [20 empty] | -0.091942 | 5.5467 | 17.6421 | -36.623 | 56.724 | -65.1187 | 55.3777 | -29.6088 | 7.0199 | [12 empty] | 1.2716 | 3.495 | 8.4352 | -2.3556 | 0.020904
[7 empty] | 2028 | 10.8737 | [20 empty] | -0.091942 | 5.5494 | 17.6595 | -36.6772 | 56.836 | -65.2795 | 55.5418 | -29.7112 | 7.0476 | [12 empty] | 1.2716 | 3.4967 | 8.4435 | -2.359 | 0.020945
[7 empty] | 2029 | 10.8804 | [20 empty] | -0.091942 | 5.5521 | 17.6769 | -36.7315 | 56.9482 | -65.4406 | 55.7064 | -29.8139 | 7.0755 | [12 empty] | 1.2716 | 3.4985 | 8.4518 | -2.3625 | 0.020987
[7 empty] | 2030 | 10.887 | [20 empty] | -0.091942 | 5.5549 | 17.6944 | -36.7859 | 57.0606 | -65.602 | 55.8713 | -29.9169 | 7.1034 | [12 empty] | 1.2716 | 3.5002 | 8.4601 | -2.366 | 0.021028
[7 empty] | 2031 | 10.8936 | [20 empty] | -0.091942 | 5.5576 | 17.7118 | -36.8403 | 57.1731 | -65.7638 | 56.0366 | -30.0202 | 7.1315 | [12 empty] | 1.2716 | 3.5019 | 8.4685 | -2.3695 | 0.02107
[7 empty] | 2032 | 10.9002 | [20 empty] | -0.091942 | 5.5603 | 17.7292 | -36.8947 | 57.2858 | -65.9258 | 56.2024 | -30.1238 | 7.1596 | [12 empty] | 1.2716 | 3.5036 | 8.4768 | -2.373 | 0.021111
[7 empty] | 2033 | 10.9068 | [20 empty] | -0.091942 | 5.5631 | 17.7467 | -36.9492 | 57.3986 | -66.0882 | 56.3685 | -30.2278 | 7.1878 | [12 empty] | 1.2716 | 3.5054 | 8.4852 | -2.3765 | 0.021153
[7 empty] | 2034 | 10.9134 | [20 empty] | -0.091942 | 5.5658 | 17.7642 | -37.0038 | 57.5116 | -66.2509 | 56.5351 | -30.332 | 7.2162 | [12 empty] | 1.2716 | 3.5071 | 8.4935 | -2.38 | 0.021194
[7 empty] | 2035 | 10.92 | [20 empty] | -0.091942 | 5.5685 | 17.7816 | -37.0584 | 57.6248 | -66.4139 | 56.7021 | -30.4365 | 7.2446 | [12 empty] | 1.2716 | 3.5088 | 8.5019 | -2.3836 | 0.021236
[7 empty] | 2036 | 10.9266 | [20 empty] | -0.091942 | 5.5713 | 17.7991 | -37.113 | 57.7382 | -66.5772 | 56.8695 | -30.5414 | 7.2731 | [12 empty] | 1.2716 | 3.5105 | 8.5102 | -2.3871 | 0.021278
[7 empty] | 2037 | 10.9332 | [20 empty] | -0.091942 | 5.574 | 17.8166 | -37.1677 | 57.8517 | -66.7409 | 57.0373 | -30.6465 | 7.3018 | [12 empty] | 1.2716 | 3.5123 | 8.5186 | -2.3906 | 0.02132
[7 empty] | 2038 | 10.9398 | [20 empty] | -0.091942 | 5.5768 | 17.8341 | -37.2225 | 57.9654 | -66.9049 | 57.2055 | -30.752 | 7.3305 | [12 empty] | 1.2716 | 3.514 | 8.527 | -2.3941 | 0.021362
[7 empty] | 2039 | 10.9464 | [20 empty] | -0.091942 | 5.5795 | 17.8516 | -37.2773 | 58.0792 | -67.0692 | 57.3741 | -30.8578 | 7.3593 | [12 empty] | 1.2716 | 3.5157 | 8.5353 | -2.3976 | 0.021404
[7 empty] | 2040 | 10.953 | [20 empty] | -0.091942 | 5.5822 | 17.8691 | -37.3322 | 58.1933 | -67.2338 | 57.5431 | -30.9639 | 7.3882 | [12 empty] | 1.2716 | 3.5174 | 8.5437 | -2.4012 | 0.021446
[7 empty] | 2041 | 10.9596 | [20 empty] | -0.091942 | 5.585 | 17.8866 | -37.3871 | 58.3074 | -67.3988 | 57.7126 | -31.0703 | 7.4173 | [12 empty] | 1.2716 | 3.5192 | 8.5521 | -2.4047 | 0.021488
[7 empty] | 2042 | 10.9663 | [20 empty] | -0.091942 | 5.5877 | 17.9042 | -37.4421 | 58.4218 | -67.564 | 57.8825 | -31.177 | 7.4464 | [12 empty] | 1.2716 | 3.5209 | 8.5605 | -2.4082 | 0.02153
[7 empty] | 2043 | 10.9729 | [20 empty] | -0.091942 | 5.5904 | 17.9217 | -37.4971 | 58.5363 | -67.7296 | 58.0527 | -31.284 | 7.4756 | [12 empty] | 1.2716 | 3.5226 | 8.5688 | -2.4118 | 0.021572
[7 empty] | 2044 | 10.9795 | [20 empty] | -0.091942 | 5.5932 | 17.9393 | -37.5522 | 58.651 | -67.8956 | 58.2234 | -31.3914 | 7.5049 | [12 empty] | 1.2716 | 3.5243 | 8.5772 | -2.4153 | 0.021614
[7 empty] | 2045 | 10.9861 | [20 empty] | -0.091942 | 5.5959 | 17.9568 | -37.6074 | 58.7659 | -68.0618 | 58.3946 | -31.4991 | 7.5343 | [12 empty] | 1.2716 | 3.5261 | 8.5856 | -2.4189 | 0.021657
[7 empty] | 2046 | 10.9927 | [20 empty] | -0.091942 | 5.5986 | 17.9744 | -37.6626 | 58.8809 | -68.2284 | 58.5661 | -31.607 | 7.5639 | [12 empty] | 1.2716 | 3.5278 | 8.594 | -2.4224 | 0.021699
[7 empty] | 2047 | 10.9993 | [20 empty] | -0.091942 | 5.6014 | 17.992 | -37.7178 | 58.9961 | -68.3953 | 58.7381 | -31.7153 | 7.5935 | [12 empty] | 1.2716 | 3.5295 | 8.6024 | -2.426 | 0.021742
[7 empty] | 2048 | 11.0059 | [20 empty] | -0.091942 | 5.6041 | 18.0095 | -37.7731 | 59.1115 | -68.5625 | 58.9104 | -31.8239 | 7.6232 | [12 empty] | 1.2716 | 3.5312 | 8.6108 | -2.4295 | 0.021784
[7 empty] | 2049 | 11.0126 | [20 empty] | -0.091942 | 5.6069 | 18.0271 | -37.8285 | 59.227 | -68.7301 | 59.0832 | -31.9329 | 7.6531 | [12 empty] | 1.2716 | 3.533 | 8.6193 | -2.4331 | 0.021827
[7 empty] | 2050 | 11.0192 | [20 empty] | -0.091942 | 5.6096 | 18.0447 | -37.8839 | 59.3427 | -68.8979 | 59.2565 | -32.0421 | 7.683 | [12 empty] | 1.2716 | 3.5347 | 8.6277 | -2.4366 | 0.021869
[7 empty] | 2051 | 11.0258 | [20 empty] | -0.091942 | 5.6123 | 18.0623 | -37.9394 | 59.4586 | -69.0662 | 59.4301 | -32.1517 | 7.713 | [12 empty] | 1.2716 | 3.5364 | 8.6361 | -2.4402 | 0.021912
[7 empty] | 2052 | 11.0324 | [20 empty] | -0.091942 | 5.6151 | 18.08 | -37.9949 | 59.5746 | -69.2347 | 59.6042 | -32.2616 | 7.7432 | [12 empty] | 1.2716 | 3.5381 | 8.6445 | -2.4438 | 0.021955
[7 empty] | 2053 | 11.039 | [20 empty] | -0.091942 | 5.6178 | 18.0976 | -38.0504 | 59.6908 | -69.4036 | 59.7787 | -32.3718 | 7.7734 | [12 empty] | 1.2716 | 3.5399 | 8.6529 | -2.4474 | 0.021998
[7 empty] | 2054 | 11.0457 | [20 empty] | -0.091942 | 5.6205 | 18.1152 | -38.1061 | 59.8072 | -69.5727 | 59.9536 | -32.4823 | 7.8037 | [12 empty] | 1.2716 | 3.5416 | 8.6614 | -2.4509 | 0.02204
[7 empty] | 2055 | 11.0523 | [20 empty] | -0.091942 | 5.6233 | 18.1329 | -38.1618 | 59.9238 | -69.7423 | 60.1289 | -32.5932 | 7.8342 | [12 empty] | 1.2716 | 3.5433 | 8.6698 | -2.4545 | 0.022083
[7 empty] | 2056 | 11.0589 | [20 empty] | -0.091942 | 5.626 | 18.1505 | -38.2175 | 60.0405 | -69.9121 | 60.3047 | -32.7044 | 7.8647 | [12 empty] | 1.2716 | 3.545 | 8.6782 | -2.4581 | 0.022126
[7 empty] | 2057 | 11.0655 | [20 empty] | -0.091942 | 5.6287 | 18.1682 | -38.2733 | 60.1574 | -70.0823 | 60.4809 | -32.8159 | 7.8954 | [12 empty] | 1.2716 | 3.5468 | 8.6867 | -2.4617 | 0.022169
[7 empty] | 2058 | 11.0722 | [20 empty] | -0.091942 | 5.6315 | 18.1859 | -38.3291 | 60.2745 | -70.2528 | 60.6575 | -32.9277 | 7.9261 | [12 empty] | 1.2716 | 3.5485 | 8.6951 | -2.4653 | 0.022213
[7 empty] | 2059 | 11.0788 | [20 empty] | -0.091942 | 5.6342 | 18.2035 | -38.385 | 60.3917 | -70.4237 | 60.8346 | -33.0399 | 7.957 | [12 empty] | 1.2716 | 3.5502 | 8.7036 | -2.4689 | 0.022256
[7 empty] | 2060 | 11.0854 | [20 empty] | -0.091942 | 5.637 | 18.2212 | -38.441 | 60.5091 | -70.5949 | 61.0121 | -33.1524 | 7.988 | [12 empty] | 1.2716 | 3.5519 | 8.712 | -2.4725 | 0.022299
[7 empty] | 2061 | 11.092 | [20 empty] | -0.091942 | 5.6397 | 18.2389 | -38.497 | 60.6267 | -70.7664 | 61.19 | -33.2652 | 8.019 | [12 empty] | 1.2716 | 3.5536 | 8.7205 | -2.4761 | 0.022342
[7 empty] | 2062 | 11.0987 | [20 empty] | -0.091942 | 5.6424 | 18.2566 | -38.5531 | 60.7444 | -70.9382 | 61.3684 | -33.3784 | 8.0502 | [12 empty] | 1.2716 | 3.5554 | 8.729 | -2.4797 | 0.022386
[7 empty] | 2063 | 11.1053 | [20 empty] | -0.091942 | 5.6452 | 18.2743 | -38.6092 | 60.8624 | -71.1104 | 61.5471 | -33.4918 | 8.0815 | [12 empty] | 1.2716 | 3.5571 | 8.7374 | -2.4833 | 0.022429
[7 empty] | 2064 | 11.1119 | [20 empty] | -0.091942 | 5.6479 | 18.292 | -38.6654 | 60.9805 | -71.2829 | 61.7264 | -33.6056 | 8.1129 | [12 empty] | 1.2716 | 3.5588 | 8.7459 | -2.4869 | 0.022473
[7 empty] | 2065 | 11.1185 | [20 empty] | -0.091942 | 5.6506 | 18.3098 | -38.7216 | 61.0987 | -71.4558 | 61.906 | -33.7198 | 8.1444 | [12 empty] | 1.2716 | 3.5605 | 8.7544 | -2.4905 | 0.022516
[7 empty] | 2066 | 11.1252 | [20 empty] | -0.091942 | 5.6534 | 18.3275 | -38.7779 | 61.2172 | -71.6289 | 62.0861 | -33.8343 | 8.176 | [12 empty] | 1.2716 | 3.5623 | 8.7629 | -2.4941 | 0.02256
[7 empty] | 2067 | 11.1318 | [20 empty] | -0.091942 | 5.6561 | 18.3453 | -38.8342 | 61.3358 | -71.8025 | 62.2666 | -33.9491 | 8.2077 | [12 empty] | 1.2716 | 3.564 | 8.7714 | -2.4978 | 0.022604
[7 empty] | 2068 | 11.1384 | [20 empty] | -0.091942 | 5.6588 | 18.363 | -38.8906 | 61.4546 | -71.9763 | 62.4476 | -34.0642 | 8.2395 | [12 empty] | 1.2716 | 3.5657 | 8.7798 | -2.5014 | 0.022648
[7 empty] | 2069 | 11.1451 | [20 empty] | -0.091942 | 5.6616 | 18.3808 | -38.947 | 61.5735 | -72.1505 | 62.629 | -34.1797 | 8.2715 | [12 empty] | 1.2716 | 3.5674 | 8.7883 | -2.505 | 0.022691
[7 empty] | 2070 | 11.1517 | [20 empty] | -0.091942 | 5.6643 | 18.3986 | -39.0035 | 61.6926 | -72.325 | 62.8108 | -34.2955 | 8.3035 | [12 empty] | 1.2716 | 3.5692 | 8.7968 | -2.5087 | 0.022735
[7 empty] | 2071 | 11.1583 | [20 empty] | -0.091942 | 5.6671 | 18.4163 | -39.0601 | 61.8119 | -72.4999 | 62.9931 | -34.4116 | 8.3357 | [12 empty] | 1.2716 | 3.5709 | 8.8053 | -2.5123 | 0.022779
[7 empty] | 2072 | 11.1649 | [20 empty] | -0.091942 | 5.6698 | 18.4341 | -39.1167 | 61.9314 | -72.6751 | 63.1758 | -34.5281 | 8.3679 | [12 empty] | 1.2716 | 3.5726 | 8.8138 | -2.5159 | 0.022823
[7 empty] | 2073 | 11.1716 | [20 empty] | -0.091942 | 5.6725 | 18.4519 | -39.1734 | 62.0511 | -72.8507 | 63.359 | -34.6449 | 8.4003 | [12 empty] | 1.2716 | 3.5743 | 8.8223 | -2.5196 | 0.022867
[7 empty] | 2074 | 11.1782 | [20 empty] | -0.091942 | 5.6753 | 18.4697 | -39.2301 | 62.1709 | -73.0265 | 63.5426 | -34.7621 | 8.4327 | [12 empty] | 1.2716 | 3.5761 | 8.8309 | -2.5232 | 0.022911
[7 empty] | 2075 | 11.1849 | [20 empty] | -0.091942 | 5.678 | 18.4875 | -39.2869 | 62.2909 | -73.2028 | 63.7267 | -34.8796 | 8.4653 | [12 empty] | 1.2716 | 3.5778 | 8.8394 | -2.5269 | 0.022956
[7 empty] | 2076 | 11.1915 | [20 empty] | -0.091942 | 5.6807 | 18.5054 | -39.3437 | 62.411 | -73.3793 | 63.9111 | -34.9974 | 8.498 | [12 empty] | 1.2716 | 3.5795 | 8.8479 | -2.5305 | 0.023
[7 empty] | 2077 | 11.1981 | [20 empty] | -0.091942 | 5.6835 | 18.5232 | -39.4006 | 62.5314 | -73.5562 | 64.0961 | -35.1156 | 8.5308 | [12 empty] | 1.2716 | 3.5812 | 8.8564 | -2.5342 | 0.023044
[7 empty] | 2078 | 11.2048 | [20 empty] | -0.091942 | 5.6862 | 18.541 | -39.4575 | 62.6519 | -73.7335 | 64.2815 | -35.2341 | 8.5637 | [12 empty] | 1.2716 | 3.583 | 8.865 | -2.5379 | 0.023089
[7 empty] | 2079 | 11.2114 | [20 empty] | -0.091942 | 5.6889 | 18.5589 | -39.5145 | 62.7726 | -73.9111 | 64.4673 | -35.353 | 8.5968 | [12 empty] | 1.2716 | 3.5847 | 8.8735 | -2.5415 | 0.023133
[7 empty] | 2080 | 11.218 | [20 empty] | -0.091942 | 5.6917 | 18.5767 | -39.5715 | 62.8934 | -74.089 | 64.6536 | -35.4722 | 8.6299 | [12 empty] | 1.2716 | 3.5864 | 8.882 | -2.5452 | 0.023178
[7 empty] | 2081 | 11.2247 | [20 empty] | -0.091942 | 5.6944 | 18.5946 | -39.6286 | 63.0145 | -74.2673 | 64.8403 | -35.5917 | 8.6631 | [12 empty] | 1.2716 | 3.5881 | 8.8906 | -2.5489 | 0.023222
[7 empty] | 2082 | 11.2313 | [20 empty] | -0.091942 | 5.6972 | 18.6125 | -39.6858 | 63.1357 | -74.4459 | 65.0275 | -35.7116 | 8.6965 | [12 empty] | 1.2716 | 3.5899 | 8.8991 | -2.5525 | 0.023267
[7 empty] | 2083 | 11.238 | [20 empty] | -0.091942 | 5.6999 | 18.6304 | -39.743 | 63.2571 | -74.6248 | 65.2151 | -35.8319 | 8.73 | [12 empty] | 1.2716 | 3.5916 | 8.9077 | -2.5562 | 0.023312
[7 empty] | 2084 | 11.2446 | [20 empty] | -0.091942 | 5.7026 | 18.6483 | -39.8003 | 63.3786 | -74.8041 | 65.4032 | -35.9524 | 8.7636 | [12 empty] | 1.2716 | 3.5933 | 8.9162 | -2.5599 | 0.023357
[7 empty] | 2085 | 11.2512 | [20 empty] | -0.091942 | 5.7054 | 18.6662 | -39.8576 | 63.5004 | -74.9838 | 65.5917 | -36.0734 | 8.7973 | [12 empty] | 1.2716 | 3.595 | 8.9248 | -2.5636 | 0.023401
[7 empty] | 2086 | 11.2579 | [20 empty] | -0.091942 | 5.7081 | 18.6841 | -39.915 | 63.6223 | -75.1638 | 65.7807 | -36.1947 | 8.8311 | [12 empty] | 1.2716 | 3.5968 | 8.9333 | -2.5673 | 0.023446
[7 empty] | 2087 | 11.2645 | [20 empty] | -0.091942 | 5.7108 | 18.702 | -39.9724 | 63.7444 | -75.3441 | 65.9701 | -36.3163 | 8.865 | [12 empty] | 1.2716 | 3.5985 | 8.9419 | -2.571 | 0.023491
[7 empty] | 2088 | 11.2712 | [20 empty] | -0.091942 | 5.7136 | 18.7199 | -40.0299 | 63.8666 | -75.5248 | 66.16 | -36.4383 | 8.899 | [12 empty] | 1.2716 | 3.6002 | 8.9505 | -2.5747 | 0.023536
[7 empty] | 2089 | 11.2778 | [20 empty] | -0.091942 | 5.7163 | 18.7379 | -40.0874 | 63.9891 | -75.7058 | 66.3503 | -36.5606 | 8.9332 | [12 empty] | 1.2716 | 3.6019 | 8.9591 | -2.5784 | 0.023582
[7 empty] | 2090 | 11.2844 | [20 empty] | -0.091942 | 5.719 | 18.7558 | -40.145 | 64.1117 | -75.8872 | 66.5411 | -36.6833 | 8.9675 | [12 empty] | 1.2716 | 3.6037 | 8.9676 | -2.5821 | 0.023627
[7 empty] | 2091 | 11.2911 | [20 empty] | -0.091942 | 5.7218 | 18.7737 | -40.2027 | 64.2345 | -76.0689 | 66.7324 | -36.8063 | 9.0018 | [12 empty] | 1.2716 | 3.6054 | 8.9762 | -2.5858 | 0.023672
[7 empty] | 2092 | 11.2977 | [20 empty] | -0.091942 | 5.7245 | 18.7917 | -40.2604 | 64.3574 | -76.251 | 66.9241 | -36.9297 | 9.0363 | [12 empty] | 1.2716 | 3.6071 | 8.9848 | -2.5895 | 0.023717
[7 empty] | 2093 | 11.3044 | [20 empty] | -0.091942 | 5.7273 | 18.8097 | -40.3181 | 64.4806 | -76.4334 | 67.1163 | -37.0535 | 9.071 | [12 empty] | 1.2716 | 3.6088 | 8.9934 | -2.5932 | 0.023763
[7 empty] | 2094 | 11.311 | [20 empty] | -0.091942 | 5.73 | 18.8277 | -40.376 | 64.6039 | -76.6162 | 67.3089 | -37.1776 | 9.1057 | [12 empty] | 1.2716 | 3.6105 | 9.002 | -2.5969 | 0.023808
[7 empty] | 2095 | 11.3177 | [20 empty] | -0.091942 | 5.7327 | 18.8456 | -40.4338 | 64.7274 | -76.7993 | 67.502 | -37.302 | 9.1405 | [12 empty] | 1.2716 | 3.6123 | 9.0106 | -2.6007 | 0.023854
[7 empty] | 2096 | 11.3243 | [20 empty] | -0.091942 | 5.7355 | 18.8636 | -40.4918 | 64.8511 | -76.9827 | 67.6956 | -37.4269 | 9.1755 | [12 empty] | 1.2716 | 3.614 | 9.0192 | -2.6044 | 0.023899
[7 empty] | 2097 | 11.331 | [20 empty] | -0.091942 | 5.7382 | 18.8816 | -40.5498 | 64.9749 | -77.1666 | 67.8896 | -37.552 | 9.2106 | [12 empty] | 1.2716 | 3.6157 | 9.0278 | -2.6081 | 0.023945
[7 empty] | 2098 | 11.3376 | [20 empty] | -0.091942 | 5.7409 | 18.8997 | -40.6078 | 65.0989 | -77.3507 | 68.0841 | -37.6776 | 9.2458 | [12 empty] | 1.2716 | 3.6174 | 9.0364 | -2.6118 | 0.023991
[7 empty] | 2099 | 11.3443 | [20 empty] | -0.091942 | 5.7437 | 18.9177 | -40.6659 | 65.2231 | -77.5352 | 68.279 | -37.8034 | 9.2811 | [12 empty] | 1.2716 | 3.6192 | 9.045 | -2.6156 | 0.024036
[7 empty] | 2100 | 11.3509 | [20 empty] | -0.091942 | 5.7464 | 18.9357 | -40.724 | 65.3475 | -77.7201 | 68.4744 | -37.9297 | 9.3165 | [12 empty] | 1.2716 | 3.6209 | 9.0537 | -2.6193 | 0.024082
[7 empty] | 2101 | 11.3576 | [20 empty] | -0.091942 | 5.7491 | 18.9537 | -40.7822 | 65.4721 | -77.9053 | 68.6703 | -38.0563 | 9.3521 | [12 empty] | 1.2716 | 3.6226 | 9.0623 | -2.6231 | 0.024128
[7 empty] | 2102 | 11.3642 | [20 empty] | -0.091942 | 5.7519 | 18.9718 | -40.8405 | 65.5968 | -78.0909 | 68.8666 | -38.1833 | 9.3877 | [12 empty] | 1.2716 | 3.6243 | 9.0709 | -2.6268 | 0.024174
[7 empty] | 2103 | 11.3709 | [20 empty] | -0.091942 | 5.7546 | 18.9898 | -40.8988 | 65.7217 | -78.2768 | 69.0634 | -38.3106 | 9.4235 | [12 empty] | 1.2716 | 3.6261 | 9.0795 | -2.6306 | 0.02422
[7 empty] | 2104 | 11.3775 | [20 empty] | -0.091942 | 5.7574 | 19.0079 | -40.9572 | 65.8468 | -78.4631 | 69.2607 | -38.4383 | 9.4594 | [12 empty] | 1.2716 | 3.6278 | 9.0882 | -2.6343 | 0.024266
[7 empty] | 2105 | 11.3842 | [20 empty] | -0.091942 | 5.7601 | 19.026 | -41.0156 | 65.9721 | -78.6498 | 69.4585 | -38.5664 | 9.4955 | [12 empty] | 1.2716 | 3.6295 | 9.0968 | -2.6381 | 0.024312
[7 empty] | 2106 | 11.3908 | [20 empty] | -0.091942 | 5.7628 | 19.0441 | -41.0741 | 66.0976 | -78.8368 | 69.6567 | -38.6948 | 9.5316 | [12 empty] | 1.2716 | 3.6312 | 9.1055 | -2.6418 | 0.024359
[7 empty] | 2107 | 11.3975 | [20 empty] | -0.091942 | 5.7656 | 19.0622 | -41.1326 | 66.2232 | -79.0241 | 69.8554 | -38.8236 | 9.5679 | [12 empty] | 1.2716 | 3.633 | 9.1141 | -2.6456 | 0.024405
[7 empty] | 2108 | 11.4042 | [20 empty] | -0.091942 | 5.7683 | 19.0803 | -41.1912 | 66.349 | -79.2118 | 70.0545 | -38.9528 | 9.6043 | [12 empty] | 1.2716 | 3.6347 | 9.1228 | -2.6494 | 0.024451
[7 empty] | 2109 | 11.4108 | [20 empty] | -0.091942 | 5.771 | 19.0984 | -41.2499 | 66.475 | -79.3999 | 70.2542 | -39.0823 | 9.6408 | [12 empty] | 1.2716 | 3.6364 | 9.1314 | -2.6531 | 0.024498
[7 empty] | 2110 | 11.4175 | [20 empty] | -0.091942 | 5.7738 | 19.1165 | -41.3086 | 66.6012 | -79.5883 | 70.4543 | -39.2122 | 9.6774 | [12 empty] | 1.2716 | 3.6381 | 9.1401 | -2.6569 | 0.024544
[7 empty] | 2111 | 11.4241 | [20 empty] | -0.091942 | 5.7765 | 19.1346 | -41.3673 | 66.7275 | -79.7771 | 70.6548 | -39.3425 | 9.7142 | [12 empty] | 1.2716 | 3.6399 | 9.1488 | -2.6607 | 0.024591
[7 empty] | 2112 | 11.4308 | [20 empty] | -0.091942 | 5.7792 | 19.1527 | -41.4262 | 66.854 | -79.9662 | 70.8559 | -39.4731 | 9.751 | [12 empty] | 1.2716 | 3.6416 | 9.1574 | -2.6645 | 0.024637
[7 empty] | 2113 | 11.4374 | [20 empty] | -0.091942 | 5.782 | 19.1709 | -41.485 | 66.9807 | -80.1557 | 71.0574 | -39.6042 | 9.788 | [12 empty] | 1.2716 | 3.6433 | 9.1661 | -2.6683 | 0.024684
[7 empty] | 2114 | 11.4441 | [20 empty] | -0.091942 | 5.7847 | 19.189 | -41.544 | 67.1076 | -80.3456 | 71.2594 | -39.7355 | 9.8251 | [12 empty] | 1.2716 | 3.645 | 9.1748 | -2.6721 | 0.024731
[7 empty] | 2115 | 11.4508 | [20 empty] | -0.091942 | 5.7875 | 19.2072 | -41.6029 | 67.2347 | -80.5358 | 71.4619 | -39.8673 | 9.8624 | [12 empty] | 1.2716 | 3.6468 | 9.1835 | -2.6759 | 0.024778
[7 empty] | 2116 | 11.4574 | [20 empty] | -0.091942 | 5.7902 | 19.2253 | -41.662 | 67.3619 | -80.7263 | 71.6649 | -39.9994 | 9.8998 | [12 empty] | 1.2716 | 3.6485 | 9.1921 | -2.6797 | 0.024825
[7 empty] | 2117 | 11.4641 | [20 empty] | -0.091942 | 5.7929 | 19.2435 | -41.7211 | 67.4894 | -80.9173 | 71.8684 | -40.132 | 9.9372 | [12 empty] | 1.2716 | 3.6502 | 9.2008 | -2.6835 | 0.024871
[7 empty] | 2118 | 11.4707 | [20 empty] | -0.091942 | 5.7957 | 19.2617 | -41.7802 | 67.617 | -81.1086 | 72.0723 | -40.2648 | 9.9749 | [12 empty] | 1.2716 | 3.6519 | 9.2095 | -2.6873 | 0.024919
[7 empty] | 2119 | 11.4774 | [20 empty] | -0.091942 | 5.7984 | 19.2799 | -41.8394 | 67.7448 | -81.3002 | 72.2767 | -40.3981 | 10.0126 | [12 empty] | 1.2716 | 3.6537 | 9.2182 | -2.6911 | 0.024966
[7 empty] | 2120 | 11.4841 | [20 empty] | -0.091942 | 5.8011 | 19.2981 | -41.8987 | 67.8727 | -81.4922 | 72.4816 | -40.5318 | 10.0505 | [12 empty] | 1.2716 | 3.6554 | 9.2269 | -2.6949 | 0.025013
[7 empty] | 2121 | 11.4907 | [20 empty] | -0.091942 | 5.8039 | 19.3163 | -41.958 | 68.0009 | -81.6846 | 72.687 | -40.6658 | 10.0884 | [12 empty] | 1.2716 | 3.6571 | 9.2356 | -2.6987 | 0.02506
[7 empty] | 2122 | 11.4974 | [20 empty] | -0.091942 | 5.8066 | 19.3345 | -42.0174 | 68.1292 | -81.8774 | 72.8928 | -40.8002 | 10.1266 | [12 empty] | 1.2716 | 3.6588 | 9.2443 | -2.7025 | 0.025107
[7 empty] | 2123 | 11.5041 | [20 empty] | -0.091942 | 5.8093 | 19.3528 | -42.0768 | 68.2577 | -82.0705 | 73.0992 | -40.935 | 10.1648 | [12 empty] | 1.2716 | 3.6606 | 9.2531 | -2.7063 | 0.025155
[7 empty] | 2124 | 11.5107 | [20 empty] | -0.091942 | 5.8121 | 19.371 | -42.1363 | 68.3864 | -82.2639 | 73.306 | -41.0701 | 10.2032 | [12 empty] | 1.2716 | 3.6623 | 9.2618 | -2.7102 | 0.025202
[7 empty] | 2125 | 11.5174 | [20 empty] | -0.091942 | 5.8148 | 19.3892 | -42.1958 | 68.5153 | -82.4578 | 73.5133 | -41.2057 | 10.2417 | [12 empty] | 1.2716 | 3.664 | 9.2705 | -2.714 | 0.02525
[7 empty] | 2126 | 11.524 | [20 empty] | -0.091942 | 5.8176 | 19.4075 | -42.2554 | 68.6444 | -82.652 | 73.7211 | -41.3416 | 10.2803 | [12 empty] | 1.2716 | 3.6657 | 9.2792 | -2.7178 | 0.025297
[7 empty] | 2127 | 11.5307 | [20 empty] | -0.091942 | 5.8203 | 19.4258 | -42.3151 | 68.7736 | -82.8465 | 73.9294 | -41.4779 | 10.319 | [12 empty] | 1.2716 | 3.6674 | 9.288 | -2.7217 | 0.025345
[7 empty] | 2128 | 11.5374 | [20 empty] | -0.091942 | 5.823 | 19.444 | -42.3748 | 68.903 | -83.0415 | 74.1382 | -41.6146 | 10.3579 | [12 empty] | 1.2716 | 3.6692 | 9.2967 | -2.7255 | 0.025392
[7 empty] | 2129 | 11.544 | [20 empty] | -0.091942 | 5.8258 | 19.4623 | -42.4346 | 69.0327 | -83.2368 | 74.3475 | -41.7517 | 10.3969 | [12 empty] | 1.2716 | 3.6709 | 9.3054 | -2.7293 | 0.02544
[7 empty] | 2130 | 11.5507 | [20 empty] | -0.091942 | 5.8285 | 19.4806 | -42.4944 | 69.1624 | -83.4324 | 74.5573 | -41.8892 | 10.436 | [12 empty] | 1.2716 | 3.6726 | 9.3142 | -2.7332 | 0.025488
[7 empty] | 2131 | 11.5574 | [20 empty] | -0.091942 | 5.8312 | 19.4989 | -42.5543 | 69.2924 | -83.6285 | 74.7676 | -42.027 | 10.4753 | [12 empty] | 1.2716 | 3.6743 | 9.3229 | -2.737 | 0.025536
[7 empty] | 2132 | 11.5641 | [20 empty] | -0.091942 | 5.834 | 19.5172 | -42.6142 | 69.4226 | -83.8249 | 74.9783 | -42.1653 | 10.5147 | [12 empty] | 1.2716 | 3.6761 | 9.3317 | -2.7409 | 0.025584
[7 empty] | 2133 | 11.5707 | [20 empty] | -0.091942 | 5.8367 | 19.5355 | -42.6742 | 69.5529 | -84.0217 | 75.1896 | -42.3039 | 10.5542 | [12 empty] | 1.2716 | 3.6778 | 9.3404 | -2.7448 | 0.025632
[7 empty] | 2134 | 11.5774 | [20 empty] | -0.091942 | 5.8394 | 19.5538 | -42.7343 | 69.6834 | -84.2188 | 75.4013 | -42.4429 | 10.5939 | [12 empty] | 1.2716 | 3.6795 | 9.3492 | -2.7486 | 0.02568
[7 empty] | 2135 | 11.5841 | [20 empty] | -0.091942 | 5.8422 | 19.5722 | -42.7944 | 69.8142 | -84.4163 | 75.6136 | -42.5823 | 10.6336 | [12 empty] | 1.2716 | 3.6812 | 9.358 | -2.7525 | 0.025728
[7 empty] | 2136 | 11.5907 | [20 empty] | -0.091942 | 5.8449 | 19.5905 | -42.8545 | 69.945 | -84.6142 | 75.8263 | -42.7221 | 10.6736 | [12 empty] | 1.2716 | 3.683 | 9.3667 | -2.7564 | 0.025776
[7 empty] | 2137 | 11.5974 | [20 empty] | -0.091942 | 5.8477 | 19.6088 | -42.9147 | 70.0761 | -84.8124 | 76.0396 | -42.8624 | 10.7136 | [12 empty] | 1.2716 | 3.6847 | 9.3755 | -2.7602 | 0.025825
[7 empty] | 2138 | 11.6041 | [20 empty] | -0.091942 | 5.8504 | 19.6272 | -42.975 | 70.2074 | -85.0111 | 76.2533 | -43.003 | 10.7538 | [12 empty] | 1.2716 | 3.6864 | 9.3843 | -2.7641 | 0.025873
[7 empty] | 2139 | 11.6108 | [20 empty] | -0.091942 | 5.8531 | 19.6456 | -43.0353 | 70.3388 | -85.2101 | 76.4676 | -43.1439 | 10.7941 | [12 empty] | 1.2716 | 3.6881 | 9.3931 | -2.768 | 0.025922
[7 empty] | 2140 | 11.6174 | [20 empty] | -0.091942 | 5.8559 | 19.6639 | -43.0957 | 70.4704 | -85.4094 | 76.6823 | -43.2853 | 10.8345 | [12 empty] | 1.2716 | 3.6899 | 9.4018 | -2.7719 | 0.02597
[7 empty] | 2141 | 11.6241 | [20 empty] | -0.091942 | 5.8586 | 19.6823 | -43.1562 | 70.6023 | -85.6092 | 76.8976 | -43.4271 | 10.8751 | [12 empty] | 1.2716 | 3.6916 | 9.4106 | -2.7758 | 0.026019
[7 empty] | 2142 | 11.6308 | [20 empty] | -0.091942 | 5.8613 | 19.7007 | -43.2167 | 70.7343 | -85.8093 | 77.1133 | -43.5693 | 10.9158 | [12 empty] | 1.2716 | 3.6933 | 9.4194 | -2.7796 | 0.026067
[7 empty] | 2143 | 11.6374 | [20 empty] | -0.091942 | 5.8641 | 19.7191 | -43.2772 | 70.8664 | -86.0098 | 77.3296 | -43.7119 | 10.9566 | [12 empty] | 1.2716 | 3.695 | 9.4282 | -2.7835 | 0.026116
[7 empty] | 2144 | 11.6441 | [20 empty] | -0.091942 | 5.8668 | 19.7375 | -43.3378 | 70.9988 | -86.2106 | 77.5463 | -43.8549 | 10.9976 | [12 empty] | 1.2716 | 3.6968 | 9.437 | -2.7874 | 0.026165
[7 empty] | 2145 | 11.6508 | [20 empty] | -0.091942 | 5.8696 | 19.7559 | -43.3985 | 71.1314 | -86.4119 | 77.7636 | -43.9983 | 11.0387 | [12 empty] | 1.2716 | 3.6985 | 9.4458 | -2.7913 | 0.026214
[7 empty] | 2146 | 11.6575 | [20 empty] | -0.091942 | 5.8723 | 19.7744 | -43.4592 | 71.2641 | -86.6135 | 77.9814 | -44.142 | 11.0799 | [12 empty] | 1.2716 | 3.7002 | 9.4546 | -2.7952 | 0.026263
[7 empty] | 2147 | 11.6642 | [20 empty] | -0.091942 | 5.875 | 19.7928 | -43.52 | 71.397 | -86.8155 | 78.1996 | -44.2862 | 11.1213 | [12 empty] | 1.2716 | 3.7019 | 9.4635 | -2.7992 | 0.026312
[7 empty] | 2148 | 11.6708 | [20 empty] | -0.091942 | 5.8778 | 19.8112 | -43.5809 | 71.5301 | -87.0178 | 78.4184 | -44.4308 | 11.1628 | [12 empty] | 1.2716 | 3.7037 | 9.4723 | -2.8031 | 0.026361
[7 empty] | 2149 | 11.6775 | [20 empty] | -0.091942 | 5.8805 | 19.8297 | -43.6417 | 71.6634 | -87.2206 | 78.6377 | -44.5758 | 11.2044 | [12 empty] | 1.2716 | 3.7054 | 9.4811 | -2.807 | 0.02641
[7 empty] | 2150 | 11.6842 | [20 empty] | -0.091942 | 5.8832 | 19.8481 | -43.7027 | 71.7969 | -87.4237 | 78.8576 | -44.7212 | 11.2462 | [12 empty] | 1.2716 | 3.7071 | 9.4899 | -2.8109 | 0.026459
[7 empty] | 2151 | 11.6909 | [20 empty] | -0.091942 | 5.886 | 19.8666 | -43.7637 | 71.9306 | -87.6272 | 79.0779 | -44.867 | 11.2881 | [12 empty] | 1.2716 | 3.7088 | 9.4987 | -2.8148 | 0.026508
[7 empty] | 2152 | 11.6975 | [20 empty] | -0.091942 | 5.8887 | 19.8851 | -43.8248 | 72.0644 | -87.8311 | 79.2987 | -45.0132 | 11.3302 | [12 empty] | 1.2716 | 3.7106 | 9.5076 | -2.8188 | 0.026558
[7 empty] | 2153 | 11.7042 | [20 empty] | -0.091942 | 5.8914 | 19.9036 | -43.8859 | 72.1985 | -88.0354 | 79.5201 | -45.1599 | 11.3724 | [12 empty] | 1.2716 | 3.7123 | 9.5164 | -2.8227 | 0.026607
[7 empty] | 2154 | 11.7109 | [20 empty] | -0.091942 | 5.8942 | 19.9221 | -43.9471 | 72.3327 | -88.24 | 79.7419 | -45.3069 | 11.4147 | [12 empty] | 1.2716 | 3.714 | 9.5253 | -2.8266 | 0.026656
[7 empty] | 2155 | 11.7176 | [20 empty] | -0.091942 | 5.8969 | 19.9406 | -44.0083 | 72.4671 | -88.445 | 79.9643 | -45.4543 | 11.4572 | [12 empty] | 1.2716 | 3.7157 | 9.5341 | -2.8306 | 0.026706
[7 empty] | 2156 | 11.7243 | [20 empty] | -0.091942 | 5.8997 | 19.9591 | -44.0696 | 72.6017 | -88.6504 | 80.1872 | -45.6022 | 11.4998 | [12 empty] | 1.2716 | 3.7175 | 9.543 | -2.8345 | 0.026756
[7 empty] | 2157 | 11.731 | [20 empty] | -0.091942 | 5.9024 | 19.9776 | -44.131 | 72.7365 | -88.8562 | 80.4106 | -45.7504 | 11.5425 | [12 empty] | 1.2716 | 3.7192 | 9.5518 | -2.8385 | 0.026805
[7 empty] | 2158 | 11.7376 | [20 empty] | -0.091942 | 5.9051 | 19.9961 | -44.1924 | 72.8715 | -89.0624 | 80.6346 | -45.8991 | 11.5854 | [12 empty] | 1.2716 | 3.7209 | 9.5607 | -2.8424 | 0.026855
[7 empty] | 2159 | 11.7443 | [20 empty] | -0.091942 | 5.9079 | 20.0147 | -44.2538 | 73.0067 | -89.2689 | 80.859 | -46.0482 | 11.6284 | [12 empty] | 1.2716 | 3.7226 | 9.5695 | -2.8464 | 0.026905
[7 empty] | 2160 | 11.751 | [20 empty] | -0.091942 | 5.9106 | 20.0332 | -44.3153 | 73.142 | -89.4758 | 81.084 | -46.1977 | 11.6716 | [12 empty] | 1.2716 | 3.7243 | 9.5784 | -2.8503 | 0.026955
[7 empty] | 2161 | 11.7577 | [20 empty] | -0.091942 | 5.9133 | 20.0518 | -44.3769 | 73.2776 | -89.6831 | 81.3095 | -46.3476 | 11.7149 | [12 empty] | 1.2716 | 3.7261 | 9.5873 | -2.8543 | 0.027005
[7 empty] | 2162 | 11.7644 | [20 empty] | -0.091942 | 5.9161 | 20.0703 | -44.4386 | 73.4133 | -89.8908 | 81.5355 | -46.498 | 11.7583 | [12 empty] | 1.2716 | 3.7278 | 9.5961 | -2.8582 | 0.027055
[7 empty] | 2163 | 11.7711 | [20 empty] | -0.091942 | 5.9188 | 20.0889 | -44.5003 | 73.5492 | -90.0989 | 81.762 | -46.6487 | 11.8019 | [12 empty] | 1.2716 | 3.7295 | 9.605 | -2.8622 | 0.027105
[7 empty] | 2164 | 11.7777 | [20 empty] | -0.091942 | 5.9215 | 20.1075 | -44.562 | 73.6853 | -90.3074 | 81.9891 | -46.7999 | 11.8456 | [12 empty] | 1.2716 | 3.7312 | 9.6139 | -2.8662 | 0.027155
[7 empty] | 2165 | 11.7844 | [20 empty] | -0.091942 | 5.9243 | 20.1261 | -44.6238 | 73.8216 | -90.5162 | 82.2167 | -46.9515 | 11.8895 | [12 empty] | 1.2716 | 3.733 | 9.6228 | -2.8702 | 0.027205
[7 empty] | 2166 | 11.7911 | [20 empty] | -0.091942 | 5.927 | 20.1447 | -44.6857 | 73.9581 | -90.7255 | 82.4448 | -47.1035 | 11.9335 | [12 empty] | 1.2716 | 3.7347 | 9.6317 | -2.8741 | 0.027255
[7 empty] | 2167 | 11.7978 | [20 empty] | -0.091942 | 5.9298 | 20.1633 | -44.7476 | 74.0948 | -90.9351 | 82.6734 | -47.256 | 11.9776 | [12 empty] | 1.2716 | 3.7364 | 9.6406 | -2.8781 | 0.027306
[7 empty] | 2168 | 11.8045 | [20 empty] | -0.091942 | 5.9325 | 20.1819 | -44.8096 | 74.2316 | -91.1451 | 82.9026 | -47.4088 | 12.0219 | [12 empty] | 1.2716 | 3.7381 | 9.6495 | -2.8821 | 0.027356
[7 empty] | 2169 | 11.8112 | [20 empty] | -0.091942 | 5.9352 | 20.2005 | -44.8716 | 74.3687 | -91.3555 | 83.1323 | -47.5621 | 12.0663 | [12 empty] | 1.2716 | 3.7399 | 9.6584 | -2.8861 | 0.027407
[7 empty] | 2170 | 11.8179 | [20 empty] | -0.091942 | 5.938 | 20.2191 | -44.9337 | 74.5059 | -91.5663 | 83.3625 | -47.7158 | 12.1109 | [12 empty] | 1.2716 | 3.7416 | 9.6673 | -2.8901 | 0.027457
[7 empty] | 2171 | 11.8246 | [20 empty] | -0.091942 | 5.9407 | 20.2378 | -44.9958 | 74.6434 | -91.7775 | 83.5933 | -47.87 | 12.1556 | [12 empty] | 1.2716 | 3.7433 | 9.6762 | -2.8941 | 0.027508
[7 empty] | 2172 | 11.8313 | [20 empty] | -0.091942 | 5.9434 | 20.2564 | -45.058 | 74.781 | -91.989 | 83.8246 | -48.0245 | 12.2005 | [12 empty] | 1.2716 | 3.745 | 9.6851 | -2.8981 | 0.027559
[7 empty] | 2173 | 11.8379 | [20 empty] | -0.091942 | 5.9462 | 20.2751 | -45.1203 | 74.9188 | -92.201 | 84.0564 | -48.1795 | 12.2455 | [12 empty] | 1.2716 | 3.7468 | 9.694 | -2.9021 | 0.027609
[7 empty] | 2174 | 11.8446 | [20 empty] | -0.091942 | 5.9489 | 20.2937 | -45.1826 | 75.0568 | -92.4133 | 84.2888 | -48.3349 | 12.2907 | [12 empty] | 1.2716 | 3.7485 | 9.703 | -2.9061 | 0.02766
[7 empty] | 2175 | 11.8513 | [20 empty] | -0.091942 | 5.9516 | 20.3124 | -45.245 | 75.195 | -92.6261 | 84.5217 | -48.4908 | 12.336 | [12 empty] | 1.2716 | 3.7502 | 9.7119 | -2.9101 | 0.027711
[7 empty] | 2176 | 11.858 | [20 empty] | -0.091942 | 5.9544 | 20.3311 | -45.3074 | 75.3334 | -92.8392 | 84.7551 | -48.6471 | 12.3814 | [12 empty] | 1.2716 | 3.7519 | 9.7208 | -2.9141 | 0.027762
[7 empty] | 2177 | 11.8647 | [20 empty] | -0.091942 | 5.9571 | 20.3498 | -45.3699 | 75.472 | -93.0527 | 84.9891 | -48.8038 | 12.427 | [12 empty] | 1.2716 | 3.7537 | 9.7298 | -2.9181 | 0.027813
[7 empty] | 2178 | 11.8714 | [20 empty] | -0.091942 | 5.9599 | 20.3685 | -45.4325 | 75.6107 | -93.2666 | 85.2236 | -48.9609 | 12.4727 | [12 empty] | 1.2716 | 3.7554 | 9.7387 | -2.9222 | 0.027864
[7 empty] | 2179 | 11.8781 | [20 empty] | -0.091942 | 5.9626 | 20.3872 | -45.4951 | 75.7497 | -93.4809 | 85.4586 | -49.1185 | 12.5186 | [12 empty] | 1.2716 | 3.7571 | 9.7477 | -2.9262 | 0.027916
[7 empty] | 2180 | 11.8848 | [20 empty] | -0.091942 | 5.9653 | 20.4059 | -45.5578 | 75.8888 | -93.6957 | 85.6942 | -49.2765 | 12.5647 | [12 empty] | 1.2716 | 3.7588 | 9.7566 | -2.9302 | 0.027967
[7 empty] | 2181 | 11.8915 | [20 empty] | -0.091942 | 5.9681 | 20.4246 | -45.6205 | 76.0282 | -93.9107 | 85.9303 | -49.4349 | 12.6108 | [12 empty] | 1.2716 | 3.7606 | 9.7656 | -2.9343 | 0.028018
[7 empty] | 2182 | 11.8982 | [20 empty] | -0.091942 | 5.9708 | 20.4434 | -45.6833 | 76.1677 | -94.1262 | 86.167 | -49.5938 | 12.6572 | [12 empty] | 1.2716 | 3.7623 | 9.7745 | -2.9383 | 0.02807
[7 empty] | 2183 | 11.9049 | [20 empty] | -0.091942 | 5.9735 | 20.4621 | -45.7461 | 76.3074 | -94.3421 | 86.4042 | -49.7531 | 12.7037 | [12 empty] | 1.2716 | 3.764 | 9.7835 | -2.9423 | 0.028121
[7 empty] | 2184 | 11.9116 | [20 empty] | -0.091942 | 5.9763 | 20.4809 | -45.809 | 76.4474 | -94.5584 | 86.642 | -49.9129 | 12.7503 | [12 empty] | 1.2716 | 3.7657 | 9.7924 | -2.9464 | 0.028173
[7 empty] | 2185 | 11.9183 | [20 empty] | -0.091942 | 5.979 | 20.4996 | -45.872 | 76.5875 | -94.7751 | 86.8803 | -50.0731 | 12.7971 | [12 empty] | 1.2716 | 3.7675 | 9.8014 | -2.9504 | 0.028224
[7 empty] | 2186 | 11.925 | [20 empty] | -0.091942 | 5.9817 | 20.5184 | -45.935 | 76.7278 | -94.9922 | 87.1191 | -50.2337 | 12.844 | [12 empty] | 1.2716 | 3.7692 | 9.8104 | -2.9545 | 0.028276
[7 empty] | 2187 | 11.9317 | [20 empty] | -0.091942 | 5.9845 | 20.5372 | -45.998 | 76.8683 | -95.2096 | 87.3585 | -50.3948 | 12.8911 | [12 empty] | 1.2716 | 3.7709 | 9.8194 | -2.9585 | 0.028328
[7 empty] | 2188 | 11.9384 | [20 empty] | -0.091942 | 5.9872 | 20.556 | -46.0612 | 77.009 | -95.4275 | 87.5985 | -50.5563 | 12.9383 | [12 empty] | 1.2716 | 3.7726 | 9.8283 | -2.9626 | 0.02838
[7 empty] | 2189 | 11.9451 | [20 empty] | -0.091942 | 5.99 | 20.5747 | -46.1243 | 77.1498 | -95.6458 | 87.8389 | -50.7183 | 12.9857 | [12 empty] | 1.2716 | 3.7744 | 9.8373 | -2.9667 | 0.028432
[7 empty] | 2190 | 11.9518 | [20 empty] | -0.091942 | 5.9927 | 20.5935 | -46.1876 | 77.2909 | -95.8644 | 88.08 | -50.8807 | 13.0332 | [12 empty] | 1.2716 | 3.7761 | 9.8463 | -2.9707 | 0.028484
[7 empty] | 2191 | 11.9585 | [20 empty] | -0.091942 | 5.9954 | 20.6124 | -46.2509 | 77.4322 | -96.0835 | 88.3216 | -51.0436 | 13.0809 | [12 empty] | 1.2716 | 3.7778 | 9.8553 | -2.9748 | 0.028536
[7 empty] | 2192 | 11.9652 | [20 empty] | -0.091942 | 5.9982 | 20.6312 | -46.3142 | 77.5736 | -96.303 | 88.5637 | -51.2069 | 13.1287 | [12 empty] | 1.2716 | 3.7795 | 9.8643 | -2.9789 | 0.028588
[7 empty] | 2193 | 11.9719 | [20 empty] | -0.091942 | 6.0009 | 20.65 | -46.3777 | 77.7153 | -96.5228 | 88.8064 | -51.3706 | 13.1767 | [12 empty] | 1.2716 | 3.7812 | 9.8733 | -2.983 | 0.02864
[7 empty] | 2194 | 11.9786 | [20 empty] | -0.091942 | 6.0036 | 20.6688 | -46.4411 | 77.8571 | -96.7431 | 89.0497 | -51.5348 | 13.2249 | [12 empty] | 1.2716 | 3.783 | 9.8823 | -2.987 | 0.028692
[7 empty] | 2195 | 11.9853 | [20 empty] | -0.091942 | 6.0064 | 20.6877 | -46.5047 | 77.9992 | -96.9638 | 89.2935 | -51.6995 | 13.2732 | [12 empty] | 1.2716 | 3.7847 | 9.8913 | -2.9911 | 0.028745
[7 empty] | 2196 | 11.992 | [20 empty] | -0.091942 | 6.0091 | 20.7065 | -46.5683 | 78.1414 | -97.1849 | 89.5378 | -51.8646 | 13.3216 | [12 empty] | 1.2716 | 3.7864 | 9.9003 | -2.9952 | 0.028797
[7 empty] | 2197 | 11.9987 | [20 empty] | -0.091942 | 6.0118 | 20.7254 | -46.6319 | 78.2838 | -97.4063 | 89.7827 | -52.0301 | 13.3702 | [12 empty] | 1.2716 | 3.7881 | 9.9094 | -2.9993 | 0.02885
[7 empty] | 2198 | 12.0054 | [20 empty] | -0.091942 | 6.0146 | 20.7443 | -46.6956 | 78.4265 | -97.6282 | 90.0282 | -52.1961 | 13.419 | [12 empty] | 1.2716 | 3.7899 | 9.9184 | -3.0034 | 0.028902
[7 empty] | 2199 | 12.0121 | [20 empty] | -0.091942 | 6.0173 | 20.7632 | -46.7594 | 78.5693 | -97.8505 | 90.2743 | -52.3626 | 13.4679 | [12 empty] | 1.2716 | 3.7916 | 9.9274 | -3.0075 | 0.028955
[7 empty] | 2200 | 12.0188 | [20 empty] | -0.091942 | 6.0201 | 20.782 | -46.8232 | 78.7123 | -98.0732 | 90.5208 | -52.5295 | 13.517 | [12 empty] | 1.2716 | 3.7933 | 9.9364 | -3.0116 | 0.029007
[7 empty] | 2201 | 12.0255 | [20 empty] | -0.091942 | 6.0228 | 20.8009 | -46.8871 | 78.8555 | -98.2963 | 90.768 | -52.6968 | 13.5662 | [12 empty] | 1.2716 | 3.795 | 9.9455 | -3.0157 | 0.02906
[7 empty] | 2202 | 12.0322 | [20 empty] | -0.091942 | 6.0255 | 20.8198 | -46.951 | 78.9989 | -98.5198 | 91.0157 | -52.8647 | 13.6156 | [12 empty] | 1.2716 | 3.7968 | 9.9545 | -3.0198 | 0.029113
[7 empty] | 2203 | 12.0389 | [20 empty] | -0.091942 | 6.0283 | 20.8388 | -47.015 | 79.1425 | -98.7437 | 91.264 | -53.033 | 13.6652 | [12 empty] | 1.2716 | 3.7985 | 9.9636 | -3.024 | 0.029166
[7 empty] | 2204 | 12.0456 | [20 empty] | -0.091942 | 6.031 | 20.8577 | -47.079 | 79.2863 | -98.968 | 91.5128 | -53.2017 | 13.7149 | [12 empty] | 1.2716 | 3.8002 | 9.9726 | -3.0281 | 0.029219
[7 empty] | 2205 | 12.0523 | [20 empty] | -0.091942 | 6.0337 | 20.8766 | -47.1432 | 79.4303 | -99.1928 | 91.7623 | -53.3709 | 13.7647 | [12 empty] | 1.2716 | 3.8019 | 9.9817 | -3.0322 | 0.029272
[7 empty] | 2206 | 12.059 | [20 empty] | -0.091942 | 6.0365 | 20.8956 | -47.2073 | 79.5745 | -99.4179 | 92.0122 | -53.5406 | 13.8148 | [12 empty] | 1.2716 | 3.8037 | 9.9907 | -3.0363 | 0.029325
[7 empty] | 2207 | 12.0657 | [20 empty] | -0.091942 | 6.0392 | 20.9145 | -47.2716 | 79.7189 | -99.6434 | 92.2628 | -53.7107 | 13.8649 | [12 empty] | 1.2716 | 3.8054 | 9.9998 | -3.0405 | 0.029378
[7 empty] | 2208 | 12.0724 | [20 empty] | -0.091942 | 6.0419 | 20.9335 | -47.3358 | 79.8635 | -99.8694 | 92.5139 | -53.8813 | 13.9153 | [12 empty] | 1.2716 | 3.8071 | 10.0088 | -3.0446 | 0.029432
[7 empty] | 2209 | 12.0791 | [20 empty] | -0.091942 | 6.0447 | 20.9524 | -47.4002 | 80.0083 | -100.0957 | 92.7656 | -54.0523 | 13.9658 | [12 empty] | 1.2716 | 3.8088 | 10.0179 | -3.0487 | 0.029485
[7 empty] | 2210 | 12.0858 | [20 empty] | -0.091942 | 6.0474 | 20.9714 | -47.4646 | 80.1532 | -100.3225 | 93.0178 | -54.2238 | 14.0164 | [12 empty] | 1.2716 | 3.8106 | 10.027 | -3.0529 | 0.029538
[7 empty] | 2211 | 12.0925 | [20 empty] | -0.091942 | 6.0502 | 20.9904 | -47.5291 | 80.2984 | -100.5497 | 93.2706 | -54.3958 | 14.0672 | [12 empty] | 1.2716 | 3.8123 | 10.0361 | -3.057 | 0.029592
[7 empty] | 2212 | 12.0992 | [20 empty] | -0.091942 | 6.0529 | 21.0094 | -47.5936 | 80.4438 | -100.7773 | 93.524 | -54.5683 | 14.1182 | [12 empty] | 1.2716 | 3.814 | 10.0451 | -3.0612 | 0.029646
[7 empty] | 2213 | 12.1059 | [20 empty] | -0.091942 | 6.0556 | 21.0284 | -47.6581 | 80.5893 | -101.0053 | 93.778 | -54.7412 | 14.1694 | [12 empty] | 1.2716 | 3.8157 | 10.0542 | -3.0653 | 0.029699
[7 empty] | 2214 | 12.1127 | [20 empty] | -0.091942 | 6.0584 | 21.0474 | -47.7228 | 80.7351 | -101.2337 | 94.0326 | -54.9146 | 14.2207 | [12 empty] | 1.2716 | 3.8175 | 10.0633 | -3.0695 | 0.029753
[7 empty] | 2215 | 12.1194 | [20 empty] | -0.091942 | 6.0611 | 21.0664 | -47.7875 | 80.8811 | -101.4625 | 94.2877 | -55.0884 | 14.2721 | [12 empty] | 1.2716 | 3.8192 | 10.0724 | -3.0736 | 0.029807
[7 empty] | 2216 | 12.1261 | [20 empty] | -0.091942 | 6.0638 | 21.0854 | -47.8522 | 81.0272 | -101.6918 | 94.5434 | -55.2628 | 14.3238 | [12 empty] | 1.2716 | 3.8209 | 10.0815 | -3.0778 | 0.029861
[7 empty] | 2217 | 12.1328 | [20 empty] | -0.091942 | 6.0666 | 21.1045 | -47.917 | 81.1736 | -101.9214 | 94.7996 | -55.4376 | 14.3756 | [12 empty] | 1.2716 | 3.8226 | 10.0906 | -3.082 | 0.029914
[7 empty] | 2218 | 12.1395 | [20 empty] | -0.091942 | 6.0693 | 21.1235 | -47.9819 | 81.3201 | -102.1515 | 95.0565 | -55.6128 | 14.4275 | [12 empty] | 1.2716 | 3.8244 | 10.0997 | -3.0861 | 0.029968
[7 empty] | 2219 | 12.1462 | [20 empty] | -0.091942 | 6.072 | 21.1426 | -48.0468 | 81.4669 | -102.382 | 95.3139 | -55.7886 | 14.4796 | [12 empty] | 1.2716 | 3.8261 | 10.1088 | -3.0903 | 0.030023
[7 empty] | 2220 | 12.1529 | [20 empty] | -0.091942 | 6.0748 | 21.1616 | -48.1118 | 81.6138 | -102.6129 | 95.5719 | -55.9648 | 14.5319 | [12 empty] | 1.2716 | 3.8278 | 10.1179 | -3.0945 | 0.030077
[7 empty] | 2221 | 12.1596 | [20 empty] | -0.091942 | 6.0775 | 21.1807 | -48.1769 | 81.761 | -102.8442 | 95.8305 | -56.1415 | 14.5844 | [12 empty] | 1.2716 | 3.8295 | 10.127 | -3.0987 | 0.030131
[7 empty] | 2222 | 12.1663 | [20 empty] | -0.091942 | 6.0803 | 21.1998 | -48.242 | 81.9083 | -103.0759 | 96.0897 | -56.3187 | 14.637 | [12 empty] | 1.2716 | 3.8313 | 10.1362 | -3.1029 | 0.030185
[7 empty] | 2223 | 12.1731 | [20 empty] | -0.091942 | 6.083 | 21.2189 | -48.3071 | 82.0559 | -103.3081 | 96.3495 | -56.4964 | 14.6898 | [12 empty] | 1.2716 | 3.833 | 10.1453 | -3.1071 | 0.03024
[7 empty] | 2224 | 12.1798 | [20 empty] | -0.091942 | 6.0857 | 21.2379 | -48.3724 | 82.2036 | -103.5406 | 96.6098 | -56.6745 | 14.7427 | [12 empty] | 1.2716 | 3.8347 | 10.1544 | -3.1113 | 0.030294
[7 empty] | 2225 | 12.1865 | [20 empty] | -0.091942 | 6.0885 | 21.257 | -48.4376 | 82.3516 | -103.7736 | 96.8707 | -56.8531 | 14.7958 | [12 empty] | 1.2716 | 3.8364 | 10.1636 | -3.1155 | 0.030349
[7 empty] | 2226 | 12.1932 | [20 empty] | -0.091942 | 6.0912 | 21.2762 | -48.503 | 82.4997 | -104.007 | 97.1323 | -57.0322 | 14.8491 | [12 empty] | 1.2716 | 3.8381 | 10.1727 | -3.1197 | 0.030403
[7 empty] | 2227 | 12.1999 | [20 empty] | -0.091942 | 6.0939 | 21.2953 | -48.5684 | 82.6481 | -104.2409 | 97.3944 | -57.2118 | 14.9026 | [12 empty] | 1.2716 | 3.8399 | 10.1818 | -3.1239 | 0.030458
[7 empty] | 2228 | 12.2066 | [20 empty] | -0.091942 | 6.0967 | 21.3144 | -48.6338 | 82.7966 | -104.4751 | 97.6571 | -57.3919 | 14.9562 | [12 empty] | 1.2716 | 3.8416 | 10.191 | -3.1281 | 0.030513
[7 empty] | 2229 | 12.2133 | [20 empty] | -0.091942 | 6.0994 | 21.3335 | -48.6993 | 82.9454 | -104.7098 | 97.9203 | -57.5725 | 15.01 | [12 empty] | 1.2716 | 3.8433 | 10.2001 | -3.1323 | 0.030567
[7 empty] | 2230 | 12.2201 | [20 empty] | -0.091942 | 6.1021 | 21.3527 | -48.7649 | 83.0943 | -104.9449 | 98.1842 | -57.7535 | 15.0639 | [12 empty] | 1.2716 | 3.845 | 10.2093 | -3.1365 | 0.030622
[7 empty] | 2231 | 12.2268 | [20 empty] | -0.091942 | 6.1049 | 21.3718 | -48.8306 | 83.2435 | -105.1804 | 98.4487 | -57.935 | 15.118 | [12 empty] | 1.2716 | 3.8468 | 10.2184 | -3.1407 | 0.030677
[7 empty] | 2232 | 12.2335 | [20 empty] | -0.091942 | 6.1076 | 21.391 | -48.8962 | 83.3928 | -105.4163 | 98.7138 | -58.1171 | 15.1723 | [12 empty] | 1.2716 | 3.8485 | 10.2276 | -3.145 | 0.030732
[7 empty] | 2233 | 12.2402 | [20 empty] | -0.091942 | 6.1104 | 21.4102 | -48.962 | 83.5424 | -105.6527 | 98.9794 | -58.2996 | 15.2268 | [12 empty] | 1.2716 | 3.8502 | 10.2368 | -3.1492 | 0.030787
[7 empty] | 2234 | 12.2469 | [20 empty] | -0.091942 | 6.1131 | 21.4294 | -49.0278 | 83.6921 | -105.8895 | 99.2457 | -58.4826 | 15.2814 | [12 empty] | 1.2716 | 3.8519 | 10.2459 | -3.1534 | 0.030843
[7 empty] | 2235 | 12.2536 | [20 empty] | -0.091942 | 6.1158 | 21.4486 | -49.0937 | 83.8421 | -106.1267 | 99.5125 | -58.6661 | 15.3363 | [12 empty] | 1.2716 | 3.8537 | 10.2551 | -3.1576 | 0.030898
[7 empty] | 2236 | 12.2604 | [20 empty] | -0.091942 | 6.1186 | 21.4677 | -49.1596 | 83.9922 | -106.3643 | 99.78 | -58.8501 | 15.3912 | [12 empty] | 1.2716 | 3.8554 | 10.2643 | -3.1619 | 0.030953
[7 empty] | 2237 | 12.2671 | [20 empty] | -0.091942 | 6.1213 | 21.487 | -49.2256 | 84.1426 | -106.6024 | 100.048 | -59.0345 | 15.4464 | [12 empty] | 1.2716 | 3.8571 | 10.2735 | -3.1661 | 0.031009
[7 empty] | 2238 | 12.2738 | [20 empty] | -0.091942 | 6.124 | 21.5062 | -49.2916 | 84.2931 | -106.8409 | 100.3167 | -59.2195 | 15.5017 | [12 empty] | 1.2716 | 3.8588 | 10.2827 | -3.1704 | 0.031064
[7 empty] | 2239 | 12.2805 | [20 empty] | -0.091942 | 6.1268 | 21.5254 | -49.3577 | 84.4439 | -107.0798 | 100.5859 | -59.405 | 15.5572 | [12 empty] | 1.2716 | 3.8606 | 10.2919 | -3.1746 | 0.03112
[7 empty] | 2240 | 12.2872 | [20 empty] | -0.091942 | 6.1295 | 21.5446 | -49.4239 | 84.5949 | -107.3191 | 100.8557 | -59.591 | 15.6129 | [12 empty] | 1.2716 | 3.8623 | 10.3011 | -3.1789 | 0.031175
[7 empty] | 2241 | 12.294 | [20 empty] | -0.091942 | 6.1322 | 21.5639 | -49.4901 | 84.746 | -107.5589 | 101.1262 | -59.7774 | 15.6687 | [12 empty] | 1.2716 | 3.864 | 10.3103 | -3.1831 | 0.031231
[7 empty] | 2242 | 12.3007 | [20 empty] | -0.091942 | 6.135 | 21.5831 | -49.5564 | 84.8974 | -107.7991 | 101.3973 | -59.9644 | 15.7248 | [12 empty] | 1.2716 | 3.8657 | 10.3195 | -3.1874 | 0.031287
[7 empty] | 2243 | 12.3074 | [20 empty] | -0.091942 | 6.1377 | 21.6024 | -49.6227 | 85.049 | -108.0397 | 101.6689 | -60.1519 | 15.781 | [12 empty] | 1.2716 | 3.8675 | 10.3287 | -3.1917 | 0.031343
[7 empty] | 2244 | 12.3141 | [20 empty] | -0.091942 | 6.1405 | 21.6216 | -49.6891 | 85.2007 | -108.2807 | 101.9412 | -60.3399 | 15.8373 | [12 empty] | 1.2716 | 3.8692 | 10.3379 | -3.1959 | 0.031399
[7 empty] | 2245 | 12.3208 | [20 empty] | -0.091942 | 6.1432 | 21.6409 | -49.7556 | 85.3527 | -108.5222 | 102.2141 | -60.5283 | 15.8939 | [12 empty] | 1.2716 | 3.8709 | 10.3471 | -3.2002 | 0.031455
[7 empty] | 2246 | 12.3276 | [20 empty] | -0.091942 | 6.1459 | 21.6602 | -49.8221 | 85.5049 | -108.7641 | 102.4875 | -60.7173 | 15.9506 | [12 empty] | 1.2716 | 3.8726 | 10.3563 | -3.2045 | 0.031511
[7 empty] | 2247 | 12.3343 | [20 empty] | -0.091942 | 6.1487 | 21.6795 | -49.8887 | 85.6573 | -109.0065 | 102.7616 | -60.9068 | 16.0075 | [12 empty] | 1.2716 | 3.8744 | 10.3655 | -3.2088 | 0.031567
[7 empty] | 2248 | 12.341 | [20 empty] | -0.091942 | 6.1514 | 21.6988 | -49.9553 | 85.8099 | -109.2493 | 103.0363 | -61.0968 | 16.0646 | [12 empty] | 1.2716 | 3.8761 | 10.3748 | -3.2131 | 0.031623
[7 empty] | 2249 | 12.3477 | [20 empty] | -0.091942 | 6.1541 | 21.7181 | -50.022 | 85.9626 | -109.4925 | 103.3116 | -61.2873 | 16.1218 | [12 empty] | 1.2716 | 3.8778 | 10.384 | -3.2174 | 0.031679
[7 empty] | 2250 | 12.3545 | [20 empty] | -0.091942 | 6.1569 | 21.7374 | -50.0888 | 86.1156 | -109.7361 | 103.5876 | -61.4783 | 16.1793 | [12 empty] | 1.2716 | 3.8795 | 10.3932 | -3.2217 | 0.031736
[7 empty] | 2251 | 12.3612 | [20 empty] | -0.091942 | 6.1596 | 21.7567 | -50.1556 | 86.2688 | -109.9802 | 103.8641 | -61.6698 | 16.2369 | [12 empty] | 1.2716 | 3.8813 | 10.4025 | -3.226 | 0.031792
[7 empty] | 2252 | 12.3679 | [20 empty] | -0.091942 | 6.1623 | 21.7761 | -50.2225 | 86.4222 | -110.2247 | 104.1413 | -61.8619 | 16.2947 | [12 empty] | 1.2716 | 3.883 | 10.4117 | -3.2303 | 0.031849
[7 empty] | 2253 | 12.3746 | [20 empty] | -0.091942 | 6.1651 | 21.7954 | -50.2894 | 86.5758 | -110.4696 | 104.419 | -62.0544 | 16.3527 | [12 empty] | 1.2716 | 3.8847 | 10.421 | -3.2346 | 0.031905
[7 empty] | 2254 | 12.3814 | [20 empty] | -0.091942 | 6.1678 | 21.8148 | -50.3564 | 86.7296 | -110.715 | 104.6974 | -62.2475 | 16.4108 | [12 empty] | 1.2716 | 3.8864 | 10.4302 | -3.2389 | 0.031962
[7 empty] | 2255 | 12.3881 | [20 empty] | -0.091942 | 6.1706 | 21.8341 | -50.4234 | 86.8837 | -110.9608 | 104.9764 | -62.441 | 16.4692 | [12 empty] | 1.2716 | 3.8881 | 10.4395 | -3.2432 | 0.032019
[7 empty] | 2256 | 12.3948 | [20 empty] | -0.091942 | 6.1733 | 21.8535 | -50.4906 | 87.0379 | -111.2071 | 105.2561 | -62.6351 | 16.5277 | [12 empty] | 1.2716 | 3.8899 | 10.4487 | -3.2475 | 0.032076
[7 empty] | 2257 | 12.4015 | [20 empty] | -0.091942 | 6.176 | 21.8729 | -50.5577 | 87.1923 | -111.4538 | 105.5363 | -62.8297 | 16.5864 | [12 empty] | 1.2716 | 3.8916 | 10.458 | -3.2518 | 0.032133
[7 empty] | 2258 | 12.4083 | [20 empty] | -0.091942 | 6.1788 | 21.8923 | -50.625 | 87.3469 | -111.7009 | 105.8172 | -63.0249 | 16.6453 | [12 empty] | 1.2716 | 3.8933 | 10.4673 | -3.2561 | 0.03219
[7 empty] | 2259 | 12.415 | [20 empty] | -0.091942 | 6.1815 | 21.9117 | -50.6923 | 87.5018 | -111.9485 | 106.0987 | -63.2205 | 16.7043 | [12 empty] | 1.2716 | 3.895 | 10.4765 | -3.2605 | 0.032247
[7 empty] | 2260 | 12.4217 | [20 empty] | -0.091942 | 6.1842 | 21.9311 | -50.7596 | 87.6568 | -112.1965 | 106.3808 | -63.4167 | 16.7636 | [12 empty] | 1.2716 | 3.8968 | 10.4858 | -3.2648 | 0.032304
[7 empty] | 2261 | 12.4284 | [20 empty] | -0.091942 | 6.187 | 21.9505 | -50.827 | 87.8121 | -112.4449 | 106.6635 | -63.6134 | 16.823 | [12 empty] | 1.2716 | 3.8985 | 10.4951 | -3.2691 | 0.032361
[7 empty] | 2262 | 12.4352 | [20 empty] | -0.091942 | 6.1897 | 21.9699 | -50.8945 | 87.9675 | -112.6938 | 106.9469 | -63.8106 | 16.8826 | [12 empty] | 1.2716 | 3.9002 | 10.5044 | -3.2735 | 0.032418
[7 empty] | 2263 | 12.4419 | [20 empty] | -0.091942 | 6.1924 | 21.9893 | -50.962 | 88.1232 | -112.9431 | 107.2309 | -64.0083 | 16.9424 | [12 empty] | 1.2716 | 3.9019 | 10.5137 | -3.2778 | 0.032476
[7 empty] | 2264 | 12.4486 | [20 empty] | -0.091942 | 6.1952 | 22.0088 | -51.0296 | 88.279 | -113.1929 | 107.5155 | -64.2065 | 17.0024 | [12 empty] | 1.2716 | 3.9037 | 10.523 | -3.2822 | 0.032533
[7 empty] | 2265 | 12.4554 | [20 empty] | -0.091942 | 6.1979 | 22.0282 | -51.0973 | 88.4351 | -113.4431 | 107.8007 | -64.4053 | 17.0626 | [12 empty] | 1.2716 | 3.9054 | 10.5323 | -3.2865 | 0.032591
[7 empty] | 2266 | 12.4621 | [20 empty] | -0.091942 | 6.2007 | 22.0477 | -51.165 | 88.5914 | -113.6937 | 108.0866 | -64.6046 | 17.1229 | [12 empty] | 1.2716 | 3.9071 | 10.5416 | -3.2909 | 0.032648
[7 empty] | 2267 | 12.4688 | [20 empty] | -0.091942 | 6.2034 | 22.0671 | -51.2327 | 88.7479 | -113.9448 | 108.3731 | -64.8045 | 17.1835 | [12 empty] | 1.2716 | 3.9088 | 10.5509 | -3.2952 | 0.032706
[7 empty] | 2268 | 12.4755 | [20 empty] | -0.091942 | 6.2061 | 22.0866 | -51.3006 | 88.9046 | -114.1963 | 108.6603 | -65.0048 | 17.2442 | [12 empty] | 1.2716 | 3.9106 | 10.5602 | -3.2996 | 0.032764
[7 empty] | 2269 | 12.4823 | [20 empty] | -0.091942 | 6.2089 | 22.1061 | -51.3684 | 89.0615 | -114.4483 | 108.9481 | -65.2057 | 17.3051 | [12 empty] | 1.2716 | 3.9123 | 10.5695 | -3.304 | 0.032821
[7 empty] | 2270 | 12.489 | [20 empty] | -0.091942 | 6.2116 | 22.1256 | -51.4364 | 89.2186 | -114.7007 | 109.2365 | -65.4072 | 17.3662 | [12 empty] | 1.2716 | 3.914 | 10.5788 | -3.3083 | 0.032879
[7 empty] | 2271 | 12.4957 | [20 empty] | -0.091942 | 6.2143 | 22.1451 | -51.5044 | 89.3759 | -114.9536 | 109.5255 | -65.6091 | 17.4275 | [12 empty] | 1.2716 | 3.9157 | 10.5881 | -3.3127 | 0.032937
[7 empty] | 2272 | 12.5025 | [20 empty] | -0.091942 | 6.2171 | 22.1646 | -51.5725 | 89.5334 | -115.2069 | 109.8152 | -65.8116 | 17.489 | [12 empty] | 1.2716 | 3.9175 | 10.5975 | -3.3171 | 0.032995
[7 empty] | 2273 | 12.5092 | [20 empty] | -0.091942 | 6.2198 | 22.1841 | -51.6406 | 89.6912 | -115.4607 | 110.1055 | -66.0147 | 17.5507 | [12 empty] | 1.2716 | 3.9192 | 10.6068 | -3.3215 | 0.033053
[7 empty] | 2274 | 12.5159 | [20 empty] | -0.091942 | 6.2225 | 22.2036 | -51.7088 | 89.8491 | -115.7149 | 110.3965 | -66.2182 | 17.6126 | [12 empty] | 1.2716 | 3.9209 | 10.6161 | -3.3259 | 0.033112
[7 empty] | 2275 | 12.5227 | [20 empty] | -0.091942 | 6.2253 | 22.2232 | -51.777 | 90.0073 | -115.9695 | 110.6881 | -66.4223 | 17.6746 | [12 empty] | 1.2716 | 3.9226 | 10.6255 | -3.3302 | 0.03317
[7 empty] | 2276 | 12.5294 | [20 empty] | -0.091942 | 6.228 | 22.2427 | -51.8453 | 90.1656 | -116.2247 | 110.9803 | -66.627 | 17.7369 | [12 empty] | 1.2716 | 3.9244 | 10.6348 | -3.3346 | 0.033228
[7 empty] | 2277 | 12.5361 | [20 empty] | -0.091942 | 6.2308 | 22.2622 | -51.9137 | 90.3242 | -116.4802 | 111.2732 | -66.8322 | 17.7993 | [12 empty] | 1.2716 | 3.9261 | 10.6442 | -3.339 | 0.033287
[7 empty] | 2278 | 12.5429 | [20 empty] | -0.091942 | 6.2335 | 22.2818 | -51.9821 | 90.483 | -116.7362 | 111.5668 | -67.0379 | 17.8619 | [12 empty] | 1.2716 | 3.9278 | 10.6535 | -3.3434 | 0.033345
[7 empty] | 2279 | 12.5496 | [20 empty] | -0.091942 | 6.2362 | 22.3014 | -52.0506 | 90.642 | -116.9927 | 111.8609 | -67.2442 | 17.9248 | [12 empty] | 1.2716 | 3.9295 | 10.6629 | -3.3478 | 0.033404
[7 empty] | 2280 | 12.5563 | [20 empty] | -0.091942 | 6.239 | 22.3209 | -52.1192 | 90.8011 | -117.2496 | 112.1558 | -67.451 | 17.9878 | [12 empty] | 1.2716 | 3.9313 | 10.6722 | -3.3522 | 0.033462
[7 empty] | 2281 | 12.5631 | [20 empty] | -0.091942 | 6.2417 | 22.3405 | -52.1878 | 90.9606 | -117.5069 | 112.4512 | -67.6583 | 18.051 | [12 empty] | 1.2716 | 3.933 | 10.6816 | -3.3567 | 0.033521
[7 empty] | 2282 | 12.5698 | [20 empty] | -0.091942 | 6.2444 | 22.3601 | -52.2564 | 91.1202 | -117.7647 | 112.7474 | -67.8662 | 18.1144 | [12 empty] | 1.2716 | 3.9347 | 10.691 | -3.3611 | 0.03358
[7 empty] | 2283 | 12.5765 | [20 empty] | -0.091942 | 6.2472 | 22.3797 | -52.3252 | 91.28 | -118.023 | 113.0441 | -68.0747 | 18.178 | [12 empty] | 1.2716 | 3.9364 | 10.7003 | -3.3655 | 0.033639
[7 empty] | 2284 | 12.5833 | [20 empty] | -0.091942 | 6.2499 | 22.3993 | -52.394 | 91.44 | -118.2817 | 113.3415 | -68.2837 | 18.2418 | [12 empty] | 1.2716 | 3.9382 | 10.7097 | -3.3699 | 0.033698
[7 empty] | 2285 | 12.59 | [20 empty] | -0.091942 | 6.2526 | 22.419 | -52.4628 | 91.6003 | -118.5408 | 113.6396 | -68.4933 | 18.3058 | [12 empty] | 1.2716 | 3.9399 | 10.7191 | -3.3743 | 0.033757
[7 empty] | 2286 | 12.5967 | [20 empty] | -0.091942 | 6.2554 | 22.4386 | -52.5317 | 91.7607 | -118.8005 | 113.9383 | -68.7034 | 18.37 | [12 empty] | 1.2716 | 3.9416 | 10.7285 | -3.3788 | 0.033816
[7 empty] | 2287 | 12.6035 | [20 empty] | -0.091942 | 6.2581 | 22.4582 | -52.6007 | 91.9214 | -119.0605 | 114.2377 | -68.914 | 18.4344 | [12 empty] | 1.2716 | 3.9433 | 10.7379 | -3.3832 | 0.033875
[7 empty] | 2288 | 12.6102 | [20 empty] | -0.091942 | 6.2609 | 22.4779 | -52.6697 | 92.0823 | -119.321 | 114.5378 | -69.1252 | 18.499 | [12 empty] | 1.2716 | 3.945 | 10.7473 | -3.3877 | 0.033935
[7 empty] | 2289 | 12.6169 | [20 empty] | -0.091942 | 6.2636 | 22.4975 | -52.7388 | 92.2434 | -119.582 | 114.8384 | -69.337 | 18.5637 | [12 empty] | 1.2716 | 3.9468 | 10.7567 | -3.3921 | 0.033994
[7 empty] | 2290 | 12.6237 | [20 empty] | -0.091942 | 6.2663 | 22.5172 | -52.808 | 92.4047 | -119.8435 | 115.1398 | -69.5493 | 18.6287 | [12 empty] | 1.2716 | 3.9485 | 10.7661 | -3.3965 | 0.034053
[7 empty] | 2291 | 12.6304 | [20 empty] | -0.091942 | 6.2691 | 22.5368 | -52.8772 | 92.5662 | -120.1054 | 115.4418 | -69.7622 | 18.6939 | [12 empty] | 1.2716 | 3.9502 | 10.7755 | -3.401 | 0.034113
[7 empty] | 2292 | 12.6371 | [20 empty] | -0.091942 | 6.2718 | 22.5565 | -52.9464 | 92.7279 | -120.3677 | 115.7445 | -69.9756 | 18.7593 | [12 empty] | 1.2716 | 3.9519 | 10.7849 | -3.4055 | 0.034173
[7 empty] | 2293 | 12.6439 | [20 empty] | -0.091942 | 6.2745 | 22.5762 | -53.0158 | 92.8898 | -120.6305 | 116.0478 | -70.1896 | 18.8248 | [12 empty] | 1.2716 | 3.9537 | 10.7943 | -3.4099 | 0.034232
[7 empty] | 2294 | 12.6506 | [20 empty] | -0.091942 | 6.2773 | 22.5959 | -53.0852 | 93.052 | -120.8938 | 116.3518 | -70.4041 | 18.8906 | [12 empty] | 1.2716 | 3.9554 | 10.8037 | -3.4144 | 0.034292
[7 empty] | 2295 | 12.6574 | [20 empty] | -0.091942 | 6.28 | 22.6156 | -53.1546 | 93.2143 | -121.1575 | 116.6564 | -70.6193 | 18.9566 | [12 empty] | 1.2716 | 3.9571 | 10.8131 | -3.4188 | 0.034352
[7 empty] | 2296 | 12.6641 | [20 empty] | -0.091942 | 6.2827 | 22.6353 | -53.2241 | 93.3769 | -121.4217 | 116.9617 | -70.8349 | 19.0228 | [12 empty] | 1.2716 | 3.9588 | 10.8225 | -3.4233 | 0.034412
[7 empty] | 2297 | 12.6708 | [20 empty] | -0.091942 | 6.2855 | 22.655 | -53.2937 | 93.5397 | -121.6864 | 117.2677 | -71.0512 | 19.0892 | [12 empty] | 1.2716 | 3.9606 | 10.832 | -3.4278 | 0.034472
[7 empty] | 2298 | 12.6776 | [20 empty] | -0.091942 | 6.2882 | 22.6748 | -53.3633 | 93.7027 | -121.9515 | 117.5744 | -71.268 | 19.1558 | [12 empty] | 1.2716 | 3.9623 | 10.8414 | -3.4323 | 0.034532
[7 empty] | 2299 | 12.6843 | [20 empty] | -0.091942 | 6.291 | 22.6945 | -53.433 | 93.8659 | -122.2171 | 117.8817 | -71.4854 | 19.2225 | [12 empty] | 1.2716 | 3.964 | 10.8508 | -3.4368 | 0.034592
[7 empty] | 2300 | 12.6911 | [20 empty] | -0.091942 | 6.2937 | 22.7143 | -53.5028 | 94.0293 | -122.4831 | 118.1897 | -71.7033 | 19.2895 | [12 empty] | 1.2716 | 3.9657 | 10.8603 | -3.4412 | 0.034652
[7 empty] | 2301 | 12.6978 | [20 empty] | -0.091942 | 6.2964 | 22.734 | -53.5726 | 94.193 | -122.7496 | 118.4983 | -71.9218 | 19.3567 | [12 empty] | 1.2716 | 3.9675 | 10.8697 | -3.4457 | 0.034712
[7 empty] | 2302 | 12.7045 | [20 empty] | -0.091942 | 6.2992 | 22.7538 | -53.6425 | 94.3568 | -123.0166 | 118.8077 | -72.1409 | 19.4241 | [12 empty] | 1.2716 | 3.9692 | 10.8792 | -3.4502 | 0.034773
[7 empty] | 2303 | 12.7113 | [20 empty] | -0.091942 | 6.3019 | 22.7736 | -53.7124 | 94.5209 | -123.284 | 119.1177 | -72.3606 | 19.4917 | [12 empty] | 1.2716 | 3.9709 | 10.8886 | -3.4547 | 0.034833
[7 empty] | 2304 | 12.718 | [20 empty] | -0.091942 | 6.3046 | 22.7933 | -53.7824 | 94.6851 | -123.5519 | 119.4283 | -72.5808 | 19.5595 | [12 empty] | 1.2716 | 3.9726 | 10.8981 | -3.4592 | 0.034894
[7 empty] | 2305 | 12.7248 | [20 empty] | -0.091942 | 6.3074 | 22.8131 | -53.8525 | 94.8496 | -123.8202 | 119.7397 | -72.8016 | 19.6276 | [12 empty] | 1.2716 | 3.9744 | 10.9076 | -3.4637 | 0.034954
[7 empty] | 2306 | 12.7315 | [20 empty] | -0.091942 | 6.3101 | 22.8329 | -53.9226 | 95.0143 | -124.0891 | 120.0517 | -73.023 | 19.6958 | [12 empty] | 1.2716 | 3.9761 | 10.917 | -3.4682 | 0.035015
[7 empty] | 2307 | 12.7382 | [20 empty] | -0.091942 | 6.3128 | 22.8527 | -53.9928 | 95.1793 | -124.3583 | 120.3644 | -73.2449 | 19.7642 | [12 empty] | 1.2716 | 3.9778 | 10.9265 | -3.4728 | 0.035076
[7 empty] | 2308 | 12.745 | [20 empty] | -0.091942 | 6.3156 | 22.8725 | -54.063 | 95.3444 | -124.6281 | 120.6778 | -73.4675 | 19.8329 | [12 empty] | 1.2716 | 3.9795 | 10.936 | -3.4773 | 0.035137
[7 empty] | 2309 | 12.7517 | [20 empty] | -0.091942 | 6.3183 | 22.8924 | -54.1333 | 95.5097 | -124.8983 | 120.9919 | -73.6906 | 19.9017 | [12 empty] | 1.2716 | 3.9813 | 10.9454 | -3.4818 | 0.035198
[7 empty] | 2310 | 12.7585 | [20 empty] | -0.091942 | 6.3211 | 22.9122 | -54.2037 | 95.6753 | -125.169 | 121.3066 | -73.9143 | 19.9708 | [12 empty] | 1.2716 | 3.983 | 10.9549 | -3.4863 | 0.035259
[7 empty] | 2311 | 12.7652 | [20 empty] | -0.091942 | 6.3238 | 22.932 | -54.2741 | 95.8411 | -125.4402 | 121.622 | -74.1385 | 20.04 | [12 empty] | 1.2716 | 3.9847 | 10.9644 | -3.4908 | 0.03532
[7 empty] | 2312 | 12.7719 | [20 empty] | -0.091942 | 6.3265 | 22.9519 | -54.3446 | 96.0071 | -125.7118 | 121.9381 | -74.3634 | 20.1095 | [12 empty] | 1.2716 | 3.9864 | 10.9739 | -3.4954 | 0.035381
[7 empty] | 2313 | 12.7787 | [20 empty] | -0.091942 | 6.3293 | 22.9718 | -54.4151 | 96.1733 | -125.9839 | 122.2549 | -74.5888 | 20.1792 | [12 empty] | 1.2716 | 3.9882 | 10.9834 | -3.4999 | 0.035442
[7 empty] | 2314 | 12.7854 | [20 empty] | -0.091942 | 6.332 | 22.9916 | -54.4858 | 96.3397 | -126.2565 | 122.5724 | -74.8149 | 20.2491 | [12 empty] | 1.2716 | 3.9899 | 10.9929 | -3.5045 | 0.035504
[7 empty] | 2315 | 12.7922 | [20 empty] | -0.091942 | 6.3347 | 23.0115 | -54.5564 | 96.5064 | -126.5295 | 122.8906 | -75.0415 | 20.3192 | [12 empty] | 1.2716 | 3.9916 | 11.0024 | -3.509 | 0.035565
[7 empty] | 2316 | 12.7989 | [20 empty] | -0.091942 | 6.3375 | 23.0314 | -54.6272 | 96.6732 | -126.8031 | 123.2094 | -75.2687 | 20.3895 | [12 empty] | 1.2716 | 3.9933 | 11.0119 | -3.5136 | 0.035626
[7 empty] | 2317 | 12.8057 | [20 empty] | -0.091942 | 6.3402 | 23.0513 | -54.6979 | 96.8403 | -127.0771 | 123.5289 | -75.4965 | 20.4601 | [12 empty] | 1.2716 | 3.9951 | 11.0214 | -3.5181 | 0.035688
[7 empty] | 2318 | 12.8124 | [20 empty] | -0.091942 | 6.3429 | 23.0712 | -54.7688 | 97.0076 | -127.3515 | 123.8492 | -75.7248 | 20.5308 | [12 empty] | 1.2716 | 3.9968 | 11.0309 | -3.5227 | 0.03575
[7 empty] | 2319 | 12.8192 | [20 empty] | -0.091942 | 6.3457 | 23.0911 | -54.8397 | 97.1751 | -127.6265 | 124.1701 | -75.9538 | 20.6018 | [12 empty] | 1.2716 | 3.9985 | 11.0405 | -3.5272 | 0.035811
[7 empty] | 2320 | 12.8259 | [20 empty] | -0.091942 | 6.3484 | 23.111 | -54.9107 | 97.3428 | -127.9019 | 124.4917 | -76.1834 | 20.673 | [12 empty] | 1.2716 | 4.0002 | 11.05 | -3.5318 | 0.035873
[7 empty] | 2321 | 12.8326 | [20 empty] | -0.091942 | 6.3512 | 23.1309 | -54.9817 | 97.5108 | -128.1778 | 124.814 | -76.4135 | 20.7444 | [12 empty] | 1.2716 | 4.0019 | 11.0595 | -3.5364 | 0.035935
[7 empty] | 2322 | 12.8394 | [20 empty] | -0.091942 | 6.3539 | 23.1509 | -55.0528 | 97.6789 | -128.4541 | 125.137 | -76.6443 | 20.816 | [12 empty] | 1.2716 | 4.0037 | 11.069 | -3.5409 | 0.035997
[7 empty] | 2323 | 12.8461 | [20 empty] | -0.091942 | 6.3566 | 23.1708 | -55.124 | 97.8473 | -128.731 | 125.4607 | -76.8757 | 20.8878 | [12 empty] | 1.2716 | 4.0054 | 11.0786 | -3.5455 | 0.036059
[7 empty] | 2324 | 12.8529 | [20 empty] | -0.091942 | 6.3594 | 23.1908 | -55.1952 | 98.0159 | -129.0083 | 125.7851 | -77.1076 | 20.9598 | [12 empty] | 1.2716 | 4.0071 | 11.0881 | -3.5501 | 0.036121
[7 empty] | 2325 | 12.8596 | [20 empty] | -0.091942 | 6.3621 | 23.2107 | -55.2665 | 98.1847 | -129.2861 | 126.1102 | -77.3402 | 21.0321 | [12 empty] | 1.2716 | 4.0088 | 11.0977 | -3.5547 | 0.036183
[7 empty] | 2326 | 12.8664 | [20 empty] | -0.091942 | 6.3648 | 23.2307 | -55.3378 | 98.3537 | -129.5643 | 126.436 | -77.5733 | 21.1046 | [12 empty] | 1.2716 | 4.0106 | 11.1072 | -3.5593 | 0.036246
[7 empty] | 2327 | 12.8731 | [20 empty] | -0.091942 | 6.3676 | 23.2507 | -55.4092 | 98.523 | -129.8431 | 126.7625 | -77.8071 | 21.1773 | [12 empty] | 1.2716 | 4.0123 | 11.1168 | -3.5639 | 0.036308
[7 empty] | 2328 | 12.8799 | [20 empty] | -0.091942 | 6.3703 | 23.2707 | -55.4807 | 98.6924 | -130.1223 | 127.0897 | -78.0414 | 21.2502 | [12 empty] | 1.2716 | 4.014 | 11.1263 | -3.5685 | 0.036371
[7 empty] | 2329 | 12.8866 | [20 empty] | -0.091942 | 6.373 | 23.2907 | -55.5522 | 98.8621 | -130.402 | 127.4176 | -78.2764 | 21.3233 | [12 empty] | 1.2716 | 4.0157 | 11.1359 | -3.5731 | 0.036433
[7 empty] | 2330 | 12.8934 | [20 empty] | -0.091942 | 6.3758 | 23.3107 | -55.6238 | 99.032 | -130.6822 | 127.7462 | -78.512 | 21.3967 | [12 empty] | 1.2716 | 4.0175 | 11.1454 | -3.5777 | 0.036496
[7 empty] | 2331 | 12.9001 | [20 empty] | -0.091942 | 6.3785 | 23.3307 | -55.6955 | 99.2021 | -130.9629 | 128.0755 | -78.7481 | 21.4703 | [12 empty] | 1.2716 | 4.0192 | 11.155 | -3.5823 | 0.036558
[7 empty] | 2332 | 12.9069 | [20 empty] | -0.091942 | 6.3813 | 23.3507 | -55.7672 | 99.3725 | -131.2441 | 128.4056 | -78.9849 | 21.544 | [12 empty] | 1.2716 | 4.0209 | 11.1646 | -3.5869 | 0.036621
[7 empty] | 2333 | 12.9136 | [20 empty] | -0.091942 | 6.384 | 23.3707 | -55.8389 | 99.543 | -131.5257 | 128.7363 | -79.2223 | 21.6181 | [12 empty] | 1.2716 | 4.0226 | 11.1742 | -3.5915 | 0.036684
[7 empty] | 2334 | 12.9204 | [20 empty] | -0.091942 | 6.3867 | 23.3908 | -55.9108 | 99.7138 | -131.8078 | 129.0677 | -79.4603 | 21.6923 | [12 empty] | 1.2716 | 4.0244 | 11.1837 | -3.5961 | 0.036747
[7 empty] | 2335 | 12.9271 | [20 empty] | -0.091942 | 6.3895 | 23.4108 | -55.9827 | 99.8848 | -132.0904 | 129.3999 | -79.6989 | 21.7668 | [12 empty] | 1.2716 | 4.0261 | 11.1933 | -3.6007 | 0.03681
[7 empty] | 2336 | 12.9339 | [20 empty] | -0.091942 | 6.3922 | 23.4309 | -56.0546 | 100.056 | -132.3735 | 129.7327 | -79.9382 | 21.8415 | [12 empty] | 1.2716 | 4.0278 | 11.2029 | -3.6054 | 0.036873
[7 empty] | 2337 | 12.9406 | [20 empty] | -0.091942 | 6.3949 | 23.4509 | -56.1266 | 100.2275 | -132.6571 | 130.0663 | -80.178 | 21.9164 | [12 empty] | 1.2716 | 4.0295 | 11.2125 | -3.61 | 0.036936
[7 empty] | 2338 | 12.9473 | [20 empty] | -0.091942 | 6.3977 | 23.471 | -56.1987 | 100.3991 | -132.9412 | 130.4006 | -80.4185 | 21.9915 | [12 empty] | 1.2716 | 4.0313 | 11.2221 | -3.6146 | 0.037
[7 empty] | 2339 | 12.9541 | [20 empty] | -0.091942 | 6.4004 | 23.4911 | -56.2709 | 100.571 | -133.2257 | 130.7356 | -80.6596 | 22.0669 | [12 empty] | 1.2716 | 4.033 | 11.2317 | -3.6193 | 0.037063
[7 empty] | 2340 | 12.9608 | [20 empty] | -0.091942 | 6.4031 | 23.5112 | -56.3431 | 100.7431 | -133.5108 | 131.0713 | -80.9013 | 22.1425 | [12 empty] | 1.2716 | 4.0347 | 11.2413 | -3.6239 | 0.037126
[7 empty] | 2341 | 12.9676 | [20 empty] | -0.091942 | 6.4059 | 23.5313 | -56.4153 | 100.9154 | -133.7963 | 131.4078 | -81.1436 | 22.2183 | [12 empty] | 1.2716 | 4.0364 | 11.2509 | -3.6286 | 0.03719
[7 empty] | 2342 | 12.9743 | [20 empty] | -0.091942 | 6.4086 | 23.5514 | -56.4877 | 101.088 | -134.0823 | 131.7449 | -81.3865 | 22.2943 | [12 empty] | 1.2716 | 4.0382 | 11.2605 | -3.6332 | 0.037253
[7 empty] | 2343 | 12.9811 | [20 empty] | -0.091942 | 6.4114 | 23.5715 | -56.56 | 101.2607 | -134.3688 | 132.0828 | -81.6301 | 22.3706 | [12 empty] | 1.2716 | 4.0399 | 11.2702 | -3.6379 | 0.037317
[7 empty] | 2344 | 12.9878 | [20 empty] | -0.091942 | 6.4141 | 23.5916 | -56.6325 | 101.4337 | -134.6558 | 132.4214 | -81.8743 | 22.4471 | [12 empty] | 1.2716 | 4.0416 | 11.2798 | -3.6425 | 0.037381
[7 empty] | 2345 | 12.9946 | [20 empty] | -0.091942 | 6.4168 | 23.6118 | -56.705 | 101.6069 | -134.9433 | 132.7607 | -82.1191 | 22.5238 | [12 empty] | 1.2716 | 4.0433 | 11.2894 | -3.6472 | 0.037445
[7 empty] | 2346 | 13.0013 | [20 empty] | -0.091942 | 6.4196 | 23.6319 | -56.7776 | 101.7804 | -135.2312 | 133.1008 | -82.3646 | 22.6008 | [12 empty] | 1.2716 | 4.0451 | 11.299 | -3.6519 | 0.037509
[7 empty] | 2347 | 13.0081 | [20 empty] | -0.091942 | 6.4223 | 23.6521 | -56.8502 | 101.954 | -135.5197 | 133.4415 | -82.6106 | 22.6779 | [12 empty] | 1.2716 | 4.0468 | 11.3087 | -3.6565 | 0.037573
[7 empty] | 2348 | 13.0149 | [20 empty] | -0.091942 | 6.425 | 23.6722 | -56.9229 | 102.1279 | -135.8086 | 133.783 | -82.8573 | 22.7554 | [12 empty] | 1.2716 | 4.0485 | 11.3183 | -3.6612 | 0.037637
[7 empty] | 2349 | 13.0216 | [20 empty] | -0.091942 | 6.4278 | 23.6924 | -56.9957 | 102.302 | -136.0981 | 134.1253 | -83.1047 | 22.833 | [12 empty] | 1.2716 | 4.0502 | 11.328 | -3.6659 | 0.037701
[7 empty] | 2350 | 13.0284 | [20 empty] | -0.091942 | 6.4305 | 23.7126 | -57.0685 | 102.4763 | -136.388 | 134.4682 | -83.3526 | 22.9109 | [12 empty] | 1.2716 | 4.052 | 11.3376 | -3.6706 | 0.037765
[7 empty] | 2351 | 13.0351 | [20 empty] | -0.091942 | 6.4332 | 23.7328 | -57.1414 | 102.6508 | -136.6785 | 134.8119 | -83.6012 | 22.989 | [12 empty] | 1.2716 | 4.0537 | 11.3473 | -3.6753 | 0.037829
[7 empty] | 2352 | 13.0419 | [20 empty] | -0.091942 | 6.436 | 23.753 | -57.2143 | 102.8256 | -136.9694 | 135.1563 | -83.8505 | 23.0673 | [12 empty] | 1.2716 | 4.0554 | 11.3569 | -3.68 | 0.037894
[7 empty] | 2353 | 13.0486 | [20 empty] | -0.091942 | 6.4387 | 23.7732 | -57.2873 | 103.0006 | -137.2608 | 135.5015 | -84.1004 | 23.1459 | [12 empty] | 1.2716 | 4.0571 | 11.3666 | -3.6847 | 0.037958
[7 empty] | 2354 | 13.0554 | [20 empty] | -0.091942 | 6.4415 | 23.7934 | -57.3604 | 103.1758 | -137.5527 | 135.8474 | -84.3509 | 23.2247 | [12 empty] | 1.2716 | 4.0588 | 11.3762 | -3.6894 | 0.038023
[7 empty] | 2355 | 13.0621 | [20 empty] | -0.091942 | 6.4442 | 23.8136 | -57.4335 | 103.3512 | -137.8452 | 136.194 | -84.602 | 23.3038 | [12 empty] | 1.2716 | 4.0606 | 11.3859 | -3.6941 | 0.038087
[7 empty] | 2356 | 13.0689 | [20 empty] | -0.091942 | 6.4469 | 23.8338 | -57.5067 | 103.5269 | -138.1381 | 136.5414 | -84.8538 | 23.3831 | [12 empty] | 1.2716 | 4.0623 | 11.3956 | -3.6988 | 0.038152
[7 empty] | 2357 | 13.0756 | [20 empty] | -0.091942 | 6.4497 | 23.854 | -57.58 | 103.7028 | -138.4315 | 136.8895 | -85.1062 | 23.4626 | [12 empty] | 1.2716 | 4.064 | 11.4052 | -3.7035 | 0.038217
[7 empty] | 2358 | 13.0824 | [20 empty] | -0.091942 | 6.4524 | 23.8743 | -57.6533 | 103.8789 | -138.7254 | 137.2383 | -85.3593 | 23.5423 | [12 empty] | 1.2716 | 4.0657 | 11.4149 | -3.7082 | 0.038282
[7 empty] | 2359 | 13.0891 | [20 empty] | -0.091942 | 6.4551 | 23.8945 | -57.7267 | 104.0552 | -139.0198 | 137.5879 | -85.613 | 23.6223 | [12 empty] | 1.2716 | 4.0675 | 11.4246 | -3.7129 | 0.038347
[7 empty] | 2360 | 13.0959 | [20 empty] | -0.091942 | 6.4579 | 23.9148 | -57.8001 | 104.2317 | -139.3147 | 137.9382 | -85.8674 | 23.7025 | [12 empty] | 1.2716 | 4.0692 | 11.4343 | -3.7176 | 0.038412
[7 empty] | 2361 | 13.1026 | [20 empty] | -0.091942 | 6.4606 | 23.9351 | -57.8736 | 104.4085 | -139.6101 | 138.2893 | -86.1224 | 23.783 | [12 empty] | 1.2716 | 4.0709 | 11.444 | -3.7224 | 0.038477
[7 empty] | 2362 | 13.1094 | [20 empty] | -0.091942 | 6.4633 | 23.9554 | -57.9472 | 104.5855 | -139.906 | 138.6411 | -86.3781 | 23.8637 | [12 empty] | 1.2716 | 4.0726 | 11.4537 | -3.7271 | 0.038542
[7 empty] | 2363 | 13.1161 | [20 empty] | -0.091942 | 6.4661 | 23.9756 | -58.0208 | 104.7627 | -140.2024 | 138.9936 | -86.6344 | 23.9447 | [12 empty] | 1.2716 | 4.0744 | 11.4634 | -3.7318 | 0.038608
[7 empty] | 2364 | 13.1229 | [20 empty] | -0.091942 | 6.4688 | 23.9959 | -58.0945 | 104.9402 | -140.4994 | 139.3469 | -86.8914 | 24.0258 | [12 empty] | 1.2716 | 4.0761 | 11.4731 | -3.7366 | 0.038673
[7 empty] | 2365 | 13.1297 | [20 empty] | -0.091942 | 6.4716 | 24.0163 | -58.1683 | 105.1179 | -140.7968 | 139.701 | -87.149 | 24.1073 | [12 empty] | 1.2716 | 4.0778 | 11.4828 | -3.7413 | 0.038739
[7 empty] | 2366 | 13.1364 | [20 empty] | -0.091942 | 6.4743 | 24.0366 | -58.2421 | 105.2958 | -141.0947 | 140.0558 | -87.4073 | 24.1889 | [12 empty] | 1.2716 | 4.0795 | 11.4925 | -3.7461 | 0.038804
[7 empty] | 2367 | 13.1432 | [20 empty] | -0.091942 | 6.477 | 24.0569 | -58.316 | 105.4739 | -141.3931 | 140.4113 | -87.6662 | 24.2709 | [12 empty] | 1.2716 | 4.0813 | 11.5022 | -3.7508 | 0.03887
[7 empty] | 2368 | 13.1499 | [20 empty] | -0.091942 | 6.4798 | 24.0772 | -58.3899 | 105.6522 | -141.692 | 140.7676 | -87.9258 | 24.353 | [12 empty] | 1.2716 | 4.083 | 11.512 | -3.7556 | 0.038935
[7 empty] | 2369 | 13.1567 | [20 empty] | -0.091942 | 6.4825 | 24.0976 | -58.4639 | 105.8308 | -141.9915 | 141.1247 | -88.186 | 24.4354 | [12 empty] | 1.2716 | 4.0847 | 11.5217 | -3.7603 | 0.039001
[7 empty] | 2370 | 13.1634 | [20 empty] | -0.091942 | 6.4852 | 24.1179 | -58.538 | 106.0096 | -142.2914 | 141.4825 | -88.4469 | 24.518 | [12 empty] | 1.2716 | 4.0864 | 11.5314 | -3.7651 | 0.039067
[7 empty] | 2371 | 13.1702 | [20 empty] | -0.091942 | 6.488 | 24.1383 | -58.6121 | 106.1887 | -142.5919 | 141.841 | -88.7085 | 24.6009 | [12 empty] | 1.2716 | 4.0882 | 11.5411 | -3.7699 | 0.039133
[7 empty] | 2372 | 13.1769 | [20 empty] | -0.091942 | 6.4907 | 24.1586 | -58.6863 | 106.3679 | -142.8928 | 142.2003 | -88.9707 | 24.6841 | [12 empty] | 1.2716 | 4.0899 | 11.5509 | -3.7746 | 0.039199
[7 empty] | 2373 | 13.1837 | [20 empty] | -0.091942 | 6.4934 | 24.179 | -58.7606 | 106.5474 | -143.1943 | 142.5604 | -89.2336 | 24.7674 | [12 empty] | 1.2716 | 4.0916 | 11.5606 | -3.7794 | 0.039265
[7 empty] | 2374 | 13.1904 | [20 empty] | -0.091942 | 6.4962 | 24.1994 | -58.8349 | 106.7271 | -143.4962 | 142.9213 | -89.4972 | 24.851 | [12 empty] | 1.2716 | 4.0933 | 11.5704 | -3.7842 | 0.039332
[7 empty] | 2375 | 13.1972 | [20 empty] | -0.091942 | 6.4989 | 24.2198 | -58.9093 | 106.9071 | -143.7987 | 143.2829 | -89.7614 | 24.9349 | [12 empty] | 1.2716 | 4.0951 | 11.5801 | -3.789 | 0.039398
[7 empty] | 2376 | 13.204 | [20 empty] | -0.091942 | 6.5017 | 24.2402 | -58.9837 | 107.0872 | -144.1017 | 143.6452 | -90.0263 | 25.019 | [12 empty] | 1.2716 | 4.0968 | 11.5899 | -3.7938 | 0.039464
[7 empty] | 2377 | 13.2107 | [20 empty] | -0.091942 | 6.5044 | 24.2606 | -59.0582 | 107.2676 | -144.4052 | 144.0083 | -90.2919 | 25.1034 | [12 empty] | 1.2716 | 4.0985 | 11.5996 | -3.7986 | 0.039531
[7 empty] | 2378 | 13.2175 | [20 empty] | -0.091942 | 6.5071 | 24.281 | -59.1328 | 107.4482 | -144.7092 | 144.3722 | -90.5581 | 25.188 | [12 empty] | 1.2716 | 4.1002 | 11.6094 | -3.8034 | 0.039597
[7 empty] | 2379 | 13.2242 | [20 empty] | -0.091942 | 6.5099 | 24.3014 | -59.2074 | 107.6291 | -145.0137 | 144.7369 | -90.825 | 25.2729 | [12 empty] | 1.2716 | 4.102 | 11.6192 | -3.8082 | 0.039664
[7 empty] | 2380 | 13.231 | [20 empty] | -0.091942 | 6.5126 | 24.3219 | -59.2821 | 107.8102 | -145.3188 | 145.1023 | -91.0926 | 25.358 | [12 empty] | 1.2716 | 4.1037 | 11.6289 | -3.813 | 0.039731
[7 empty] | 2381 | 13.2377 | [20 empty] | -0.091942 | 6.5153 | 24.3423 | -59.3569 | 107.9915 | -145.6243 | 145.4685 | -91.3608 | 25.4433 | [12 empty] | 1.2716 | 4.1054 | 11.6387 | -3.8178 | 0.039798
[7 empty] | 2382 | 13.2445 | [20 empty] | -0.091942 | 6.5181 | 24.3628 | -59.4317 | 108.173 | -145.9304 | 145.8354 | -91.6298 | 25.529 | [12 empty] | 1.2716 | 4.1071 | 11.6485 | -3.8226 | 0.039864
[7 empty] | 2383 | 13.2513 | [20 empty] | -0.091942 | 6.5208 | 24.3832 | -59.5066 | 108.3548 | -146.237 | 146.2032 | -91.8994 | 25.6148 | [12 empty] | 1.2716 | 4.1089 | 11.6583 | -3.8274 | 0.039931
[7 empty] | 2384 | 13.258 | [20 empty] | -0.091942 | 6.5235 | 24.4037 | -59.5815 | 108.5368 | -146.5441 | 146.5717 | -92.1697 | 25.7009 | [12 empty] | 1.2716 | 4.1106 | 11.668 | -3.8322 | 0.039998
[7 empty] | 2385 | 13.2648 | [20 empty] | -0.091942 | 6.5263 | 24.4242 | -59.6565 | 108.719 | -146.8517 | 146.9409 | -92.4407 | 25.7873 | [12 empty] | 1.2716 | 4.1123 | 11.6778 | -3.837 | 0.040066
[7 empty] | 2386 | 13.2715 | [20 empty] | -0.091942 | 6.529 | 24.4446 | -59.7316 | 108.9015 | -147.1598 | 147.311 | -92.7123 | 25.8739 | [12 empty] | 1.2716 | 4.114 | 11.6876 | -3.8419 | 0.040133
[7 empty] | 2387 | 13.2783 | [20 empty] | -0.091942 | 6.5318 | 24.4651 | -59.8067 | 109.0841 | -147.4684 | 147.6818 | -92.9847 | 25.9608 | [12 empty] | 1.2716 | 4.1157 | 11.6974 | -3.8467 | 0.0402
[7 empty] | 2388 | 13.285 | [20 empty] | -0.091942 | 6.5345 | 24.4856 | -59.8819 | 109.2671 | -147.7776 | 148.0534 | -93.2577 | 26.048 | [12 empty] | 1.2716 | 4.1175 | 11.7072 | -3.8515 | 0.040268
[7 empty] | 2389 | 13.2918 | [20 empty] | -0.091942 | 6.5372 | 24.5062 | -59.9572 | 109.4502 | -148.0873 | 148.4258 | -93.5314 | 26.1353 | [12 empty] | 1.2716 | 4.1192 | 11.717 | -3.8564 | 0.040335
[7 empty] | 2390 | 13.2986 | [20 empty] | -0.091942 | 6.54 | 24.5267 | -60.0325 | 109.6336 | -148.3975 | 148.799 | -93.8058 | 26.223 | [12 empty] | 1.2716 | 4.1209 | 11.7268 | -3.8612 | 0.040403
[7 empty] | 2391 | 13.3053 | [20 empty] | -0.091942 | 6.5427 | 24.5472 | -60.1079 | 109.8172 | -148.7082 | 149.1729 | -94.0809 | 26.3109 | [12 empty] | 1.2716 | 4.1226 | 11.7367 | -3.8661 | 0.04047
[7 empty] | 2392 | 13.3121 | [20 empty] | -0.091942 | 6.5454 | 24.5677 | -60.1834 | 110.001 | -149.0194 | 149.5476 | -94.3567 | 26.3991 | [12 empty] | 1.2716 | 4.1244 | 11.7465 | -3.8709 | 0.040538
[7 empty] | 2393 | 13.3188 | [20 empty] | -0.091942 | 6.5482 | 24.5883 | -60.2589 | 110.1851 | -149.3312 | 149.9232 | -94.6331 | 26.4875 | [12 empty] | 1.2716 | 4.1261 | 11.7563 | -3.8758 | 0.040606
[7 empty] | 2394 | 13.3256 | [20 empty] | -0.091942 | 6.5509 | 24.6088 | -60.3344 | 110.3694 | -149.6434 | 150.2995 | -94.9103 | 26.5762 | [12 empty] | 1.2716 | 4.1278 | 11.7661 | -3.8806 | 0.040674
[7 empty] | 2395 | 13.3324 | [20 empty] | -0.091942 | 6.5536 | 24.6294 | -60.4101 | 110.5539 | -149.9562 | 150.6765 | -95.1882 | 26.6651 | [12 empty] | 1.2716 | 4.1295 | 11.776 | -3.8855 | 0.040742
[7 empty] | 2396 | 13.3391 | [20 empty] | -0.091942 | 6.5564 | 24.65 | -60.4858 | 110.7386 | -150.2696 | 151.0544 | -95.4667 | 26.7543 | [12 empty] | 1.2716 | 4.1313 | 11.7858 | -3.8904 | 0.04081
[7 empty] | 2397 | 13.3459 | [20 empty] | -0.091942 | 6.5591 | 24.6706 | -60.5616 | 110.9236 | -150.5834 | 151.4331 | -95.746 | 26.8438 | [12 empty] | 1.2716 | 4.133 | 11.7956 | -3.8953 | 0.040878
[7 empty] | 2398 | 13.3526 | [20 empty] | -0.091942 | 6.5619 | 24.6911 | -60.6374 | 111.1089 | -150.8978 | 151.8125 | -96.0259 | 26.9335 | [12 empty] | 1.2716 | 4.1347 | 11.8055 | -3.9001 | 0.040946
[7 empty] | 2399 | 13.3594 | [20 empty] | -0.091942 | 6.5646 | 24.7117 | -60.7133 | 111.2943 | -151.2127 | 152.1928 | -96.3066 | 27.0235 | [12 empty] | 1.2716 | 4.1364 | 11.8153 | -3.905 | 0.041015
[7 empty] | 2400 | 13.3662 | [20 empty] | -0.091942 | 6.5673 | 24.7324 | -60.7892 | 111.48 | -151.5281 | 152.5738 | -96.588 | 27.1137 | [12 empty] | 1.2716 | 4.1382 | 11.8252 | -3.9099 | 0.041083
[7 empty] | 2401 | 13.3729 | [20 empty] | -0.091942 | 6.5701 | 24.753 | -60.8652 | 111.6659 | -151.844 | 152.9556 | -96.87 | 27.2042 | [12 empty] | 1.2716 | 4.1399 | 11.835 | -3.9148 | 0.041152
[7 empty] | 2402 | 13.3797 | [20 empty] | -0.091942 | 6.5728 | 24.7736 | -60.9413 | 111.8521 | -152.1605 | 153.3383 | -97.1528 | 27.295 | [12 empty] | 1.2716 | 4.1416 | 11.8449 | -3.9197 | 0.04122
[7 empty] | 2403 | 13.3864 | [20 empty] | -0.091942 | 6.5755 | 24.7942 | -61.0175 | 112.0384 | -152.4775 | 153.7217 | -97.4363 | 27.386 | [12 empty] | 1.2716 | 4.1433 | 11.8548 | -3.9246 | 0.041289
[7 empty] | 2404 | 13.3932 | [20 empty] | -0.091942 | 6.5783 | 24.8149 | -61.0937 | 112.225 | -152.795 | 154.1059 | -97.7205 | 27.4773 | [12 empty] | 1.2716 | 4.1451 | 11.8646 | -3.9295 | 0.041358
[7 empty] | 2405 | 13.4 | [20 empty] | -0.091942 | 6.581 | 24.8355 | -61.17 | 112.4119 | -153.1131 | 154.4909 | -98.0054 | 27.5689 | [12 empty] | 1.2716 | 4.1468 | 11.8745 | -3.9344 | 0.041427
[7 empty] | 2406 | 13.4067 | [20 empty] | -0.091942 | 6.5837 | 24.8562 | -61.2463 | 112.599 | -153.4317 | 154.8768 | -98.291 | 27.6608 | [12 empty] | 1.2716 | 4.1485 | 11.8844 | -3.9393 | 0.041495
[7 empty] | 2407 | 13.4135 | [20 empty] | -0.091942 | 6.5865 | 24.8768 | -61.3227 | 112.7863 | -153.7508 | 155.2634 | -98.5773 | 27.7529 | [12 empty] | 1.2716 | 4.1502 | 11.8943 | -3.9442 | 0.041565
[7 empty] | 2408 | 13.4202 | [20 empty] | -0.091942 | 6.5892 | 24.8975 | -61.3991 | 112.9738 | -154.0705 | 155.6508 | -98.8644 | 27.8452 | [12 empty] | 1.2716 | 4.152 | 11.9042 | -3.9491 | 0.041634
[7 empty] | 2409 | 13.427 | [20 empty] | -0.091942 | 6.592 | 24.9182 | -61.4757 | 113.1616 | -154.3906 | 156.0391 | -99.1521 | 27.9379 | [12 empty] | 1.2716 | 4.1537 | 11.914 | -3.954 | 0.041703
[7 empty] | 2410 | 13.4338 | [20 empty] | -0.091942 | 6.5947 | 24.9389 | -61.5523 | 113.3496 | -154.7113 | 156.4281 | -99.4406 | 28.0308 | [12 empty] | 1.2716 | 4.1554 | 11.9239 | -3.959 | 0.041772
[7 empty] | 2411 | 13.4405 | [20 empty] | -0.091942 | 6.5974 | 24.9596 | -61.6289 | 113.5379 | -155.0326 | 156.818 | -99.7298 | 28.124 | [12 empty] | 1.2716 | 4.1571 | 11.9338 | -3.9639 | 0.041842
[7 empty] | 2412 | 13.4473 | [20 empty] | -0.091942 | 6.6002 | 24.9803 | -61.7056 | 113.7264 | -155.3544 | 157.2086 | -100.0197 | 28.2174 | [12 empty] | 1.2716 | 4.1589 | 11.9437 | -3.9688 | 0.041911
[7 empty] | 2413 | 13.454 | [20 empty] | -0.091942 | 6.6029 | 25.001 | -61.7824 | 113.9151 | -155.6767 | 157.6001 | -100.3103 | 28.3112 | [12 empty] | 1.2716 | 4.1606 | 11.9536 | -3.9738 | 0.041981
[7 empty] | 2414 | 13.4608 | [20 empty] | -0.091942 | 6.6056 | 25.0217 | -61.8593 | 114.104 | -155.9995 | 157.9924 | -100.6017 | 28.4052 | [12 empty] | 1.2716 | 4.1623 | 11.9636 | -3.9787 | 0.04205
[7 empty] | 2415 | 13.4676 | [20 empty] | -0.091942 | 6.6084 | 25.0425 | -61.9362 | 114.2932 | -156.3229 | 158.3855 | -100.8938 | 28.4994 | [12 empty] | 1.2716 | 4.164 | 11.9735 | -3.9837 | 0.04212
[7 empty] | 2416 | 13.4743 | [20 empty] | -0.091942 | 6.6111 | 25.0632 | -62.0131 | 114.4826 | -156.6468 | 158.7794 | -101.1866 | 28.594 | [12 empty] | 1.2716 | 4.1658 | 11.9834 | -3.9886 | 0.04219
[7 empty] | 2417 | 13.4811 | [20 empty] | -0.091942 | 6.6138 | 25.084 | -62.0902 | 114.6723 | -156.9713 | 159.1741 | -101.4801 | 28.6888 | [12 empty] | 1.2716 | 4.1675 | 11.9933 | -3.9936 | 0.04226
[7 empty] | 2418 | 13.4878 | [20 empty] | -0.091942 | 6.6166 | 25.1047 | -62.1673 | 114.8622 | -157.2963 | 159.5697 | -101.7744 | 28.7839 | [12 empty] | 1.2716 | 4.1692 | 12.0032 | -3.9985 | 0.04233
[7 empty] | 2419 | 13.4946 | [20 empty] | -0.091942 | 6.6193 | 25.1255 | -62.2444 | 115.0523 | -157.6218 | 159.966 | -102.0694 | 28.8793 | [12 empty] | 1.2716 | 4.1709 | 12.0132 | -4.0035 | 0.0424
[7 empty] | 2420 | 13.5014 | [20 empty] | -0.091942 | 6.6221 | 25.1463 | -62.3217 | 115.2427 | -157.9479 | 160.3632 | -102.3651 | 28.9749 | [12 empty] | 1.2716 | 4.1726 | 12.0231 | -4.0085 | 0.04247
[7 empty] | 2421 | 13.5081 | [20 empty] | -0.091942 | 6.6248 | 25.1671 | -62.3989 | 115.4333 | -158.2745 | 160.7612 | -102.6616 | 29.0708 | [12 empty] | 1.2716 | 4.1744 | 12.033 | -4.0134 | 0.04254
[7 empty] | 2422 | 13.5149 | [20 empty] | -0.091942 | 6.6275 | 25.1879 | -62.4763 | 115.6241 | -158.6016 | 161.16 | -102.9588 | 29.167 | [12 empty] | 1.2716 | 4.1761 | 12.043 | -4.0184 | 0.04261
[7 empty] | 2423 | 13.5217 | [20 empty] | -0.091942 | 6.6303 | 25.2087 | -62.5537 | 115.8152 | -158.9293 | 161.5597 | -103.2567 | 29.2635 | [12 empty] | 1.2716 | 4.1778 | 12.0529 | -4.0234 | 0.042681
[7 empty] | 2424 | 13.5284 | [20 empty] | -0.091942 | 6.633 | 25.2295 | -62.6312 | 116.0065 | -159.2576 | 161.9602 | -103.5554 | 29.3603 | [12 empty] | 1.2716 | 4.1795 | 12.0629 | -4.0284 | 0.042751
[7 empty] | 2425 | 13.5352 | [20 empty] | -0.091942 | 6.6357 | 25.2503 | -62.7087 | 116.1981 | -159.5863 | 162.3615 | -103.8548 | 29.4573 | [12 empty] | 1.2716 | 4.1813 | 12.0728 | -4.0334 | 0.042822
[7 empty] | 2426 | 13.5419 | [20 empty] | -0.091942 | 6.6385 | 25.2711 | -62.7864 | 116.3899 | -159.9156 | 162.7636 | -104.155 | 29.5546 | [12 empty] | 1.2716 | 4.183 | 12.0828 | -4.0383 | 0.042893
[7 empty] | 2427 | 13.5487 | [20 empty] | -0.091942 | 6.6412 | 25.292 | -62.864 | 116.5819 | -160.2455 | 163.1666 | -104.4559 | 29.6522 | [12 empty] | 1.2716 | 4.1847 | 12.0928 | -4.0433 | 0.042963
[7 empty] | 2428 | 13.5555 | [20 empty] | -0.091942 | 6.6439 | 25.3128 | -62.9418 | 116.7741 | -160.5759 | 163.5704 | -104.7575 | 29.7501 | [12 empty] | 1.2716 | 4.1864 | 12.1027 | -4.0483 | 0.043034
[7 empty] | 2429 | 13.5622 | [20 empty] | -0.091942 | 6.6467 | 25.3337 | -63.0196 | 116.9666 | -160.9069 | 163.975 | -105.0599 | 29.8483 | [12 empty] | 1.2716 | 4.1882 | 12.1127 | -4.0533 | 0.043105
[7 empty] | 2430 | 13.569 | [20 empty] | -0.091942 | 6.6494 | 25.3545 | -63.0974 | 117.1594 | -161.2383 | 164.3805 | -105.363 | 29.9467 | [12 empty] | 1.2716 | 4.1899 | 12.1227 | -4.0584 | 0.043176
[7 empty] | 2431 | 13.5758 | [20 empty] | -0.091942 | 6.6522 | 25.3754 | -63.1754 | 117.3523 | -161.5704 | 164.7867 | -105.6669 | 30.0454 | [12 empty] | 1.2716 | 4.1916 | 12.1326 | -4.0634 | 0.043247
[7 empty] | 2432 | 13.5825 | [20 empty] | -0.091942 | 6.6549 | 25.3963 | -63.2534 | 117.5456 | -161.903 | 165.1939 | -105.9716 | 30.1445 | [12 empty] | 1.2716 | 4.1933 | 12.1426 | -4.0684 | 0.043318
[7 empty] | 2433 | 13.5893 | [20 empty] | -0.091942 | 6.6576 | 25.4172 | -63.3314 | 117.739 | -162.2361 | 165.6019 | -106.277 | 30.2438 | [12 empty] | 1.2716 | 4.1951 | 12.1526 | -4.0734 | 0.04339
[7 empty] | 2434 | 13.596 | [20 empty] | -0.091942 | 6.6604 | 25.4381 | -63.4095 | 117.9327 | -162.5698 | 166.0107 | -106.5831 | 30.3434 | [12 empty] | 1.2716 | 4.1968 | 12.1626 | -4.0784 | 0.043461
[7 empty] | 2435 | 13.6028 | [20 empty] | -0.091942 | 6.6631 | 25.459 | -63.4877 | 118.1266 | -162.904 | 166.4203 | -106.89 | 30.4432 | [12 empty] | 1.2716 | 4.1985 | 12.1726 | -4.0835 | 0.043533
[7 empty] | 2436 | 13.6096 | [20 empty] | -0.091942 | 6.6658 | 25.4799 | -63.566 | 118.3208 | -163.2388 | 166.8308 | -107.1977 | 30.5434 | [12 empty] | 1.2716 | 4.2002 | 12.1826 | -4.0885 | 0.043604
[7 empty] | 2437 | 13.6163 | [20 empty] | -0.091942 | 6.6686 | 25.5008 | -63.6443 | 118.5152 | -163.5741 | 167.2421 | -107.5061 | 30.6438 | [12 empty] | 1.2716 | 4.202 | 12.1926 | -4.0935 | 0.043676
[7 empty] | 2438 | 13.6231 | [20 empty] | -0.091942 | 6.6713 | 25.5217 | -63.7227 | 118.7098 | -163.91 | 167.6543 | -107.8153 | 30.7446 | [12 empty] | 1.2716 | 4.2037 | 12.2026 | -4.0986 | 0.043748
[7 empty] | 2439 | 13.6299 | [20 empty] | -0.091942 | 6.674 | 25.5427 | -63.8011 | 118.9047 | -164.2464 | 168.0673 | -108.1252 | 30.8456 | [12 empty] | 1.2716 | 4.2054 | 12.2126 | -4.1036 | 0.043819
[7 empty] | 2440 | 13.6366 | [20 empty] | -0.091942 | 6.6768 | 25.5636 | -63.8796 | 119.0999 | -164.5834 | 168.4812 | -108.4359 | 30.9469 | [12 empty] | 1.2716 | 4.2071 | 12.2226 | -4.1087 | 0.043891
[7 empty] | 2441 | 13.6434 | [20 empty] | -0.091942 | 6.6795 | 25.5846 | -63.9582 | 119.2952 | -164.921 | 168.8959 | -108.7474 | 31.0485 | [12 empty] | 1.2716 | 4.2089 | 12.2327 | -4.1137 | 0.043963
[7 empty] | 2442 | 13.6501 | [20 empty] | -0.091942 | 6.6823 | 25.6056 | -64.0368 | 119.4908 | -165.2591 | 169.3115 | -109.0596 | 31.1504 | [12 empty] | 1.2716 | 4.2106 | 12.2427 | -4.1188 | 0.044035
[7 empty] | 2443 | 13.6569 | [20 empty] | -0.091942 | 6.685 | 25.6265 | -64.1155 | 119.6867 | -165.5977 | 169.7279 | -109.3726 | 31.2526 | [12 empty] | 1.2716 | 4.2123 | 12.2527 | -4.1238 | 0.044107
[7 empty] | 2444 | 13.6637 | [20 empty] | -0.091942 | 6.6877 | 25.6475 | -64.1943 | 119.8828 | -165.9369 | 170.1452 | -109.6864 | 31.3551 | [12 empty] | 1.2716 | 4.214 | 12.2628 | -4.1289 | 0.04418
[7 empty] | 2445 | 13.6704 | [20 empty] | -0.091942 | 6.6905 | 25.6685 | -64.2731 | 120.0791 | -166.2767 | 170.5633 | -110.0009 | 31.4579 | [12 empty] | 1.2716 | 4.2158 | 12.2728 | -4.134 | 0.044252
[7 empty] | 2446 | 13.6772 | [20 empty] | -0.091942 | 6.6932 | 25.6895 | -64.352 | 120.2757 | -166.617 | 170.9823 | -110.3163 | 31.561 | [12 empty] | 1.2716 | 4.2175 | 12.2828 | -4.139 | 0.044325
[7 empty] | 2447 | 13.684 | [20 empty] | -0.091942 | 6.6959 | 25.7105 | -64.431 | 120.4725 | -166.9578 | 171.4022 | -110.6323 | 31.6644 | [12 empty] | 1.2716 | 4.2192 | 12.2929 | -4.1441 | 0.044397
[7 empty] | 2448 | 13.6907 | [20 empty] | -0.091942 | 6.6987 | 25.7315 | -64.51 | 120.6695 | -167.2993 | 171.8229 | -110.9492 | 31.768 | [12 empty] | 1.2716 | 4.2209 | 12.3029 | -4.1492 | 0.04447
[7 empty] | 2449 | 13.6975 | [20 empty] | -0.091942 | 6.7014 | 25.7526 | -64.5891 | 120.8668 | -167.6412 | 172.2445 | -111.2669 | 31.872 | [12 empty] | 1.2716 | 4.2227 | 12.313 | -4.1543 | 0.044542
[7 empty] | 2450 | 13.7043 | [20 empty] | -0.091942 | 6.7041 | 25.7736 | -64.6683 | 121.0644 | -167.9838 | 172.6669 | -111.5853 | 31.9763 | [12 empty] | 1.2716 | 4.2244 | 12.323 | -4.1594 | 0.044615
[7 empty] | 2451 | 13.711 | [20 empty] | -0.091942 | 6.7069 | 25.7946 | -64.7475 | 121.2621 | -168.3269 | 173.0902 | -111.9045 | 32.0808 | [12 empty] | 1.2716 | 4.2261 | 12.3331 | -4.1645 | 0.044688
[7 empty] | 2452 | 13.7178 | [20 empty] | -0.091942 | 6.7096 | 25.8157 | -64.8268 | 121.4601 | -168.6706 | 173.5143 | -112.2245 | 32.1857 | [12 empty] | 1.2716 | 4.2278 | 12.3432 | -4.1696 | 0.044761
[7 empty] | 2453 | 13.7245 | [20 empty] | -0.091942 | 6.7124 | 25.8368 | -64.9061 | 121.6584 | -169.0148 | 173.9393 | -112.5453 | 32.2908 | [12 empty] | 1.2716 | 4.2295 | 12.3532 | -4.1747 | 0.044834
[7 empty] | 2454 | 13.7313 | [20 empty] | -0.091942 | 6.7151 | 25.8578 | -64.9855 | 121.8569 | -169.3596 | 174.3652 | -112.8668 | 32.3963 | [12 empty] | 1.2716 | 4.2313 | 12.3633 | -4.1798 | 0.044907
[7 empty] | 2455 | 13.7381 | [20 empty] | -0.091942 | 6.7178 | 25.8789 | -65.065 | 122.0557 | -169.7049 | 174.792 | -113.1892 | 32.5021 | [12 empty] | 1.2716 | 4.233 | 12.3734 | -4.1849 | 0.044981
[7 empty] | 2456 | 13.7448 | [20 empty] | -0.091942 | 6.7206 | 25.9 | -65.1445 | 122.2547 | -170.0508 | 175.2196 | -113.5123 | 32.6081 | [12 empty] | 1.2716 | 4.2347 | 12.3835 | -4.19 | 0.045054
[7 empty] | 2457 | 13.7516 | [20 empty] | -0.091942 | 6.7233 | 25.9211 | -65.2241 | 122.4539 | -170.3973 | 175.6481 | -113.8362 | 32.7145 | [12 empty] | 1.2716 | 4.2364 | 12.3936 | -4.1951 | 0.045127
[7 empty] | 2458 | 13.7584 | [20 empty] | -0.091942 | 6.726 | 25.9422 | -65.3038 | 122.6534 | -170.7444 | 176.0775 | -114.1609 | 32.8212 | [12 empty] | 1.2716 | 4.2382 | 12.4036 | -4.2003 | 0.045201
[7 empty] | 2459 | 13.7651 | [20 empty] | -0.091942 | 6.7288 | 25.9633 | -65.3835 | 122.8531 | -171.092 | 176.5077 | -114.4864 | 32.9281 | [12 empty] | 1.2716 | 4.2399 | 12.4137 | -4.2054 | 0.045274
[7 empty] | 2460 | 13.7719 | [20 empty] | -0.091942 | 6.7315 | 25.9844 | -65.4634 | 123.053 | -171.4401 | 176.9388 | -114.8127 | 33.0354 | [12 empty] | 1.2716 | 4.2416 | 12.4238 | -4.2105 | 0.045348
[7 empty] | 2461 | 13.7787 | [20 empty] | -0.091942 | 6.7342 | 26.0056 | -65.5432 | 123.2533 | -171.7889 | 177.3708 | -115.1398 | 33.143 | [12 empty] | 1.2716 | 4.2433 | 12.4339 | -4.2157 | 0.045422
[7 empty] | 2462 | 13.7854 | [20 empty] | -0.091942 | 6.737 | 26.0267 | -65.6231 | 123.4537 | -172.1382 | 177.8037 | -115.4677 | 33.2509 | [12 empty] | 1.2716 | 4.2451 | 12.444 | -4.2208 | 0.045496
[7 empty] | 2463 | 13.7922 | [20 empty] | -0.091942 | 6.7397 | 26.0478 | -65.7031 | 123.6544 | -172.4881 | 178.2375 | -115.7964 | 33.3591 | [12 empty] | 1.2716 | 4.2468 | 12.4542 | -4.226 | 0.04557
[7 empty] | 2464 | 13.7989 | [20 empty] | -0.091942 | 6.7425 | 26.069 | -65.7832 | 123.8553 | -172.8385 | 178.6721 | -116.1259 | 33.4676 | [12 empty] | 1.2716 | 4.2485 | 12.4643 | -4.2311 | 0.045644
[7 empty] | 2465 | 13.8057 | [20 empty] | -0.091942 | 6.7452 | 26.0902 | -65.8633 | 124.0565 | -173.1895 | 179.1076 | -116.4563 | 33.5764 | [12 empty] | 1.2716 | 4.2502 | 12.4744 | -4.2363 | 0.045718
[7 empty] | 2466 | 13.8125 | [20 empty] | -0.091942 | 6.7479 | 26.1113 | -65.9435 | 124.258 | -173.5411 | 179.544 | -116.7874 | 33.6855 | [12 empty] | 1.2716 | 4.252 | 12.4845 | -4.2414 | 0.045792
[7 empty] | 2467 | 13.8192 | [20 empty] | -0.091942 | 6.7507 | 26.1325 | -66.0238 | 124.4596 | -173.8932 | 179.9813 | -117.1193 | 33.795 | [12 empty] | 1.2716 | 4.2537 | 12.4946 | -4.2466 | 0.045866
[7 empty] | 2468 | 13.826 | [20 empty] | -0.091942 | 6.7534 | 26.1537 | -66.1041 | 124.6616 | -174.246 | 180.4195 | -117.452 | 33.9047 | [12 empty] | 1.2716 | 4.2554 | 12.5048 | -4.2517 | 0.045941
[7 empty] | 2469 | 13.8328 | [20 empty] | -0.091942 | 6.7561 | 26.1749 | -66.1845 | 124.8637 | -174.5993 | 180.8586 | -117.7855 | 34.0148 | [12 empty] | 1.2716 | 4.2571 | 12.5149 | -4.2569 | 0.046015
[7 empty] | 2470 | 13.8395 | [20 empty] | -0.091942 | 6.7589 | 26.1961 | -66.2649 | 125.0661 | -174.9531 | 181.2985 | -118.1199 | 34.1252 | [12 empty] | 1.2716 | 4.2589 | 12.5251 | -4.2621 | 0.04609
[7 empty] | 2471 | 13.8463 | [20 empty] | -0.091942 | 6.7616 | 26.2173 | -66.3455 | 125.2688 | -175.3076 | 181.7394 | -118.455 | 34.2358 | [12 empty] | 1.2716 | 4.2606 | 12.5352 | -4.2673 | 0.046165
[7 empty] | 2472 | 13.8531 | [20 empty] | -0.091942 | 6.7643 | 26.2386 | -66.426 | 125.4717 | -175.6626 | 182.1811 | -118.791 | 34.3468 | [12 empty] | 1.2716 | 4.2623 | 12.5453 | -4.2724 | 0.046239
[7 empty] | 2473 | 13.8598 | [20 empty] | -0.091942 | 6.7671 | 26.2598 | -66.5067 | 125.6749 | -176.0182 | 182.6237 | -119.1278 | 34.4581 | [12 empty] | 1.2716 | 4.264 | 12.5555 | -4.2776 | 0.046314
[7 empty] | 2474 | 13.8666 | [20 empty] | -0.091942 | 6.7698 | 26.281 | -66.5874 | 125.8783 | -176.3744 | 183.0673 | -119.4654 | 34.5698 | [12 empty] | 1.2716 | 4.2658 | 12.5656 | -4.2828 | 0.046389
[7 empty] | 2475 | 13.8733 | [20 empty] | -0.091942 | 6.7726 | 26.3023 | -66.6682 | 126.0819 | -176.7311 | 183.5117 | -119.8038 | 34.6817 | [12 empty] | 1.2716 | 4.2675 | 12.5758 | -4.288 | 0.046464
[7 empty] | 2476 | 13.8801 | [20 empty] | -0.091942 | 6.7753 | 26.3235 | -66.749 | 126.2858 | -177.0884 | 183.957 | -120.1431 | 34.794 | [12 empty] | 1.2716 | 4.2692 | 12.586 | -4.2932 | 0.046539
[7 empty] | 2477 | 13.8869 | [20 empty] | -0.091942 | 6.778 | 26.3448 | -66.8299 | 126.4899 | -177.4463 | 184.4033 | -120.4832 | 34.9066 | [12 empty] | 1.2716 | 4.2709 | 12.5961 | -4.2984 | 0.046615
[7 empty] | 2478 | 13.8936 | [20 empty] | -0.091942 | 6.7808 | 26.3661 | -66.9109 | 126.6943 | -177.8048 | 184.8504 | -120.8241 | 35.0195 | [12 empty] | 1.2716 | 4.2727 | 12.6063 | -4.3036 | 0.04669
[7 empty] | 2479 | 13.9004 | [20 empty] | -0.091942 | 6.7835 | 26.3874 | -66.9919 | 126.899 | -178.1639 | 185.2984 | -121.1658 | 35.1327 | [12 empty] | 1.2716 | 4.2744 | 12.6165 | -4.3088 | 0.046765
[7 empty] | 2480 | 13.9072 | [20 empty] | -0.091942 | 6.7862 | 26.4087 | -67.073 | 127.1038 | -178.5235 | 185.7473 | -121.5083 | 35.2462 | [12 empty] | 1.2716 | 4.2761 | 12.6267 | -4.3141 | 0.046841
[7 empty] | 2481 | 13.9139 | [20 empty] | -0.091942 | 6.789 | 26.43 | -67.1542 | 127.309 | -178.8837 | 186.1972 | -121.8517 | 35.3601 | [12 empty] | 1.2716 | 4.2778 | 12.6369 | -4.3193 | 0.046916
[7 empty] | 2482 | 13.9207 | [20 empty] | -0.091942 | 6.7917 | 26.4513 | -67.2354 | 127.5143 | -179.2445 | 186.6479 | -122.1959 | 35.4742 | [12 empty] | 1.2716 | 4.2796 | 12.647 | -4.3245 | 0.046992
[7 empty] | 2483 | 13.9275 | [20 empty] | -0.091942 | 6.7944 | 26.4726 | -67.3167 | 127.72 | -179.6059 | 187.0996 | -122.541 | 35.5887 | [12 empty] | 1.2716 | 4.2813 | 12.6572 | -4.3297 | 0.047068
[7 empty] | 2484 | 13.9342 | [20 empty] | -0.091942 | 6.7972 | 26.4939 | -67.3981 | 127.9258 | -179.9678 | 187.5522 | -122.8869 | 35.7036 | [12 empty] | 1.2716 | 4.283 | 12.6674 | -4.335 | 0.047144
[7 empty] | 2485 | 13.941 | [20 empty] | -0.091942 | 6.7999 | 26.5152 | -67.4795 | 128.132 | -180.3304 | 188.0056 | -123.2336 | 35.8187 | [12 empty] | 1.2716 | 4.2847 | 12.6776 | -4.3402 | 0.04722
[7 empty] | 2486 | 13.9478 | [20 empty] | -0.091942 | 6.8027 | 26.5366 | -67.561 | 128.3383 | -180.6935 | 188.46 | -123.5811 | 35.9342 | [12 empty] | 1.2716 | 4.2864 | 12.6878 | -4.3454 | 0.047296
[7 empty] | 2487 | 13.9545 | [20 empty] | -0.091942 | 6.8054 | 26.5579 | -67.6426 | 128.545 | -181.0572 | 188.9153 | -123.9295 | 36.05 | [12 empty] | 1.2716 | 4.2882 | 12.6981 | -4.3507 | 0.047372
[7 empty] | 2488 | 13.9613 | [20 empty] | -0.091942 | 6.8081 | 26.5793 | -67.7242 | 128.7518 | -181.4215 | 189.3716 | -124.2788 | 36.1661 | [12 empty] | 1.2716 | 4.2899 | 12.7083 | -4.3559 | 0.047448
[7 empty] | 2489 | 13.968 | [20 empty] | -0.091942 | 6.8109 | 26.6007 | -67.8059 | 128.959 | -181.7864 | 189.8287 | -124.6289 | 36.2826 | [12 empty] | 1.2716 | 4.2916 | 12.7185 | -4.3612 | 0.047525
[7 empty] | 2490 | 13.9748 | [20 empty] | -0.091942 | 6.8136 | 26.6221 | -67.8877 | 129.1663 | -182.1519 | 190.2868 | -124.9798 | 36.3994 | [12 empty] | 1.2716 | 4.2933 | 12.7287 | -4.3665 | 0.047601
[7 empty] | 2491 | 13.9816 | [20 empty] | -0.091942 | 6.8163 | 26.6434 | -67.9695 | 129.374 | -182.518 | 190.7458 | -125.3316 | 36.5165 | [12 empty] | 1.2716 | 4.2951 | 12.7389 | -4.3717 | 0.047677
[7 empty] | 2492 | 13.9883 | [20 empty] | -0.091942 | 6.8191 | 26.6648 | -68.0514 | 129.5818 | -182.8846 | 191.2057 | -125.6842 | 36.6339 | [12 empty] | 1.2716 | 4.2968 | 12.7492 | -4.377 | 0.047754
[7 empty] | 2493 | 13.9951 | [20 empty] | -0.091942 | 6.8218 | 26.6862 | -68.1334 | 129.7899 | -183.2518 | 191.6665 | -126.0376 | 36.7517 | [12 empty] | 1.2716 | 4.2985 | 12.7594 | -4.3823 | 0.047831
[7 empty] | 2494 | 14.0019 | [20 empty] | -0.091942 | 6.8245 | 26.7077 | -68.2154 | 129.9983 | -183.6197 | 192.1282 | -126.392 | 36.8698 | [12 empty] | 1.2716 | 4.3002 | 12.7696 | -4.3875 | 0.047908
[7 empty] | 2495 | 14.0086 | [20 empty] | -0.091942 | 6.8273 | 26.7291 | -68.2975 | 130.2069 | -183.9881 | 192.5909 | -126.7471 | 36.9882 | [12 empty] | 1.2716 | 4.302 | 12.7799 | -4.3928 | 0.047984
[7 empty] | 2496 | 14.0154 | [20 empty] | -0.091942 | 6.83 | 26.7505 | -68.3796 | 130.4158 | -184.3571 | 193.0545 | -127.1032 | 37.107 | [12 empty] | 1.2716 | 4.3037 | 12.7901 | -4.3981 | 0.048061
[7 empty] | 2497 | 14.0222 | [20 empty] | -0.091942 | 6.8328 | 26.772 | -68.4618 | 130.6249 | -184.7267 | 193.5191 | -127.4601 | 37.2261 | [12 empty] | 1.2716 | 4.3054 | 12.8004 | -4.4034 | 0.048139
[7 empty] | 2498 | 14.0289 | [20 empty] | -0.091942 | 6.8355 | 26.7934 | -68.5441 | 130.8343 | -185.0969 | 193.9845 | -127.8178 | 37.3455 | [12 empty] | 1.2716 | 4.3071 | 12.8106 | -4.4087 | 0.048216
[7 empty] | 2499 | 14.0357 | [20 empty] | -0.091942 | 6.8382 | 26.8149 | -68.6265 | 131.0439 | -185.4677 | 194.4509 | -128.1764 | 37.4653 | [12 empty] | 1.2716 | 4.3089 | 12.8209 | -4.414 | 0.048293
[7 empty] | 2500 | 14.0424 | [20 empty] | -0.091942 | 6.841 | 26.8363 | -68.7089 | 131.2538 | -185.8391 | 194.9183 | -128.5359 | 37.5854 | [12 empty] | 1.2716 | 4.3106 | 12.8311 | -4.4193 | 0.04837
[7 empty] | 2501 | 14.0492 | [20 empty] | -0.091942 | 6.8437 | 26.8578 | -68.7914 | 131.464 | -186.211 | 195.3866 | -128.8962 | 37.7058 | [12 empty] | 1.2716 | 4.3123 | 12.8414 | -4.4246 | 0.048448
[7 empty] | 2502 | 14.056 | [20 empty] | -0.091942 | 6.8464 | 26.8793 | -68.8739 | 131.6743 | -186.5836 | 195.8558 | -129.2574 | 37.8266 | [12 empty] | 1.2716 | 4.314 | 12.8517 | -4.4299 | 0.048525
[7 empty] | 2503 | 14.0627 | [20 empty] | -0.091942 | 6.8492 | 26.9008 | -68.9565 | 131.885 | -186.9568 | 196.3259 | -129.6195 | 37.9477 | [12 empty] | 1.2716 | 4.3158 | 12.862 | -4.4352 | 0.048603
[7 empty] | 2504 | 14.0695 | [20 empty] | -0.091942 | 6.8519 | 26.9223 | -69.0392 | 132.0959 | -187.3305 | 196.797 | -129.9824 | 38.0692 | [12 empty] | 1.2716 | 4.3175 | 12.8722 | -4.4405 | 0.048681
[7 empty] | 2505 | 14.0763 | [20 empty] | -0.091942 | 6.8546 | 26.9438 | -69.122 | 132.307 | -187.7049 | 197.269 | -130.3462 | 38.191 | [12 empty] | 1.2716 | 4.3192 | 12.8825 | -4.4458 | 0.048758
[7 empty] | 2506 | 14.083 | [20 empty] | -0.091942 | 6.8574 | 26.9653 | -69.2048 | 132.5184 | -188.0799 | 197.742 | -130.7109 | 38.3131 | [12 empty] | 1.2716 | 4.3209 | 12.8928 | -4.4512 | 0.048836
[7 empty] | 2507 | 14.0898 | [20 empty] | -0.091942 | 6.8601 | 26.9868 | -69.2877 | 132.7301 | -188.4554 | 198.2159 | -131.0765 | 38.4356 | [12 empty] | 1.2716 | 4.3227 | 12.9031 | -4.4565 | 0.048914
[7 empty] | 2508 | 14.0966 | [20 empty] | -0.091942 | 6.8629 | 27.0083 | -69.3706 | 132.942 | -188.8316 | 198.6908 | -131.4429 | 38.5584 | [12 empty] | 1.2716 | 4.3244 | 12.9134 | -4.4618 | 0.048992
[7 empty] | 2509 | 14.1033 | [20 empty] | -0.091942 | 6.8656 | 27.0299 | -69.4536 | 133.1541 | -189.2083 | 199.1666 | -131.8102 | 38.6816 | [12 empty] | 1.2716 | 4.3261 | 12.9237 | -4.4672 | 0.049071
[7 empty] | 2510 | 14.1101 | [20 empty] | -0.091942 | 6.8683 | 27.0514 | -69.5367 | 133.3665 | -189.5857 | 199.6434 | -132.1784 | 38.8051 | [12 empty] | 1.2716 | 4.3278 | 12.934 | -4.4725 | 0.049149
[7 empty] | 2511 | 14.1168 | [20 empty] | -0.091942 | 6.8711 | 27.073 | -69.6198 | 133.5792 | -189.9637 | 200.1211 | -132.5474 | 38.9289 | [12 empty] | 1.2716 | 4.3296 | 12.9443 | -4.4779 | 0.049227
[7 empty] | 2512 | 14.1236 | [20 empty] | -0.091942 | 6.8738 | 27.0946 | -69.7031 | 133.7921 | -190.3422 | 200.5997 | -132.9174 | 39.0531 | [12 empty] | 1.2716 | 4.3313 | 12.9546 | -4.4832 | 0.049306
[7 empty] | 2513 | 14.1304 | [20 empty] | -0.091942 | 6.8765 | 27.1161 | -69.7863 | 134.0053 | -190.7214 | 201.0793 | -133.2882 | 39.1777 | [12 empty] | 1.2716 | 4.333 | 12.9649 | -4.4886 | 0.049384
[7 empty] | 2514 | 14.1371 | [20 empty] | -0.091942 | 6.8793 | 27.1377 | -69.8697 | 134.2187 | -191.1012 | 201.5599 | -133.6599 | 39.3026 | [12 empty] | 1.2716 | 4.3347 | 12.9753 | -4.4939 | 0.049463
[7 empty] | 2515 | 14.1439 | [20 empty] | -0.091942 | 6.882 | 27.1593 | -69.9531 | 134.4324 | -191.4815 | 202.0414 | -134.0325 | 39.4278 | [12 empty] | 1.2716 | 4.3365 | 12.9856 | -4.4993 | 0.049542
[7 empty] | 2516 | 14.1507 | [20 empty] | -0.091942 | 6.8847 | 27.1809 | -70.0366 | 134.6463 | -191.8625 | 202.5239 | -134.406 | 39.5534 | [12 empty] | 1.2716 | 4.3382 | 12.9959 | -4.5047 | 0.04962
[7 empty] | 2517 | 14.1574 | [20 empty] | -0.091942 | 6.8875 | 27.2025 | -70.1201 | 134.8605 | -192.2441 | 203.0074 | -134.7804 | 39.6794 | [12 empty] | 1.2716 | 4.3399 | 13.0062 | -4.51 | 0.049699
[7 empty] | 2518 | 14.1642 | [20 empty] | -0.091942 | 6.8902 | 27.2242 | -70.2037 | 135.075 | -192.6263 | 203.4918 | -135.1557 | 39.8057 | [12 empty] | 1.2716 | 4.3416 | 13.0166 | -4.5154 | 0.049778
[7 empty] | 2519 | 14.1709 | [20 empty] | -0.091942 | 6.893 | 27.2458 | -70.2874 | 135.2897 | -193.0091 | 203.9772 | -135.5319 | 39.9323 | [12 empty] | 1.2716 | 4.3433 | 13.0269 | -4.5208 | 0.049858
[7 empty] | 2520 | 14.1777 | [20 empty] | -0.091942 | 6.8957 | 27.2674 | -70.3711 | 135.5046 | -193.3925 | 204.4635 | -135.909 | 40.0593 | [12 empty] | 1.2716 | 4.3451 | 13.0373 | -4.5262 | 0.049937
[7 empty] | 2521 | 14.1845 | [20 empty] | -0.091942 | 6.8984 | 27.2891 | -70.4549 | 135.7198 | -193.7765 | 204.9508 | -136.2869 | 40.1866 | [12 empty] | 1.2716 | 4.3468 | 13.0476 | -4.5316 | 0.050016
[7 empty] | 2522 | 14.1912 | [20 empty] | -0.091942 | 6.9012 | 27.3107 | -70.5388 | 135.9353 | -194.1612 | 205.4391 | -136.6658 | 40.3143 | [12 empty] | 1.2716 | 4.3485 | 13.058 | -4.537 | 0.050096
[7 empty] | 2523 | 14.198 | [20 empty] | -0.091942 | 6.9039 | 27.3324 | -70.6228 | 136.151 | -194.5464 | 205.9283 | -137.0456 | 40.4424 | [12 empty] | 1.2716 | 4.3502 | 13.0683 | -4.5424 | 0.050175
[7 empty] | 2524 | 14.2048 | [20 empty] | -0.091942 | 6.9066 | 27.354 | -70.7068 | 136.367 | -194.9323 | 206.4185 | -137.4263 | 40.5708 | [12 empty] | 1.2716 | 4.352 | 13.0787 | -4.5478 | 0.050255
[7 empty] | 2525 | 14.2115 | [20 empty] | -0.091942 | 6.9094 | 27.3757 | -70.7908 | 136.5833 | -195.3187 | 206.9097 | -137.8079 | 40.6996 | [12 empty] | 1.2716 | 4.3537 | 13.0891 | -4.5532 | 0.050334
[7 empty] | 2526 | 14.2183 | [20 empty] | -0.091942 | 6.9121 | 27.3974 | -70.875 | 136.7998 | -195.7058 | 207.4018 | -138.1904 | 40.8287 | [12 empty] | 1.2716 | 4.3554 | 13.0994 | -4.5586 | 0.050414
[7 empty] | 2527 | 14.225 | [20 empty] | -0.091942 | 6.9148 | 27.4191 | -70.9592 | 137.0165 | -196.0935 | 207.895 | -138.5738 | 40.9582 | [12 empty] | 1.2716 | 4.3571 | 13.1098 | -4.564 | 0.050494
[7 empty] | 2528 | 14.2318 | [20 empty] | -0.091942 | 6.9176 | 27.4408 | -71.0435 | 137.2335 | -196.4818 | 208.3891 | -138.9581 | 41.088 | [12 empty] | 1.2716 | 4.3589 | 13.1202 | -4.5694 | 0.050574
[7 empty] | 2529 | 14.2386 | [20 empty] | -0.091942 | 6.9203 | 27.4625 | -71.1278 | 137.4508 | -196.8707 | 208.8842 | -139.3433 | 41.2183 | [12 empty] | 1.2716 | 4.3606 | 13.1306 | -4.5749 | 0.050654
[7 empty] | 2530 | 14.2453 | [20 empty] | -0.091942 | 6.9231 | 27.4843 | -71.2122 | 137.6683 | -197.2602 | 209.3802 | -139.7295 | 41.3488 | [12 empty] | 1.2716 | 4.3623 | 13.1409 | -4.5803 | 0.050734
[7 empty] | 2531 | 14.2521 | [20 empty] | -0.091942 | 6.9258 | 27.506 | -71.2967 | 137.8861 | -197.6504 | 209.8773 | -140.1165 | 41.4798 | [12 empty] | 1.2716 | 4.364 | 13.1513 | -4.5857 | 0.050814
[7 empty] | 2532 | 14.2588 | [20 empty] | -0.091942 | 6.9285 | 27.5277 | -71.3812 | 138.1042 | -198.0412 | 210.3753 | -140.5045 | 41.611 | [12 empty] | 1.2716 | 4.3658 | 13.1617 | -4.5912 | 0.050895
[7 empty] | 2533 | 14.2656 | [20 empty] | -0.091942 | 6.9313 | 27.5495 | -71.4658 | 138.3225 | -198.4325 | 210.8743 | -140.8934 | 41.7427 | [12 empty] | 1.2716 | 4.3675 | 13.1721 | -4.5966 | 0.050975
[7 empty] | 2534 | 14.2724 | [20 empty] | -0.091942 | 6.934 | 27.5712 | -71.5505 | 138.541 | -198.8246 | 211.3743 | -141.2832 | 41.8747 | [12 empty] | 1.2716 | 4.3692 | 13.1825 | -4.602 | 0.051056
[7 empty] | 2535 | 14.2791 | [20 empty] | -0.091942 | 6.9367 | 27.593 | -71.6353 | 138.7598 | -199.2172 | 211.8753 | -141.674 | 42.0071 | [12 empty] | 1.2716 | 4.3709 | 13.1929 | -4.6075 | 0.051136
[7 empty] | 2536 | 14.2859 | [20 empty] | -0.091942 | 6.9395 | 27.6148 | -71.7201 | 138.9789 | -199.6104 | 212.3773 | -142.0657 | 42.1398 | [12 empty] | 1.2716 | 4.3727 | 13.2033 | -4.613 | 0.051217
[7 empty] | 2537 | 14.2926 | [20 empty] | -0.091942 | 6.9422 | 27.6366 | -71.8049 | 139.1983 | -200.0043 | 212.8802 | -142.4583 | 42.273 | [12 empty] | 1.2716 | 4.3744 | 13.2138 | -4.6184 | 0.051298
[7 empty] | 2538 | 14.2994 | [20 empty] | -0.091942 | 6.945 | 27.6583 | -71.8899 | 139.4179 | -200.3988 | 213.3842 | -142.8518 | 42.4064 | [12 empty] | 1.2716 | 4.3761 | 13.2242 | -4.6239 | 0.051379
[7 empty] | 2539 | 14.3062 | [20 empty] | -0.091942 | 6.9477 | 27.6801 | -71.9749 | 139.6377 | -200.7939 | 213.8891 | -143.2462 | 42.5403 | [12 empty] | 1.2716 | 4.3778 | 13.2346 | -4.6293 | 0.05146
[7 empty] | 2540 | 14.3129 | [20 empty] | -0.091942 | 6.9504 | 27.702 | -72.06 | 139.8578 | -201.1896 | 214.3951 | -143.6416 | 42.6745 | [12 empty] | 1.2716 | 4.3796 | 13.245 | -4.6348 | 0.051541
[7 empty] | 2541 | 14.3197 | [20 empty] | -0.091942 | 6.9532 | 27.7238 | -72.1451 | 140.0782 | -201.586 | 214.902 | -144.038 | 42.8091 | [12 empty] | 1.2716 | 4.3813 | 13.2555 | -4.6403 | 0.051622
[7 empty] | 2542 | 14.3264 | [20 empty] | -0.091942 | 6.9559 | 27.7456 | -72.2303 | 140.2989 | -201.9829 | 215.41 | -144.4352 | 42.9441 | [12 empty] | 1.2716 | 4.383 | 13.2659 | -4.6458 | 0.051704
[7 empty] | 2543 | 14.3332 | [20 empty] | -0.091942 | 6.9586 | 27.7674 | -72.3156 | 140.5198 | -202.3805 | 215.9189 | -144.8334 | 43.0794 | [12 empty] | 1.2716 | 4.3847 | 13.2763 | -4.6513 | 0.051785
[7 empty] | 2544 | 14.34 | [20 empty] | -0.091942 | 6.9614 | 27.7893 | -72.4009 | 140.7409 | -202.7788 | 216.4289 | -145.2326 | 43.2151 | [12 empty] | 1.2716 | 4.3865 | 13.2868 | -4.6567 | 0.051866
[7 empty] | 2545 | 14.3467 | [20 empty] | -0.091942 | 6.9641 | 27.8111 | -72.4864 | 140.9623 | -203.1776 | 216.9398 | -145.6327 | 43.3512 | [12 empty] | 1.2716 | 4.3882 | 13.2972 | -4.6622 | 0.051948
[7 empty] | 2546 | 14.3535 | [20 empty] | -0.091942 | 6.9668 | 27.833 | -72.5718 | 141.184 | -203.5771 | 217.4518 | -146.0337 | 43.4877 | [12 empty] | 1.2716 | 4.3899 | 13.3077 | -4.6677 | 0.05203
[7 empty] | 2547 | 14.3602 | [20 empty] | -0.091942 | 6.9696 | 27.8548 | -72.6574 | 141.406 | -203.9772 | 217.9647 | -146.4357 | 43.6245 | [12 empty] | 1.2716 | 4.3916 | 13.3181 | -4.6732 | 0.052112
[7 empty] | 2548 | 14.367 | [20 empty] | -0.091942 | 6.9723 | 27.8767 | -72.743 | 141.6282 | -204.378 | 218.4787 | -146.8386 | 43.7617 | [12 empty] | 1.2716 | 4.3934 | 13.3286 | -4.6787 | 0.052193
[7 empty] | 2549 | 14.3738 | [20 empty] | -0.091942 | 6.9751 | 27.8986 | -72.8287 | 141.8506 | -204.7793 | 218.9937 | -147.2425 | 43.8993 | [12 empty] | 1.2716 | 4.3951 | 13.3391 | -4.6843 | 0.052275
[7 empty] | 2550 | 14.3805 | [20 empty] | -0.091942 | 6.9778 | 27.9205 | -72.9144 | 142.0734 | -205.1813 | 219.5096 | -147.6473 | 44.0373 | [12 empty] | 1.2716 | 4.3968 | 13.3495 | -4.6898 | 0.052358
[7 empty] | 2551 | 14.3873 | [20 empty] | -0.091942 | 6.9805 | 27.9424 | -73.0002 | 142.2964 | -205.584 | 220.0266 | -148.0531 | 44.1756 | [12 empty] | 1.2716 | 4.3985 | 13.36 | -4.6953 | 0.05244
[7 empty] | 2552 | 14.394 | [20 empty] | -0.091942 | 6.9833 | 27.9643 | -73.0861 | 142.5196 | -205.9872 | 220.5447 | -148.4599 | 44.3143 | [12 empty] | 1.2716 | 4.4002 | 13.3705 | -4.7008 | 0.052522
[7 empty] | 2553 | 14.4008 | [20 empty] | -0.091942 | 6.986 | 27.9862 | -73.1721 | 142.7431 | -206.3911 | 221.0637 | -148.8676 | 44.4535 | [12 empty] | 1.2716 | 4.402 | 13.381 | -4.7063 | 0.052604
[7 empty] | 2554 | 14.4076 | [20 empty] | -0.091942 | 6.9887 | 28.0082 | -73.2581 | 142.9669 | -206.7957 | 221.5837 | -149.2762 | 44.5929 | [12 empty] | 1.2716 | 4.4037 | 13.3914 | -4.7119 | 0.052687
[7 empty] | 2555 | 14.4143 | [20 empty] | -0.091942 | 6.9915 | 28.0301 | -73.3442 | 143.1909 | -207.2008 | 222.1048 | -149.6858 | 44.7328 | [12 empty] | 1.2716 | 4.4054 | 13.4019 | -4.7174 | 0.052769
[7 empty] | 2556 | 14.4211 | [20 empty] | -0.091942 | 6.9942 | 28.0521 | -73.4303 | 143.4153 | -207.6066 | 222.6269 | -150.0964 | 44.8731 | [12 empty] | 1.2716 | 4.4071 | 13.4124 | -4.723 | 0.052852
[7 empty] | 2557 | 14.4278 | [20 empty] | -0.091942 | 6.9969 | 28.074 | -73.5165 | 143.6398 | -208.0131 | 223.15 | -150.5079 | 45.0137 | [12 empty] | 1.2716 | 4.4089 | 13.4229 | -4.7285 | 0.052935
[7 empty] | 2558 | 14.4346 | [20 empty] | -0.091942 | 6.9997 | 28.096 | -73.6028 | 143.8647 | -208.4201 | 223.6741 | -150.9205 | 45.1547 | [12 empty] | 1.2716 | 4.4106 | 13.4334 | -4.7341 | 0.053018
[7 empty] | 2559 | 14.4414 | [20 empty] | -0.091942 | 7.0024 | 28.1179 | -73.6892 | 144.0897 | -208.8278 | 224.1993 | -151.3339 | 45.2961 | [12 empty] | 1.2716 | 4.4123 | 13.4439 | -4.7396 | 0.053101
[7 empty] | 2560 | 14.4481 | [20 empty] | -0.091942 | 7.0052 | 28.1399 | -73.7756 | 144.3151 | -209.2362 | 224.7255 | -151.7484 | 45.4379 | [12 empty] | 1.2716 | 4.414 | 13.4544 | -4.7452 | 0.053184
[7 empty] | 2561 | 14.4549 | [20 empty] | -0.091942 | 7.0079 | 28.1619 | -73.8621 | 144.5407 | -209.6452 | 225.2527 | -152.1638 | 45.5801 | [12 empty] | 1.2716 | 4.4158 | 13.4649 | -4.7507 | 0.053267
[7 empty] | 2562 | 14.4616 | [20 empty] | -0.091942 | 7.0106 | 28.1839 | -73.9487 | 144.7666 | -210.0548 | 225.7809 | -152.5802 | 45.7227 | [12 empty] | 1.2716 | 4.4175 | 13.4755 | -4.7563 | 0.05335
[7 empty] | 2563 | 14.4684 | [20 empty] | -0.091942 | 7.0134 | 28.2059 | -74.0353 | 144.9928 | -210.4651 | 226.3102 | -152.9976 | 45.8657 | [12 empty] | 1.2716 | 4.4192 | 13.486 | -4.7619 | 0.053433
[7 empty] | 2564 | 14.4751 | [20 empty] | -0.091942 | 7.0161 | 28.2279 | -74.122 | 145.2192 | -210.876 | 226.8405 | -153.4159 | 46.009 | [12 empty] | 1.2716 | 4.4209 | 13.4965 | -4.7674 | 0.053517
[7 empty] | 2565 | 14.4819 | [20 empty] | -0.091942 | 7.0188 | 28.2499 | -74.2087 | 145.4459 | -211.2875 | 227.3719 | -153.8353 | 46.1528 | [12 empty] | 1.2716 | 4.4227 | 13.507 | -4.773 | 0.0536
[7 empty] | 2566 | 14.4887 | [20 empty] | -0.091942 | 7.0216 | 28.272 | -74.2956 | 145.6728 | -211.6997 | 227.9043 | -154.2556 | 46.2969 | [12 empty] | 1.2716 | 4.4244 | 13.5176 | -4.7786 | 0.053684
[7 empty] | 2567 | 14.4954 | [20 empty] | -0.091942 | 7.0243 | 28.294 | -74.3825 | 145.9 | -212.1125 | 228.4377 | -154.6769 | 46.4415 | [12 empty] | 1.2716 | 4.4261 | 13.5281 | -4.7842 | 0.053768
[7 empty] | 2568 | 14.5022 | [20 empty] | -0.091942 | 7.027 | 28.3161 | -74.4694 | 146.1275 | -212.526 | 228.9721 | -155.0992 | 46.5864 | [12 empty] | 1.2716 | 4.4278 | 13.5387 | -4.7898 | 0.053852
[7 empty] | 2569 | 14.5089 | [20 empty] | -0.091942 | 7.0298 | 28.3381 | -74.5565 | 146.3553 | -212.9401 | 229.5076 | -155.5224 | 46.7317 | [12 empty] | 1.2716 | 4.4296 | 13.5492 | -4.7954 | 0.053936
[7 empty] | 2570 | 14.5157 | [20 empty] | -0.091942 | 7.0325 | 28.3602 | -74.6435 | 146.5833 | -213.3549 | 230.0442 | -155.9467 | 46.8774 | [12 empty] | 1.2716 | 4.4313 | 13.5598 | -4.801 | 0.05402
[7 empty] | 2571 | 14.5224 | [20 empty] | -0.091942 | 7.0353 | 28.3823 | -74.7307 | 146.8116 | -213.7703 | 230.5818 | -156.372 | 47.0236 | [12 empty] | 1.2716 | 4.433 | 13.5703 | -4.8066 | 0.054104
[7 empty] | 2572 | 14.5292 | [20 empty] | -0.091942 | 7.038 | 28.4043 | -74.8179 | 147.0401 | -214.1863 | 231.1204 | -156.7982 | 47.1701 | [12 empty] | 1.2716 | 4.4347 | 13.5809 | -4.8122 | 0.054188
[7 empty] | 2573 | 14.536 | [20 empty] | -0.091942 | 7.0407 | 28.4264 | -74.9053 | 147.2689 | -214.6031 | 231.6601 | -157.2255 | 47.317 | [12 empty] | 1.2716 | 4.4365 | 13.5914 | -4.8178 | 0.054272
[7 empty] | 2574 | 14.5427 | [20 empty] | -0.091942 | 7.0435 | 28.4485 | -74.9926 | 147.498 | -215.0204 | 232.2008 | -157.6537 | 47.4643 | [12 empty] | 1.2716 | 4.4382 | 13.602 | -4.8234 | 0.054357
[7 empty] | 2575 | 14.5495 | [20 empty] | -0.091942 | 7.0462 | 28.4707 | -75.0801 | 147.7273 | -215.4384 | 232.7426 | -158.0829 | 47.612 | [12 empty] | 1.2716 | 4.4399 | 13.6126 | -4.8291 | 0.054441
[7 empty] | 2576 | 14.5562 | [20 empty] | -0.091942 | 7.0489 | 28.4928 | -75.1676 | 147.957 | -215.8571 | 233.2855 | -158.5132 | 47.7602 | [12 empty] | 1.2716 | 4.4416 | 13.6231 | -4.8347 | 0.054526
[7 empty] | 2577 | 14.563 | [20 empty] | -0.091942 | 7.0517 | 28.5149 | -75.2551 | 148.1868 | -216.2764 | 233.8294 | -158.9444 | 47.9087 | [12 empty] | 1.2716 | 4.4434 | 13.6337 | -4.8403 | 0.054611
[7 empty] | 2578 | 14.5697 | [20 empty] | -0.091942 | 7.0544 | 28.537 | -75.3428 | 148.417 | -216.6963 | 234.3743 | -159.3767 | 48.0576 | [12 empty] | 1.2716 | 4.4451 | 13.6443 | -4.846 | 0.054695
[7 empty] | 2579 | 14.5765 | [20 empty] | -0.091942 | 7.0571 | 28.5592 | -75.4305 | 148.6474 | -217.1169 | 234.9203 | -159.8099 | 48.207 | [12 empty] | 1.2716 | 4.4468 | 13.6549 | -4.8516 | 0.05478
[7 empty] | 2580 | 14.5832 | [20 empty] | -0.091942 | 7.0599 | 28.5813 | -75.5183 | 148.8781 | -217.5382 | 235.4674 | -160.2442 | 48.3567 | [12 empty] | 1.2716 | 4.4485 | 13.6655 | -4.8572 | 0.054865
[7 empty] | 2581 | 14.59 | [20 empty] | -0.091942 | 7.0626 | 28.6035 | -75.6061 | 149.109 | -217.9601 | 236.0155 | -160.6795 | 48.5068 | [12 empty] | 1.2716 | 4.4503 | 13.6761 | -4.8629 | 0.05495
[7 empty] | 2582 | 14.5967 | [20 empty] | -0.091942 | 7.0654 | 28.6257 | -75.694 | 149.3403 | -218.3827 | 236.5647 | -161.1158 | 48.6574 | [12 empty] | 1.2716 | 4.452 | 13.6867 | -4.8686 | 0.055036
[7 empty] | 2583 | 14.6035 | [20 empty] | -0.091942 | 7.0681 | 28.6478 | -75.782 | 149.5717 | -218.8059 | 237.115 | -161.5531 | 48.8084 | [12 empty] | 1.2716 | 4.4537 | 13.6973 | -4.8742 | 0.055121
[7 empty] | 2584 | 14.6103 | [20 empty] | -0.091942 | 7.0708 | 28.67 | -75.8701 | 149.8035 | -219.2298 | 237.6663 | -161.9914 | 48.9597 | [12 empty] | 1.2716 | 4.4554 | 13.7079 | -4.8799 | 0.055206
[7 empty] | 2585 | 14.617 | [20 empty] | -0.091942 | 7.0736 | 28.6922 | -75.9582 | 150.0355 | -219.6543 | 238.2187 | -162.4307 | 49.1115 | [12 empty] | 1.2716 | 4.4571 | 13.7185 | -4.8855 | 0.055292
[7 empty] | 2586 | 14.6238 | [20 empty] | -0.091942 | 7.0763 | 28.7144 | -76.0464 | 150.2678 | -220.0795 | 238.7721 | -162.8711 | 49.2637 | [12 empty] | 1.2716 | 4.4589 | 13.7291 | -4.8912 | 0.055377
[7 empty] | 2587 | 14.6305 | [20 empty] | -0.091942 | 7.079 | 28.7366 | -76.1346 | 150.5004 | -220.5053 | 239.3267 | -163.3125 | 49.4163 | [12 empty] | 1.2716 | 4.4606 | 13.7397 | -4.8969 | 0.055463
[7 empty] | 2588 | 14.6373 | [20 empty] | -0.091942 | 7.0818 | 28.7588 | -76.2229 | 150.7332 | -220.9318 | 239.8823 | -163.7549 | 49.5693 | [12 empty] | 1.2716 | 4.4623 | 13.7504 | -4.9026 | 0.055549
[7 empty] | 2589 | 14.644 | [20 empty] | -0.091942 | 7.0845 | 28.7811 | -76.3113 | 150.9663 | -221.359 | 240.439 | -164.1983 | 49.7228 | [12 empty] | 1.2716 | 4.464 | 13.761 | -4.9083 | 0.055635
[7 empty] | 2590 | 14.6508 | [20 empty] | -0.091942 | 7.0872 | 28.8033 | -76.3998 | 151.1997 | -221.7868 | 240.9967 | -164.6428 | 49.8766 | [12 empty] | 1.2716 | 4.4658 | 13.7716 | -4.9139 | 0.055721
[7 empty] | 2591 | 14.6575 | [20 empty] | -0.091942 | 7.09 | 28.8256 | -76.4883 | 151.4334 | -222.2153 | 241.5555 | -165.0883 | 50.0309 | [12 empty] | 1.2716 | 4.4675 | 13.7823 | -4.9196 | 0.055807
[7 empty] | 2592 | 14.6643 | [20 empty] | -0.091942 | 7.0927 | 28.8478 | -76.5769 | 151.6673 | -222.6445 | 242.1155 | -165.5348 | 50.1856 | [12 empty] | 1.2716 | 4.4692 | 13.7929 | -4.9253 | 0.055893
[7 empty] | 2593 | 14.671 | [20 empty] | -0.091942 | 7.0955 | 28.8701 | -76.6656 | 151.9015 | -223.0743 | 242.6764 | -165.9824 | 50.3407 | [12 empty] | 1.2716 | 4.4709 | 13.8035 | -4.931 | 0.055979
[7 empty] | 2594 | 14.6778 | [20 empty] | -0.091942 | 7.0982 | 28.8924 | -76.7543 | 152.1359 | -223.5048 | 243.2385 | -166.431 | 50.4962 | [12 empty] | 1.2716 | 4.4727 | 13.8142 | -4.9368 | 0.056066
[7 empty] | 2595 | 14.6845 | [20 empty] | -0.091942 | 7.1009 | 28.9146 | -76.8431 | 152.3707 | -223.9359 | 243.8017 | -166.8806 | 50.6521 | [12 empty] | 1.2716 | 4.4744 | 13.8248 | -4.9425 | 0.056152
[7 empty] | 2596 | 14.6913 | [20 empty] | -0.091942 | 7.1037 | 28.9369 | -76.932 | 152.6057 | -224.3677 | 244.3659 | -167.3313 | 50.8085 | [12 empty] | 1.2716 | 4.4761 | 13.8355 | -4.9482 | 0.056239
[7 empty] | 2597 | 14.698 | [20 empty] | -0.091942 | 7.1064 | 28.9592 | -77.0209 | 152.841 | -224.8002 | 244.9313 | -167.783 | 50.9653 | [12 empty] | 1.2716 | 4.4778 | 13.8462 | -4.9539 | 0.056326
[7 empty] | 2598 | 14.7048 | [20 empty] | -0.091942 | 7.1091 | 28.9815 | -77.1099 | 153.0765 | -225.2333 | 245.4977 | -168.2358 | 51.1225 | [12 empty] | 1.2716 | 4.4796 | 13.8568 | -4.9596 | 0.056412
[7 empty] | 2599 | 14.7115 | [20 empty] | -0.091942 | 7.1119 | 29.0038 | -77.199 | 153.3123 | -225.6672 | 246.0652 | -168.6896 | 51.2801 | [12 empty] | 1.2716 | 4.4813 | 13.8675 | -4.9654 | 0.056499
[7 empty] | 2600 | 14.7183 | [20 empty] | -0.091942 | 7.1146 | 29.0262 | -77.2882 | 153.5484 | -226.1016 | 246.6338 | -169.1445 | 51.4382 | [12 empty] | 1.2716 | 4.483 | 13.8782 | -4.9711 | 0.056586
[7 empty] | 2601 | 14.725 | [20 empty] | -0.091942 | 7.1173 | 29.0485 | -77.3774 | 153.7848 | -226.5368 | 247.2035 | -169.6004 | 51.5967 | [12 empty] | 1.2716 | 4.4847 | 13.8888 | -4.9768 | 0.056673
[7 empty] | 2602 | 14.7318 | [20 empty] | -0.091942 | 7.1201 | 29.0708 | -77.4667 | 154.0214 | -226.9726 | 247.7743 | -170.0573 | 51.7556 | [12 empty] | 1.2716 | 4.4865 | 13.8995 | -4.9826 | 0.056761
[7 empty] | 2603 | 14.7385 | [20 empty] | -0.091942 | 7.1228 | 29.0932 | -77.556 | 154.2583 | -227.4091 | 248.3462 | -170.5154 | 51.9149 | [12 empty] | 1.2716 | 4.4882 | 13.9102 | -4.9883 | 0.056848
[7 empty] | 2604 | 14.7453 | [20 empty] | -0.091942 | 7.1256 | 29.1155 | -77.6454 | 154.4955 | -227.8462 | 248.9192 | -170.9744 | 52.0747 | [12 empty] | 1.2716 | 4.4899 | 13.9209 | -4.9941 | 0.056935
[7 empty] | 2605 | 14.752 | [20 empty] | -0.091942 | 7.1283 | 29.1379 | -77.7349 | 154.733 | -228.2841 | 249.4933 | -171.4346 | 52.2349 | [12 empty] | 1.2716 | 4.4916 | 13.9316 | -4.9998 | 0.057023
[7 empty] | 2606 | 14.7588 | [20 empty] | -0.091942 | 7.131 | 29.1603 | -77.8245 | 154.9707 | -228.7226 | 250.0685 | -171.8958 | 52.3955 | [12 empty] | 1.2716 | 4.4934 | 13.9423 | -5.0056 | 0.057111
[7 empty] | 2607 | 14.7655 | [20 empty] | -0.091942 | 7.1338 | 29.1827 | -77.9141 | 155.2087 | -229.1617 | 250.6448 | -172.358 | 52.5566 | [12 empty] | 1.2716 | 4.4951 | 13.953 | -5.0113 | 0.057198
[7 empty] | 2608 | 14.7723 | [20 empty] | -0.091942 | 7.1365 | 29.2051 | -78.0038 | 155.447 | -229.6016 | 251.2222 | -172.8214 | 52.7181 | [12 empty] | 1.2716 | 4.4968 | 13.9637 | -5.0171 | 0.057286
[7 empty] | 2609 | 14.779 | [20 empty] | -0.091942 | 7.1392 | 29.2275 | -78.0935 | 155.6855 | -230.0421 | 251.8007 | -173.2858 | 52.88 | [12 empty] | 1.2716 | 4.4985 | 13.9744 | -5.0229 | 0.057374
[7 empty] | 2610 | 14.7858 | [20 empty] | -0.091942 | 7.142 | 29.2499 | -78.1834 | 155.9244 | -230.4833 | 252.3804 | -173.7512 | 53.0424 | [12 empty] | 1.2716 | 4.5003 | 13.9851 | -5.0287 | 0.057462
[7 empty] | 2611 | 14.7925 | [20 empty] | -0.091942 | 7.1447 | 29.2723 | -78.2733 | 156.1635 | -230.9252 | 252.9611 | -174.2178 | 53.2052 | [12 empty] | 1.2716 | 4.502 | 13.9959 | -5.0344 | 0.05755
[7 empty] | 2612 | 14.7993 | [20 empty] | -0.091942 | 7.1474 | 29.2947 | -78.3632 | 156.4028 | -231.3678 | 253.543 | -174.6854 | 53.3684 | [12 empty] | 1.2716 | 4.5037 | 14.0066 | -5.0402 | 0.057638
[7 empty] | 2613 | 14.806 | [20 empty] | -0.091942 | 7.1502 | 29.3171 | -78.4533 | 156.6425 | -231.811 | 254.1259 | -175.1541 | 53.5321 | [12 empty] | 1.2716 | 4.5054 | 14.0173 | -5.046 | 0.057727
[7 empty] | 2614 | 14.8128 | [20 empty] | -0.091942 | 7.1529 | 29.3396 | -78.5434 | 156.8824 | -232.2549 | 254.71 | -175.6238 | 53.6962 | [12 empty] | 1.2716 | 4.5072 | 14.028 | -5.0518 | 0.057815
[7 empty] | 2615 | 14.8195 | [20 empty] | -0.091942 | 7.1557 | 29.362 | -78.6336 | 157.1226 | -232.6995 | 255.2952 | -176.0947 | 53.8608 | [12 empty] | 1.2716 | 4.5089 | 14.0388 | -5.0576 | 0.057904
[7 empty] | 2616 | 14.8263 | [20 empty] | -0.091942 | 7.1584 | 29.3845 | -78.7238 | 157.3631 | -233.1448 | 255.8815 | -176.5666 | 54.0258 | [12 empty] | 1.2716 | 4.5106 | 14.0495 | -5.0634 | 0.057992
[7 empty] | 2617 | 14.833 | [20 empty] | -0.091942 | 7.1611 | 29.407 | -78.8141 | 157.6039 | -233.5907 | 256.469 | -177.0396 | 54.1912 | [12 empty] | 1.2716 | 4.5123 | 14.0602 | -5.0692 | 0.058081
[7 empty] | 2618 | 14.8398 | [20 empty] | -0.091942 | 7.1639 | 29.4295 | -78.9045 | 157.8449 | -234.0374 | 257.0576 | -177.5137 | 54.3571 | [12 empty] | 1.2716 | 4.514 | 14.071 | -5.075 | 0.05817
[7 empty] | 2619 | 14.8465 | [20 empty] | -0.091942 | 7.1666 | 29.4519 | -78.995 | 158.0862 | -234.4847 | 257.6473 | -177.9889 | 54.5234 | [12 empty] | 1.2716 | 4.5158 | 14.0817 | -5.0809 | 0.058259
[7 empty] | 2620 | 14.8533 | [20 empty] | -0.091942 | 7.1693 | 29.4744 | -79.0855 | 158.3278 | -234.9327 | 258.2381 | -178.4651 | 54.6902 | [12 empty] | 1.2716 | 4.5175 | 14.0925 | -5.0867 | 0.058348
[7 empty] | 2621 | 14.86 | [20 empty] | -0.091942 | 7.1721 | 29.4969 | -79.1761 | 158.5696 | -235.3814 | 258.83 | -178.9425 | 54.8574 | [12 empty] | 1.2716 | 4.5192 | 14.1033 | -5.0925 | 0.058437
[7 empty] | 2622 | 14.8668 | [20 empty] | -0.091942 | 7.1748 | 29.5195 | -79.2667 | 158.8118 | -235.8307 | 259.4231 | -179.421 | 55.0251 | [12 empty] | 1.2716 | 4.5209 | 14.114 | -5.0983 | 0.058526
[7 empty] | 2623 | 14.8735 | [20 empty] | -0.091942 | 7.1775 | 29.542 | -79.3575 | 159.0542 | -236.2808 | 260.0173 | -179.9005 | 55.1932 | [12 empty] | 1.2716 | 4.5227 | 14.1248 | -5.1042 | 0.058615
[7 empty] | 2624 | 14.8803 | [20 empty] | -0.091942 | 7.1803 | 29.5645 | -79.4483 | 159.2969 | -236.7315 | 260.6127 | -180.3812 | 55.3617 | [12 empty] | 1.2716 | 4.5244 | 14.1356 | -5.11 | 0.058705
[7 empty] | 2625 | 14.887 | [20 empty] | -0.091942 | 7.183 | 29.587 | -79.5391 | 159.5398 | -237.183 | 261.2091 | -180.8629 | 55.5307 | [12 empty] | 1.2716 | 4.5261 | 14.1463 | -5.1159 | 0.058794
[7 empty] | 2626 | 14.8937 | [20 empty] | -0.091942 | 7.1858 | 29.6096 | -79.6301 | 159.7831 | -237.6351 | 261.8068 | -181.3458 | 55.7002 | [12 empty] | 1.2716 | 4.5278 | 14.1571 | -5.1217 | 0.058884
[7 empty] | 2627 | 14.9005 | [20 empty] | -0.091942 | 7.1885 | 29.6321 | -79.7211 | 160.0266 | -238.0879 | 262.4055 | -181.8297 | 55.8701 | [12 empty] | 1.2716 | 4.5296 | 14.1679 | -5.1276 | 0.058974
[7 empty] | 2628 | 14.9072 | [20 empty] | -0.091942 | 7.1912 | 29.6547 | -79.8121 | 160.2704 | -238.5414 | 263.0054 | -182.3148 | 56.0405 | [12 empty] | 1.2716 | 4.5313 | 14.1787 | -5.1334 | 0.059064
[7 empty] | 2629 | 14.914 | [20 empty] | -0.091942 | 7.194 | 29.6773 | -79.9033 | 160.5145 | -238.9956 | 263.6065 | -182.8009 | 56.2113 | [12 empty] | 1.2716 | 4.533 | 14.1895 | -5.1393 | 0.059154
[7 empty] | 2630 | 14.9207 | [20 empty] | -0.091942 | 7.1967 | 29.6999 | -79.9945 | 160.7589 | -239.4505 | 264.2086 | -183.2882 | 56.3826 | [12 empty] | 1.2716 | 4.5347 | 14.2003 | -5.1452 | 0.059244
[7 empty] | 2631 | 14.9275 | [20 empty] | -0.091942 | 7.1994 | 29.7225 | -80.0858 | 161.0035 | -239.9061 | 264.812 | -183.7766 | 56.5543 | [12 empty] | 1.2716 | 4.5365 | 14.2111 | -5.151 | 0.059334
[7 empty] | 2632 | 14.9342 | [20 empty] | -0.091942 | 7.2022 | 29.745 | -80.1771 | 161.2484 | -240.3623 | 265.4165 | -184.2661 | 56.7265 | [12 empty] | 1.2716 | 4.5382 | 14.2219 | -5.1569 | 0.059424
[7 empty] | 2633 | 14.941 | [20 empty] | -0.091942 | 7.2049 | 29.7677 | -80.2686 | 161.4936 | -240.8193 | 266.0221 | -184.7568 | 56.8992 | [12 empty] | 1.2716 | 4.5399 | 14.2327 | -5.1628 | 0.059514
[7 empty] | 2634 | 14.9477 | [20 empty] | -0.091942 | 7.2076 | 29.7903 | -80.3601 | 161.7391 | -241.2769 | 266.6289 | -185.2485 | 57.0723 | [12 empty] | 1.2716 | 4.5416 | 14.2435 | -5.1687 | 0.059605
[7 empty] | 2635 | 14.9544 | [20 empty] | -0.091942 | 7.2104 | 29.8129 | -80.4516 | 161.9849 | -241.7353 | 267.2368 | -185.7414 | 57.2458 | [12 empty] | 1.2716 | 4.5434 | 14.2543 | -5.1746 | 0.059695
[7 empty] | 2636 | 14.9612 | [20 empty] | -0.091942 | 7.2131 | 29.8355 | -80.5432 | 162.2309 | -242.1943 | 267.8459 | -186.2354 | 57.4199 | [12 empty] | 1.2716 | 4.5451 | 14.2652 | -5.1805 | 0.059786
[7 empty] | 2637 | 14.9679 | [20 empty] | -0.091942 | 7.2159 | 29.8582 | -80.6349 | 162.4772 | -242.6541 | 268.4561 | -186.7305 | 57.5944 | [12 empty] | 1.2716 | 4.5468 | 14.276 | -5.1863 | 0.059877
[7 empty] | 2638 | 14.9747 | [20 empty] | -0.091942 | 7.2186 | 29.8808 | -80.7267 | 162.7238 | -243.1145 | 269.0675 | -187.2267 | 57.7693 | [12 empty] | 1.2716 | 4.5485 | 14.2868 | -5.1923 | 0.059968
[7 empty] | 2639 | 14.9814 | [20 empty] | -0.091942 | 7.2213 | 29.9035 | -80.8186 | 162.9707 | -243.5757 | 269.6801 | -187.7241 | 57.9447 | [12 empty] | 1.2716 | 4.5503 | 14.2976 | -5.1982 | 0.060059
[7 empty] | 2640 | 14.9882 | [20 empty] | -0.091942 | 7.2241 | 29.9261 | -80.9105 | 163.2178 | -244.0375 | 270.2938 | -188.2226 | 58.1206 | [12 empty] | 1.2716 | 4.552 | 14.3085 | -5.2041 | 0.06015
[7 empty] | 2641 | 14.9949 | [20 empty] | -0.091942 | 7.2268 | 29.9488 | -81.0024 | 163.4653 | -244.5001 | 270.9087 | -188.7223 | 58.297 | [12 empty] | 1.2716 | 4.5537 | 14.3193 | -5.21 | 0.060241
[7 empty] | 2642 | 15.0016 | [20 empty] | -0.091942 | 7.2295 | 29.9715 | -81.0945 | 163.713 | -244.9633 | 271.5247 | -189.223 | 58.4738 | [12 empty] | 1.2716 | 4.5554 | 14.3302 | -5.2159 | 0.060332
[7 empty] | 2643 | 15.0084 | [20 empty] | -0.091942 | 7.2323 | 29.9942 | -81.1866 | 163.961 | -245.4272 | 272.142 | -189.725 | 58.6511 | [12 empty] | 1.2716 | 4.5572 | 14.341 | -5.2218 | 0.060424
[7 empty] | 2644 | 15.0151 | [20 empty] | -0.091942 | 7.235 | 30.0169 | -81.2788 | 164.2093 | -245.8919 | 272.7603 | -190.228 | 58.8289 | [12 empty] | 1.2716 | 4.5589 | 14.3519 | -5.2278 | 0.060515
[7 empty] | 2645 | 15.0219 | [20 empty] | -0.091942 | 7.2377 | 30.0396 | -81.3711 | 164.4579 | -246.3572 | 273.3799 | -190.7322 | 59.0071 | [12 empty] | 1.2716 | 4.5606 | 14.3627 | -5.2337 | 0.060607
[7 empty] | 2646 | 15.0286 | [20 empty] | -0.091942 | 7.2405 | 30.0623 | -81.4634 | 164.7067 | -246.8233 | 274.0006 | -191.2376 | 59.1858 | [12 empty] | 1.2716 | 4.5623 | 14.3736 | -5.2396 | 0.060698
[7 empty] | 2647 | 15.0353 | [20 empty] | -0.091942 | 7.2432 | 30.0851 | -81.5558 | 164.9558 | -247.2901 | 274.6225 | -191.7441 | 59.365 | [12 empty] | 1.2716 | 4.5641 | 14.3845 | -5.2456 | 0.06079
[7 empty] | 2648 | 15.0421 | [20 empty] | -0.091942 | 7.246 | 30.1078 | -81.6482 | 165.2053 | -247.7575 | 275.2456 | -192.2517 | 59.5447 | [12 empty] | 1.2716 | 4.5658 | 14.3953 | -5.2515 | 0.060882
[7 empty] | 2649 | 15.0488 | [20 empty] | -0.091942 | 7.2487 | 30.1305 | -81.7408 | 165.455 | -248.2257 | 275.8699 | -192.7605 | 59.7248 | [12 empty] | 1.2716 | 4.5675 | 14.4062 | -5.2575 | 0.060974
[7 empty] | 2650 | 15.0556 | [20 empty] | -0.091942 | 7.2514 | 30.1533 | -81.8334 | 165.7049 | -248.6946 | 276.4953 | -193.2705 | 59.9054 | [12 empty] | 1.2716 | 4.5692 | 14.4171 | -5.2634 | 0.061066
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0
[29 empty] | #NUM! | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | [12 empty] | #NUM! | 0 | 0 | 0 | 0

## interpollation formula,template
To interpolate using any number of points between two temperatures:
Solve for C(x) where
C is a known correction factor corresponding to a set point
S is a known set point
I is a desired interpollation point
A is Actual temperature at setpoint- not part of formula
C(x)= | (I1-S1)*(C2-C1)
| S2-S1 |  |  | +C1
|  | I1 | I(y) | I(z)
| S1 |  |  |  | S2
| A1 |  |  |  | A2
| C1 | C(x) | C(y) | C(z) | C2
EXAMPLE:
|  | 250 | 300 | 350
| 200 |  |  |  | 400
| 201 |  |  |  | 399
| -1 | C(x) = -0.5 |  |  | 1

## Temperatures
TTS-1 |  |  | SP-2 |  |  | SP-1
-50 | -18 | 1 | 800 | 799.8139 | 1 | 32 | 32.7333 | 1
32 | 64 | 2 | 1000 | 999.6454 | 2 | 200 | 201.0775 | 2
120 | 152 | 3 | 1200 | 1199.689 | 3 | 400 | 400.2841 | 3 | #VALUE!
[6 empty] | 600 | 599.934 | 4
NTS-11 |  |  | SP-2 |  |  | 800 | 800.0139 | 5
32 | 32.7318 | 1 | 32 | 32.7333 | 1 | 1000 | 1000.1554 | 6
70 | 70.7318 | 2 | 200 | 199.6775 | 2 | 1200 | 1200.839 | 7
200 | 200.7318 | 3 | 400 | 399.6341 | 3 | 1400 | 1400.17 | 8
400 | 32.7333 | 4 | 600 | 599.734 | 4 | 1600 | 1599.7961 | 9
|  |  | 800 | 799.8139 | 5 | 1800 | 1799.6955 | 10
SP-2 |  |  | 1000 | 999.6454 | 6 | 2000 | 2001.7175 | 11
400 | 199.6775 | 1 | 1200 | 1199.689 | 7 | 2200 | 2201.7619 | 12
600 | 399.6341 | 2 |  |  |  | 2400 | #VALUE! | 13
800 | 599.734 | 3

## Cert Page 1
Report: |  |  | PAGE | 1
Customer:
Item | #REF! |  |  | INSPECTOR: |  | 0
Type | 0 |  |  | Date Calibrated: |  | 0
Gauge: | 0 |  |  | Date Due: |  | #REF!
Temp: | 0 |  |  | Relative Humidity: |  | 0 | %
Location: | 0 |  |  | Tolerance: | 0
Controlling Specifications |  |  | ISO 10012, ANSI/NCSL Z540-1 |  |  | 0
Calibration/Inspection Procedure: |  |  | 0 | Tolerance Source: |  | 0 | See Below
Comments: | 0
Condition: | 0
Results: | 0
Calibration Standards
TEST INSTRUMENT: |  | 0 |  |  | TEST T/C:
DATE CERTIFIED: |  | 0 |  |  | CERTIFIED
DATE DUE: |  | 0 |  |  | DUE
TRACEABILITY: |  | 0 |  |  | TRACEABILITY:
TEST INSTRUMENT: |  | 0 |  |  | TEST T/C:
DATE CERTIFIED: |  | 0 |  |  | CERTIFIED
DATE DUE: |  | 0 |  |  | DUE
TRACEABILITY: |  | 0 |  |  | TRACEABILITY:

## Cert Page 2 Data
Report: |  |  | Johnson Gage & Inspection, Inc. | [5 empty] | Page 1 | Report Number | Date Due:
|  |  | 5920 W 21st St N. Wichita, Ks 67205 | [6 empty] | Customer | Relative Humidity
Thermocouple Results Report
[10 empty] | Item | Tolerance:
|  |  |  | Test Temperatures | [5 empty] | Inspector | Uncertainty Statement
|  |  | Correction Factors Shown Parenthetically | [6 empty] | Date Calibrated | Comments
T/C Type: | Serial No. | 2000.0°F |  | 2000.0°F |  | 4000.0°F | 5000.0°F | 6000.0°F |  |  | Condition
J | JGI1 | 2000.1°F | (-0.1°F) | 2000.1°F | (-0.1°F) | [5 empty] | Results
J | JGI2 | [9 empty] | Test Instrument
J Dual | JGI3A | [9 empty] | Date Certified
J Dual | JGI3B | [9 empty] | Due:
K | K1 | [9 empty] | TN
N | [10 empty] | Test TCs
N | [10 empty] | Dates
N | [10 empty] | Due:
N | [10 empty] | TN
S
T
T
T
T
T
R
R
E Dual
E Dual
K
T/C Type: | Serial No. | 7000.0°F |  | 8000.0°F |  | 9000.0°F | 100.0°F | 1100.0°F
J | JGI1 | 2000.1°F | (-0.1°F) | 2000.1°F | (-0.1°F)
J | JGI2
J Dual | JGI3A
J Dual | JGI3B
K
N
N
N
N
S
T
T
T
T
T
R
R
E Dual
E Dual
K

## Software Validation
Date: | Action
2015-03-16 | Validation record added, password protection enabled
2016-02-04 | Added Check for Difference between roll ends when doing wire per AMS2750 or BAC5621
2018-03-21 | Changed ST7 tables to be built off formula rather than imported. Verified against NIST generated tables.